## Measurement
This sheet uses correction data to fix bridge errors.
|  |  | Measured Z |  |  |  |  | Measured residual SHORT Z: |  |  |  |  | Corrected Z (results) | [5 empty] | Validation: ref data |  |  |  |  | Deviation of XLS calculated from Octave calculated: |  |  |  |  | Deviation with no correction: | [6 empty] | Measured Z (basic units) | [8 empty] | Correction data search for interpolation | [25 empty] | Interpolated correction |  |  |  | corrected Z | [13 empty] | Interpolated correction (gain-phase) |  |  |  | Corrected Z (gain-phase method)
Range |  | f | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Fixed Rs | Fixed U(Rs) | Fixed Xs | Fixed U(Xs) |  | Status | Valid Rs |  | Valid Xs |  |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | f |  | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | ua(¦Z¦) | phi | ua(phi) | lower Z ID' | lower ID' | lower Z' | lower Z ID' | lower Z ID' | lower ID'' | lower Z ID'' | lower Z ID'' | lower Z tag | higher Z tag | lower Z | higher Z | lower tag | higher tag | lower ID | higher ID | low check | low cor(Rs) | low U(cor(Rs)) | low cor(Xs) | low U(cor(Xs)) | high check | high cor(Rs) | high U(cor(Rs)) | high cor(Xs) | high U(cor(Xs)) | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | Rs | U(Rs) | Xs | U(Xs) | low cor(gain) | low U(cor(gain)) | low cor(phi) | low U(cor(phi)) | low status | hi cor(gain) | hi U(cor(gain)) | hi cor(phi) | hi U(cor(phi)) | low status | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi)) | ¦Z¦ | U(¦Z¦) | phi | U(phi) | Rs | U(Rs) | Xs | U(Xs)
Ohm |  | Hz | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm |  | - | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | - |  | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | Ohm |  | - |  |  |  |  | - | Ohm | Ohm | - | - | - | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | - | rad | rad | - | - | - | rad | rad | - | - | - | rad | rad | Ohm | Ohm | rad | rad | Ohm | Ohm | Ohm | Ohm
100 | m | 0.05 | 8.1931 | 0.0012406 | -60 | 0.00054652 | m | -0.00069503 | 7.4716e-05 | -0.0011689 | 0.00017239 | m | 8.2 | 0.012314 | -60 | 0.0020734 | m | OK | 8.2 |  | -60 |  | m | 1.114e-08 | 0.012314 | -5.7815e-08 | 0.0020734 | m | -0.0062072 |  | 0.0011323 |  | m | 50 | mHz | 0.001 | 0.0081938 | 1.2429e-06 | -0.059999 | 5.7307e-07 | 0.060556 | 5.9218e-07 | -1.4351 | 2.0376e-05 | 4 | 10m | 0.01 | 111 | 4 | 100m | 129 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | 4.7128e-07 | 0.0013032 | 9.9997e-07 | 0.0033159 | 50mHz100m100m | 2.0948e-06 | 0.0032386 | 9.8259e-06 | 0.0035729 | 1.3833e-06 | 0.0032386 | 5.9578e-06 | 0.0034602 | 0.0081952 | 0.0032386 | -0.059993 | 0.0034602 | 1 | 3.0828e-06 | 0.0001 | 0.00037978 | OK | 1 | 5.0094e-06 | 9.827e-05 | 6.1747e-05 | OK | 1 | 5.0094e-06 | 9.9028e-05 | 0.00020113 | 0.060558 | 1.2226e-06 | -1.435 | 0.00020522 | 0.0082 | 1.2314e-05 | -0.06 | 2.0734e-06
100 | m | 0.02 | 10.5004 | 0.00018086 | -1.522 | 0.0011015 | m | -0.00021085 | 0.0012594 | -0.00087241 | 0.0002868 | m | 10.5 | 0.0026187 | -1.52 | 0.0050837 | m | OK | 10.5 |  | -1.52 |  | m | 1.5489e-08 | 0.0026187 | -8.7701e-09 | 0.0050837 | m | 0.00063137 |  | -0.0011606 |  | m | 20 | mHz | 0.001 | 0.010501 | 1.2723e-06 | -0.0015212 | 1.1382e-06 | 0.01061 | 1.2697e-06 | -0.14386 | 0.00010755 | 4 | 10m | 0.01 | 110 | 4 | 100m | 128 | 4 | 10m | 100m | 0.01 | 0.1 | 20mHz100m10m | 20mHz100m100m | 110 | 128 | 20mHz100m10m | -7.5251e-07 | 0.0030346 | 1.0001e-06 | 0.0035718 | 20mHz100m100m | 2.19e-06 | 0.004003 | 6.056e-06 | 0.0038788 | -7.3256e-07 | 0.004003 | 1.0344e-06 | 0.0035739 | 0.0105 | 0.004003 | -0.0015201 | 0.0035739 | 0.99992 | 3.0673e-06 | 9.9999e-05 | 0.00043485 | OK | 1 | 4.3997e-06 | 6.0567e-05 | 6.8822e-05 | OK | 0.99993 | 4.3997e-06 | 9.9732e-05 | 0.00043237 | 0.010609 | 2.5399e-06 | -0.14376 | 0.00048292 | 0.0105 | 2.6187e-06 | -0.00152 | 5.0837e-06
3 | m | 500 | 0.68178 | 0.0019003 | -2.3408 | 0.0018747 | m | -0.00019735 | 0.00051439 | -0.00090367 | 0.00018622 | m | 0.682 | 0.0048102 | -2.34 | 0.0038794 | m | OK | 0.682 |  | -2.34 |  | m | -1.4912e-07 | 0.0048102 | -3.6277e-07 | 0.0038794 | m | -2.7173e-05 |  | 6.925e-05 |  | m | 500 | Hz | 0.001 | 0.00068197 | 1.9687e-06 | -0.0023399 | 1.8839e-06 | 0.0024373 | 1.8907e-06 | -1.2872 | 0.00080507 | 2 | 1m | 0.001 | 51 | 2 | 3m | 69 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | -8.6726e-08 | 0.00028009 | 9.9642e-08 | 0.0031946 | 500Hz3m3m | 2.2928e-07 | 0.002375 | -1.0533e-07 | 0.00067098 | 1.4037e-07 | 0.002375 | -4.7662e-08 | 0.001381 | 0.00068211 | 0.002375 | -0.00234 | 0.001381 | 0.99991 | 3.1107e-07 | 9.995e-05 | 0.0032165 | OK | 1.0001 | 2.419e-06 | -3.544e-05 | 0.0004039 | OK | 1 | 2.419e-06 | 2.6531e-06 | 0.0011953 | 0.0024374 | 3.7814e-06 | -1.2872 | 0.0020053 | 0.000682 | 4.8102e-06 | -0.00234 | 3.8794e-06
100 | m | 0.5 | -32.8015 | 0.00040882 | 18.204 | 0.00088916 | m | -0.0010127 | 2.2217e-05 | 0.0018411 | 0.00052011 | m | -32.8 | 0.0051185 | 18.2 | 0.0090337 | m | OK | -32.8 |  | 18.2 |  | m | -4.9825e-08 | 0.0051185 | -1.2193e-07 | 0.0090337 | m | -0.00046435 |  | 0.0021339 |  | m | 500 | mHz | 0.001 | -0.0328 | 4.0943e-07 | 0.018202 | 1.0301e-06 | 0.037513 | 6.1481e-07 | 2.635 | 2.4588e-05 | 4 | 10m | 0.01 | 114 | 4 | 100m | 132 | 4 | 10m | 100m | 0.01 | 0.1 | 500mHz100m10m | 500mHz100m100m | 114 | 132 | 500mHz100m10m | -5.3831e-07 | 0.0029439 | 1.0001e-06 | 0.0018521 | 500mHz100m100m | -3.4265e-07 | 0.004815 | -8.4037e-06 | 0.0015334 | -4.785e-07 | 0.004815 | -1.8746e-06 | 0.0017547 | -0.032801 | 0.004815 | 0.0182 | 0.0017547 | 0.99995 | 3.168e-06 | 9.9999e-05 | 0.00036198 | OK | 1 | 5.2289e-06 | -8.4043e-05 | 6.2294e-05 | OK | 0.99996 | 5.2289e-06 | 4.3739e-05 | 0.00027036 | 0.037511 | 1.2452e-06 | 2.635 | 0.0002748 | -0.0328 | 5.1185e-06 | 0.0182 | 9.0337e-06
10 | m | 5 | 9.4391 | 0.0019994 | -1.5901 | 2.1585e-05 | m | -9.8701e-05 | 0.0015825 | 0.00056466 | 0.00029453 | m | 9.44 | 0.005004 | -1.59 | 0.0040155 | m | OK | 9.44 |  | -1.59 |  | m | 3.5909e-08 | 0.005004 | -9.7134e-09 | 0.0040155 | m | -0.00077123 |  | -0.00071314 |  | m | 5 | Hz | 0.001 | 0.0094392 | 2.5499e-06 | -0.0015907 | 2.9532e-07 | 0.0095723 | 2.5149e-06 | -0.16695 | 5.3713e-05 | 3 | 3m | 0.003 | 81 | 3 | 10m | 99 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -1.2341e-07 | 0.00095268 | 2.9997e-07 | 0.002009 | 5Hz10m10m | 7.4124e-07 | 0.0040653 | 8.5979e-07 | 0.0011382 | 6.8842e-07 | 0.0040653 | 8.2559e-07 | 0.0011914 | 0.0094399 | 0.0040653 | -0.0015899 | 0.0011914 | 0.99996 | 3.0538e-06 | 0.0001 | 0.00085175 | OK | 1.0001 | 4.081e-06 | 8.5991e-05 | 0.00037269 | OK | 1.0001 | 4.081e-06 | 8.6847e-05 | 0.00040196 | 0.009573 | 5.03e-06 | -0.16687 | 0.00041606 | 0.00944 | 5.004e-06 | -0.00159 | 4.0155e-06
3 | m | 2 | -0.11984 | 0.00097637 | -1.0797 | 0.00067449 | m | 0.00026044 | 0.0012331 | 0.00022918 | 0.001218 | m | -0.12 | 0.0043497 | -1.08 | 0.0028137 | m | OK | -0.12 |  | -1.08 |  | m | 3.3516e-09 | 0.0043497 | -3.6516e-10 | 0.0028137 | m | -9.8708e-05 |  | 2.3376e-05 |  | m | 2 | Hz | 0.001 | -0.0001201 | 1.5728e-06 | -0.00108 | 1.3922e-06 | 0.0010866 | 1.3946e-06 | -1.6815 | 0.0014455 | 2 | 1m | 0.001 | 44 | 2 | 3m | 62 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | 1.1326e-08 | 0.0026983 | 9.9999e-08 | 0.0014561 | 2Hz3m3m | 3.5078e-08 | 0.0038536 | -2.0848e-07 | 0.0018577 | 1.2355e-08 | 0.0038536 | 8.6637e-08 | 0.0014735 | -0.00012009 | 0.0038536 | -0.0010799 | 0.0014735 | 1 | 2.7064e-06 | 0.0001 | 0.0028826 | OK | 1 | 5.491e-06 | -6.9494e-05 | 0.00073582 | OK | 1 | 5.491e-06 | 9.2661e-05 | 0.0027896 | 0.0010866 | 2.7892e-06 | -1.6815 | 0.0040174 | -0.00012 | 4.3497e-06 | -0.00108 | 2.8137e-06
10 | m | 0.05 | 2.6516 | 0.0016406 | -0.92175 | 0.00045425 | m | 0.0018352 | 0.0019786 | 0.00048296 | 0.0010709 | m | 2.65 | 0.0047466 | -0.922 | 0.0032757 | m | Extrapolated | 2.65 |  | -0.922 |  | m | -1.5433e-09 | 0.0047466 | -2.1915e-09 | 0.0032757 | m | -0.00020497 |  | -0.00023574 |  | m | 50 | mHz | 0.001 | 0.0026498 | 2.5703e-06 | -0.00092224 | 1.1633e-06 | 0.0028057 | 2.4574e-06 | -0.33493 | 0.00049397 | 2 | 1m | 0.001 | 0 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 1.2231e-07 | 0.002322 | 3.1418e-07 | 0.0012544 | 0.0026499 | 0.002322 | -0.00092192 | 0.0012544 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK | 1 | 3.906e-06 | 0.00010336 | 0.0004249 | 0.0028058 | 4.9148e-06 | -0.33482 | 0.0010754 | 0.00265 | 4.7466e-06 | -0.000922 | 3.2757e-06
3 | m | 0.05 | 0.52471 | 0.00086698 | -0.83448 | 0.00050097 | m | -0.00023638 | 0.0017056 | 0.00062783 | 0.0018335 | m | 0.525 | 0.0053053 | -0.835 | 0.0044607 | m | Extrapolated | 0.525 |  | -0.835 |  | m | 4.1392e-09 | 0.0053053 | 4.713e-09 | 0.0044607 | m | -4.8798e-05 |  | -0.00010945 |  | m | 50 | mHz | 0.001 | 0.00052495 | 1.9133e-06 | -0.00083511 | 1.9007e-06 | 0.0009864 | 1.9043e-06 | -1.0096 | 0.0019361 | 1 | 0m | 0 | 0 | 2 | 3m | 57 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | -6.7316e-08 | 0.0020391 | 1.0001e-07 | 0.0033297 | 50mHz3m3m | -7.9542e-08 | 0.0030995 | -1.1468e-07 | 0.0020253 | -6.7233e-08 | 0.0030995 | 1.0147e-07 | 0.0033386 | 0.00052488 | 0.0030995 | -0.00083501 | 0.0033386 | 0.99993 | 2.3594e-06 | 0.00010001 | 0.0043908 | OK | 0.99997 | 3.1024e-06 | -3.8223e-05 | 0.001011 | OK | 0.99993 | 3.1024e-06 | 0.00010095 | 0.0044138 | 0.00098633 | 3.8086e-06 | -1.0095 | 0.0058716 | 0.000525 | 5.3053e-06 | -0.000835 | 4.4607e-06
3 | m | 1000 | 1.4587 | 0.0014853 | 0.063331 | 0.00072724 | m | -0.0012451 | 0.0012363 | 0.0019374 | 0.00038919 | m | 1.46 | 0.0038645 | 0.0615 | 0.0049995 | m | OK | 1.46 |  | 0.0615 |  | m | -1.9145e-07 | 0.0038645 | -2.5513e-07 | 0.0049995 | m | -7.4351e-05 |  | -0.00010627 |  | m | 1 | kHz | 0.001 | 0.0014599 | 1.9325e-06 | 6.1394e-05 | 8.2483e-07 | 0.0014612 | 1.9311e-06 | 0.042028 | 0.00056672 | 2 | 1m | 0.001 | 52 | 2 | 3m | 70 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 7.748e-08 | 0.0015949 | 1.0078e-07 | 0.00016428 | 1kHz3m3m | -8.3541e-08 | 0.002271 | -8.5512e-08 | 0.0038778 | 4.0347e-08 | 0.002271 | 5.7823e-08 | 0.0010206 | 0.00146 | 0.002271 | 6.1452e-05 | 0.0010206 | 1.0001 | 3.804e-06 | 9.9887e-05 | 0.0038076 | OK | 0.99997 | 2.2768e-06 | -2.8192e-05 | 0.0013002 | OK | 1.0001 | 2.2768e-06 | 7.0351e-05 | 0.0032293 | 0.0014613 | 3.8622e-06 | 0.042098 | 0.0034225 | 0.00146 | 3.8645e-06 | 6.15e-05 | 4.9995e-06
10 | m | 10 | -6.5817 | 0.0011473 | -4.6711 | 0.00078741 | m | -0.0019414 | 0.00033177 | -0.0015154 | 5.2797e-05 | m | -6.58 | 0.0037157 | -4.67 | 0.0047799 | m | OK | -6.58 |  | -4.67 |  | m | -1.2956e-07 | 0.0037157 | 3.7746e-08 | 0.0047799 | m | 0.00020282 |  | 0.00040778 |  | m | 10 | Hz | 0.001 | -0.0065798 | 1.1943e-06 | -0.0046696 | 7.8918e-07 | 0.0080684 | 1.0758e-06 | -2.5244 | 0.00011706 | 3 | 3m | 0.003 | 82 | 3 | 10m | 100 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -7.5776e-09 | 0.0034114 | 3.0004e-07 | 0.0028606 | 10Hz10m10m | 6.9689e-07 | 0.0039654 | -1.2962e-08 | 0.0037714 | 5.025e-07 | 0.0039654 | 7.3409e-08 | 0.0035201 | -0.0065793 | 0.0039654 | -0.0046695 | 0.0035201 | 1 | 4.1881e-06 | 9.999e-05 | 0.0013259 | OK | 1.0001 | 4.322e-06 | -1.2968e-06 | 0.00040777 | OK | 1 | 4.322e-06 | 2.6653e-05 | 0.00066111 | 0.0080688 | 2.1518e-06 | -2.5244 | 0.00070134 | -0.00658 | 3.7157e-06 | -0.00467 | 4.7799e-06
100 | m | 2 | -4.5625 | 0.0017388 | -6.4785 | 0.0016676 | m | -0.0014062 | 0.001886 | 0.0016519 | 0.001138 | m | -4.56 | 0.0054893 | -6.48 | 0.0049828 | m | Extrapolated | -4.56 |  | -6.48 |  | m | 2.1008e-08 | 0.0054893 | -4.1767e-10 | 0.0049828 | m | -0.0010897 |  | -0.00014709 |  | m | 2 | Hz | 0.001 | -0.0045611 | 2.5652e-06 | -0.0064801 | 2.019e-06 | 0.0079244 | 2.2149e-06 | -2.1841 | 0.00030262 | 3 | 3m | 0.003 | 0 | 4 | 100m | 134 | 4 | 10m | 100m | 0.01 | 0.1 | 2Hz100m10m | 2Hz100m100m | 116 | 134 | 2Hz100m10m | -9.3168e-07 | 0.0019191 | 9.9982e-07 | 0.0023846 | 2Hz100m100m | -4.3209e-06 | 0.0046507 | 2.2823e-06 | 0.0014106 | -8.5352e-07 | 0.0046507 | 9.7024e-07 | 0.0024071 | -0.0045619 | 0.0046507 | -0.0064792 | 0.0024071 | 0.99991 | 3.8756e-06 | 0.0001 | 0.00043471 | OK | 0.99996 | 5.2663e-06 | 2.2821e-05 | 6.2969e-05 | OK | 0.99991 | 5.2663e-06 | 0.00010178 | 0.00044328 | 0.0079236 | 4.43e-06 | -2.184 | 0.0007502 | -0.00456 | 5.4893e-06 | -0.00648 | 4.9828e-06
3 | m | 5000 | -0.91973 | 0.0012274 | -0.034217 | 0.00026005 | m | -0.0016906 | 0.0019074 | 0.00048991 | 0.0013095 | m | -0.918 | 0.0045325 | -0.034799 | 0.0030655 | m | Extrapolated | -0.918 |  | -0.0348 |  | m | -1.5375e-07 | 0.0045325 | 5.7925e-07 | 0.0030655 | m | -4.3692e-05 |  | 9.2756e-05 |  | m | 5 | kHz | 0.001 | -0.00091804 | 2.2682e-06 | -3.4707e-05 | 1.3351e-06 | 0.0009187 | 2.2672e-06 | -3.1038 | 0.0014552 | 1 | 0m | 0 | 0 | 2 | 3m | 72 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | -4.8442e-08 | 0.0038257 | 9.9272e-08 | 0.00077836 | 5kHz3m3m | -8.919e-08 | 0.0017082 | 9.2925e-08 | 0.00082202 | -4.6786e-08 | 0.0017082 | 9.953e-08 | 0.00077659 | -0.00091809 | 0.0017082 | -3.4608e-05 | 0.00077659 | 0.99996 | 4.3489e-06 | 9.9306e-05 | 0.0015876 | OK | 0.99997 | 1.7162e-06 | 3.1633e-05 | 0.0006263 | OK | 0.99996 | 1.7162e-06 | 0.00010206 | 0.0016267 | 0.00091866 | 4.5343e-06 | -3.1037 | 0.0033341 | -0.000918 | 4.5325e-06 | -3.4799e-05 | 3.0655e-06
10 | m | 2 | -1.7805 | 0.0019778 | -9.7395 | 0.0016885 | m | -0.0011087 | 0.0015469 | -0.00011801 | 0.0019509 | m | -1.78 | 0.0062649 | -9.74 | 0.0051967 | m | OK | -1.78 |  | -9.74 |  | m | 4.6956e-07 | 0.0062649 | 1.3986e-07 | 0.0051967 | m | 0.00064871 |  | 0.0005903 |  | m | 2 | Hz | 0.001 | -0.0017794 | 2.5109e-06 | -0.0097394 | 2.5801e-06 | 0.0099006 | 2.5779e-06 | -1.7515 | 0.00025384 | 3 | 3m | 0.003 | 80 | 3 | 10m | 98 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | -2.4036e-08 | 0.0035532 | 3e-07 | 0.0027952 | 2Hz10m10m | 7.1523e-07 | 0.00022248 | -5.5883e-07 | 0.0037175 | 7.0473e-07 | 0.00022248 | -5.4664e-07 | 0.0037044 | -0.0017786 | 0.00022253 | -0.00974 | 0.0037044 | 0.99999 | 4.8867e-06 | 9.9999e-05 | 0.00094571 | OK | 1.0001 | 2.9068e-06 | -5.5905e-05 | 0.00037521 | OK | 1.0001 | 2.9068e-06 | -5.3692e-05 | 0.00038331 | 0.0099013 | 5.1559e-06 | -1.7516 | 0.00063613 | -0.00178 | 6.2649e-06 | -0.00974 | 5.1967e-06
10 | m | 100 | -5.8178 | 0.00012095 | 5.5404 | 0.00055004 | m | 0.0014359 | 0.001261 | 0.00016273 | 0.00055372 | m | -5.82 | 0.0037909 | 5.54 | 0.0039231 | m | OK | -5.82 |  | 5.54 |  | m | -1.0195e-07 | 0.0037909 | 1.192e-07 | 0.0039231 | m | 0.00074385 |  | 0.00021182 |  | m | 100 | Hz | 0.001 | -0.0058193 | 1.2668e-06 | 0.0055402 | 7.8048e-07 | 0.0080348 | 1.0637e-06 | 2.3808 | 0.00012949 | 3 | 3m | 0.003 | 85 | 3 | 10m | 103 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | -2.1019e-07 | 0.0028016 | 2.9984e-07 | 0.0031983 | 100Hz10m10m | 9.5218e-07 | 0.0036777 | 7.6347e-07 | 0.0017965 | 6.2585e-07 | 0.0036777 | 6.3331e-07 | 0.0021901 | -0.0058186 | 0.0036777 | 0.0055408 | 0.0021901 | 0.99993 | 3.6728e-06 | 9.9999e-05 | 0.0013413 | OK | 1.0001 | 5.1837e-06 | 7.6258e-05 | 0.00026755 | OK | 1 | 5.1837e-06 | 8.2923e-05 | 0.000569 | 0.0080352 | 2.1278e-06 | 2.3808 | 0.00062517 | -0.00582 | 3.7909e-06 | 0.00554 | 3.9231e-06
1 | m | 100 | -0.11929 | 0.0019787 | 0.93246 | 0.00021192 | m | 0.00068931 | 0.0008332 | -0.0015554 | 0.00058517 | m | -0.12 | 0.0050962 | 0.934 | 0.0014907 | m | OK | -0.12 |  | 0.934 |  | m | -1.5903e-06 | 0.0050962 | -2.0468e-07 | 0.0014907 | m | 2.4413e-05 |  | 1.3045e-05 |  | m | 100 | Hz | 0.001 | -0.00011998 | 2.1469e-06 | 0.00093401 | 6.2236e-07 | 0.00094169 | 6.7518e-07 | 1.6985 | 0.0022629 | 1 | 0m | 0 | 13 | 1 | 1m | 31 | 1 | 0m | 1m | 0 | 0.001 | 100Hz1m0m | 100Hz1m1m | 13 | 31 | 100Hz1m0m | 0 | 0.0013347 | 0 | 0.00032947 | 100Hz1m1m | -1.1112e-08 | 0.00031278 | 2.9172e-08 | 0.0019521 | -1.0464e-08 | 0.00031278 | 2.9172e-08 | 0.0018575 | -0.00011999 | 0.00031281 | 0.00093404 | 0.0018575 | 1 | 1.3747e-06 | 2.9182e-05 | 0.0030423 | OK | 0.99999 | 3.5858e-06 | 2.9182e-05 | 0.0030423 | OK | 0.99999 | 3.5858e-06 | 2.9182e-05 | 0.0030423 | 0.00094168 | 1.3504e-06 | 1.6986 | 0.0054532 | -0.00012 | 5.0962e-06 | 0.000934 | 1.4907e-06
10 | m | 0.05 | 2.5883 | 0.0016275 | 6.3609 | 0.0002453 | m | -0.0018899 | 0.0017122 | 0.0010901 | 0.0012872 | m | 2.59 | 0.0046464 | 6.36 | 0.0033383 | m | OK | 2.59 |  | 6.36 |  | m | 4.2994e-08 | 0.0046464 | -2.988e-08 | 0.0033383 | m | 0.00015005 |  | -0.00023447 |  | m | 50 | mHz | 0.001 | 0.0025902 | 2.3623e-06 | 0.0063598 | 1.3104e-06 | 0.006867 | 1.5056e-06 | 1.184 | 0.00032663 | 3 | 3m | 0.003 | 75 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 1.0683e-07 | 0.002322 | 1.7763e-08 | 0.00082848 | 0.0025903 | 0.002322 | 0.0063598 | 0.00082849 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK | 1 | 3.906e-06 | 3.3107e-05 | 0.00027401 | 0.0068671 | 3.0113e-06 | 1.1841 | 0.0007084 | 0.00259 | 4.6464e-06 | 0.00636 | 3.3383e-06
1 | m | 0.05 | 0.28839 | 0.00034187 | -0.24981 | 0.001954 | m | 0.00038525 | 0.001794 | 0.0011811 | 0.00012038 | m | 0.28799 | 0.0038009 | -0.25102 | 0.0038111 | m | OK | 0.288 |  | -0.251 |  | m | -1.4076e-05 | 0.0038009 | -1.615e-05 | 0.0038111 | m | 5.9299e-06 |  | 1.2402e-05 |  | m | 50 | mHz | 0.001 | 0.00028801 | 1.8263e-06 | -0.00025099 | 1.9577e-06 | 0.00038202 | 1.8841e-06 | -0.71683 | 0.004979 | 1 | 0m | 0 | 3 | 1 | 1m | 21 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.0036718 | 0 | 0.0029355 | 50mHz1m1m | 2.5191e-08 | 0.0031843 | -9.0742e-08 | 0.0010893 | 9.6236e-09 | 0.0031843 | -9.0742e-08 | 0.0022302 | 0.00028802 | 0.0031843 | -0.00025108 | 0.0022302 | 1 | 4.7009e-06 | -9.075e-05 | 0.0014316 | OK | 1 | 3.6855e-06 | -9.075e-05 | 0.0014316 | OK | 1 | 3.6855e-06 | -9.075e-05 | 0.0014316 | 0.00038203 | 3.7683e-06 | -0.71692 | 0.01006 | 0.00028799 | 3.8009e-06 | -0.00025102 | 3.8111e-06
3 | m | 0.02 | 3.022 | 0.0015901 | 0.42507 | 1.0617e-05 | m | 0.0017703 | 0.0018391 | -0.0017688 | 0.00057752 | m | 3.02 | 0.0047758 | 0.427 | 0.0018149 | m | Extrapolated | 3.02 |  | 0.427 |  | m | 1.1142e-08 | 0.0047758 | 1.8798e-08 | 0.0018149 | m | 0.00020738 |  | -0.00016525 |  | m | 20 | mHz | 0.001 | 0.0030202 | 2.4312e-06 | 0.00042683 | 5.7761e-07 | 0.0030502 | 2.4086e-06 | 0.1404 | 0.00021817 | 3 | 3m | 0.003 | 56 | 3 | 10m | 0 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | -3.0087e-08 | 0.0039189 | 1.0001e-07 | 0.0032919 | 20mHz3m3m | -1.7701e-07 | 0.00080857 | 1.9219e-07 | 0.0012775 | -1.807e-07 | 0.00080857 | 1.9451e-07 | 0.0012269 | 0.00302 | 0.00080859 | 0.00042703 | 0.0012269 | 0.99997 | 4.4502e-06 | 9.9998e-05 | 0.004916 | OK | 0.99994 | 1.8569e-06 | 6.4068e-05 | 0.00045958 | OK | 0.99994 | 1.8569e-06 | 6.3166e-05 | 0.00034768 | 0.00305 | 4.8173e-06 | 0.14046 | 0.00055792 | 0.00302 | 4.7758e-06 | 0.000427 | 1.8149e-06
3 | m | 0.1 | -0.7433 | 0.0010924 | -2.9016 | 0.0011799 | m | 0.0015234 | 0.00036051 | -0.0014141 | 0.0014087 | m | -0.745 | 0.002677 | -2.9 | 0.0035516 | m | OK | -0.745 |  | -2.9 |  | m | -5.2581e-09 | 0.002677 | 5.8392e-09 | 0.0035516 | m | 0.00017516 |  | -0.00022867 |  | m | 100 | mHz | 0.001 | -0.00074482 | 1.1503e-06 | -0.0029002 | 1.8376e-06 | 0.0029943 | 1.8027e-06 | -1.8222 | 0.00040218 | 2 | 1m | 0.001 | 40 | 2 | 3m | 58 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 6.0542e-08 | 0.0019825 | 9.9992e-08 | 0.0031027 | 100mHz3m3m | -1.7924e-07 | 0.0040053 | -2.2851e-07 | 2.3224e-05 | -1.7857e-07 | 0.0040053 | -2.2758e-07 | 3.1935e-05 | -0.000745 | 0.0040053 | -0.0029005 | 3.2145e-05 | 1.0001 | 1.9826e-06 | 9.9998e-05 | 0.0034787 | OK | 0.99994 | 4.8158e-06 | -7.6159e-05 | 0.00032186 | OK | 0.99994 | 4.8158e-06 | -7.5661e-05 | 0.00033079 | 0.0029942 | 3.6054e-06 | -1.8223 | 0.00086973 | -0.000745 | 2.677e-06 | -0.0029 | 3.5516e-06
100 | m | 0.2 | -7.2503 | 6.8366e-05 | -6.9889 | 0.0001581 | m | 0.00053643 | 0.0007482 | 0.00047424 | 0.00028902 | m | -7.25 | 0.002569 | -6.99 | 0.0026448 | m | OK | -7.25 |  | -6.99 |  | m | 3.0182e-09 | 0.002569 | -6.2762e-09 | 0.0026448 | m | -0.00084487 |  | 0.00058297 |  | m | 200 | mHz | 0.001 | -0.0072508 | 7.5131e-07 | -0.0069894 | 3.2944e-07 | 0.010071 | 5.8725e-07 | -2.3746 | 5.6879e-05 | 4 | 10m | 0.01 | 113 | 4 | 100m | 131 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | -2.0189e-07 | 0.0027242 | 9.9997e-07 | 0.0011937 | 200mHz100m100m | -6.2211e-06 | 0.004259 | -2.8108e-06 | 0.0031892 | -2.0665e-07 | 0.004259 | 9.9696e-07 | 0.0011953 | -0.0072511 | 0.004259 | -0.0069884 | 0.0011953 | 0.99998 | 3.5304e-06 | 0.0001 | 0.00032803 | OK | 0.99994 | 5.2294e-06 | -2.812e-05 | 6.3741e-05 | OK | 0.99998 | 5.2294e-06 | 9.9899e-05 | 0.00032783 | 0.010071 | 1.1757e-06 | -2.3745 | 0.000347 | -0.00725 | 2.569e-06 | -0.00699 | 2.6448e-06
100 | m | 0.05 | 49.6986 | 0.001866 | 6.4834 | 0.0016126 | m | -0.00028951 | 0.00073491 | -0.0014447 | 0.0010821 | m | 49.7 | 0.0043016 | 6.49 | 0.012452 | m | OK | 49.7 |  | 6.49 |  | m | 3.7646e-08 | 0.0043016 | 6.8966e-09 | 0.012452 | m | -0.0011175 |  | -0.0051615 |  | m | 50 | mHz | 0.001 | 0.049699 | 2.0055e-06 | 0.0064848 | 1.942e-06 | 0.05012 | 2.0045e-06 | 0.12975 | 3.8769e-05 | 4 | 10m | 0.01 | 111 | 4 | 100m | 129 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | 4.7128e-07 | 0.0013032 | 9.9997e-07 | 0.0033159 | 50mHz100m100m | 2.0948e-06 | 0.0032386 | 9.8259e-06 | 0.0035729 | 1.195e-06 | 0.0032386 | 4.9344e-06 | 0.0034304 | 0.0497 | 0.0032386 | 0.0064898 | 0.0034304 | 1 | 3.0828e-06 | 0.0001 | 0.00037978 | OK | 1 | 5.0094e-06 | 9.827e-05 | 6.1747e-05 | OK | 1 | 5.0094e-06 | 9.9229e-05 | 0.00023801 | 0.050122 | 4.0169e-06 | 0.12985 | 0.00025032 | 0.0497 | 4.3016e-06 | 0.00649 | 1.2452e-05
10 | m | 1 | -2.1518 | 0.0010805 | 2.6199 | 0.0016309 | m | -0.0019538 | 0.0011284 | -0.00050461 | 0.0012907 | m | -2.15 | 0.0052115 | 2.62 | 0.0047906 | m | OK | -2.15 |  | 2.62 |  | m | 1.5893e-09 | 0.0052115 | 1.2191e-08 | 0.0047906 | m | 0.00011332 |  | 0.0003658 |  | m | 1 | Hz | 0.001 | -0.0021499 | 1.5623e-06 | 0.0026204 | 2.0798e-06 | 0.0033894 | 1.8887e-06 | 2.2579 | 0.0005277 | 3 | 3m | 0.003 | 79 | 3 | 10m | 97 | 3 | 3m | 10m | 0.003 | 0.01 | 1Hz10m3m | 1Hz10m10m | 79 | 97 | 1Hz10m3m | -2.004e-07 | 0.0020819 | 3.0004e-07 | 0.0036539 | 1Hz10m10m | 1.5495e-07 | 0.0038498 | -2.2976e-08 | 0.0035852 | -1.8063e-07 | 0.0038498 | 2.8207e-07 | 0.00365 | -0.0021501 | 0.0038498 | 0.0026206 | 0.00365 | 0.99993 | 3.0503e-06 | 9.9996e-05 | 0.0014729 | OK | 1 | 4.2493e-06 | -2.2971e-06 | 0.00045686 | OK | 0.99994 | 4.2493e-06 | 9.4305e-05 | 0.0014164 | 0.0033892 | 3.7775e-06 | 2.258 | 0.0017664 | -0.00215 | 5.2115e-06 | 0.00262 | 4.7906e-06
10 | m | 0.02 | 1.4506 | 0.00072101 | -4.7413 | 0.00055669 | m | 0.0008715 | 0.00056669 | -0.0012048 | 0.00012926 | m | 1.45 | 0.0038255 | -4.74 | 0.0016473 | m | OK | 1.45 |  | -4.74 |  | m | -1.977e-08 | 0.0038255 | -1.4103e-09 | 0.0016473 | m | -0.00029304 |  | -5.3719e-05 |  | m | 20 | mHz | 0.001 | 0.0014497 | 9.1705e-07 | -0.0047401 | 5.715e-07 | 0.0049568 | 6.0877e-07 | -1.274 | 0.00018011 | 3 | 3m | 0.003 | 74 | 3 | 10m | 92 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | 7.0524e-08 | 0.0022397 | 2.9998e-07 | 0.002096 | 20mHz10m10m | -3.5801e-07 | 0.0034885 | -4.4193e-07 | 0.0018726 | -4.927e-08 | 0.0034885 | 9.2587e-08 | 0.0020335 | 0.0014497 | 0.0034885 | -0.00474 | 0.0020335 | 1 | 4.262e-06 | 0.00010001 | 0.00092181 | OK | 0.99996 | 3.7418e-06 | -4.418e-05 | 0.00019379 | OK | 1 | 3.7418e-06 | 5.97e-05 | 0.0007183 | 0.0049568 | 1.2177e-06 | -1.2739 | 0.00080356 | 0.00145 | 3.8255e-06 | -0.00474 | 1.6473e-06
1 | m | 1 | 0.0019677 | 0.0010937 | 0.027765 | 0.00060545 | m | -0.00048227 | 1.429e-05 | -0.00073462 | 0.00051514 | m | 0.0024526 | 0.0021825 | 0.0285 | 0.0016002 | m | OK | 0.00245 |  | 0.0285 |  | m | 2.6455e-06 | 0.0021825 | -2.2755e-07 | 0.0016002 | m | -7.7596e-08 |  | 1.0115e-08 |  | m | 1 | Hz | 0.001 | 2.4499e-06 | 1.0938e-06 | 2.85e-05 | 7.9495e-07 | 2.8605e-05 | 7.9755e-07 | 1.485 | 0.03817 | 1 | 0m | 0 | 7 | 1 | 1m | 25 | 1 | 0m | 1m | 0 | 0.001 | 1Hz1m0m | 1Hz1m1m | 7 | 25 | 1Hz1m0m | 0 | 0.0038489 | 0 | 0.00029027 | 1Hz1m1m | -4.1944e-09 | 0.0026643 | -9.5577e-08 | 0.0024785 | -1.1998e-10 | 0.0026643 | -9.5577e-08 | 0.00035286 | 2.4498e-06 | 0.0026643 | 2.8404e-05 | 0.00035287 | 1 | 3.8598e-06 | -9.556e-05 | 0.0036582 | OK | 1 | 4.496e-06 | -9.556e-05 | 0.0036582 | OK | 1 | 4.496e-06 | -9.556e-05 | 0.0036582 | 2.8605e-05 | 1.5951e-06 | 1.4849 | 0.076428 | 2.4526e-06 | 2.1825e-06 | 2.85e-05 | 1.6002e-06
100 | m | 2000 | 4.4766 | 3.8651e-05 | 79.9939 | 0.0019869 | m | 0.001887 | 0.0009222 | -0.0013872 | 0.00036292 | m | 4.47 | 0.020443 | 80 | 0.0042009 | m | OK | 4.47 |  | 80 |  | m | 1.953e-05 | 0.020443 | 2.5608e-05 | 0.0042009 | m | 0.0046798 |  | -0.0047363 |  | m | 2 | kHz | 0.001 | 0.0044747 | 9.2301e-07 | 0.079995 | 2.0197e-06 | 0.08012 | 2.0172e-06 | 1.5149 | 1.1588e-05 | 4 | 10m | 0.01 | 125 | 4 | 100m | 143 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | -2.5323e-08 | 0.00057942 | 1.0047e-06 | 0.00010185 | 2kHz100m100m | 7.1609e-06 | 0.0019925 | 5.2033e-06 | 0.0027879 | 5.5735e-06 | 0.0019925 | 4.2759e-06 | 0.0021946 | 0.0044803 | 0.0019925 | 0.08 | 0.0021946 | 1 | 1.8411e-06 | 9.9971e-05 | 0.00043125 | OK | 1.0001 | 4.2684e-06 | 5.0453e-05 | 0.00020434 | OK | 1.0001 | 4.2684e-06 | 6.1391e-05 | 0.00025446 | 0.080125 | 4.0489e-06 | 1.515 | 0.00025552 | 0.00447 | 2.0443e-05 | 0.08 | 4.2009e-06
100 | m | 10 | -15.0956 | 0.00085107 | 96.1056 | 0.0017851 | m | 0.00072313 | 0.0016943 | 0.0012708 | 0.0015106 | m | -15.1 | 0.0073608 | 96.1 | 0.0047503 | m | OK | -15.1 |  | 96.1 |  | m | -1.2385e-06 | 0.0073608 | -1.0102e-06 | 0.0047503 | m | 0.003652 |  | 0.0043198 |  | m | 10 | Hz | 0.001 | -0.015096 | 1.896e-06 | 0.096104 | 2.3385e-06 | 0.097283 | 2.3288e-06 | 1.7266 | 1.9612e-05 | 4 | 10m | 0.01 | 118 | 4 | 100m | 136 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | -7.0881e-07 | 0.0023572 | 1.0002e-06 | 0.0033346 | 10Hz100m100m | -3.7017e-06 | 0.0024586 | 4.2241e-06 | 0.0026354 | -3.6114e-06 | 0.0024586 | 4.1268e-06 | 0.0026565 | -0.0151 | 0.0024586 | 0.096108 | 0.0026565 | 0.99993 | 2.7297e-06 | 9.9999e-05 | 0.0003481 | OK | 0.99996 | 2.7476e-06 | 4.2249e-05 | 5.6555e-05 | OK | 0.99996 | 2.7476e-06 | 4.3992e-05 | 6.5357e-05 | 0.097279 | 4.6653e-06 | 1.7267 | 7.6224e-05 | -0.0151 | 7.3608e-06 | 0.0961 | 4.7503e-06
3 | m | 1 | 0.54585 | 0.0010557 | -2.4415 | 0.00083069 | m | -0.0010114 | 0.00027299 | -0.001486 | 0.00057655 | m | 0.547 | 0.0044881 | -2.44 | 0.0022197 | m | OK | 0.547 |  | -2.44 |  | m | -1.5134e-08 | 0.0044881 | -8.2483e-09 | 0.0022197 | m | -0.00013503 |  | -5.0511e-05 |  | m | 1 | Hz | 0.001 | 0.00054686 | 1.0904e-06 | -0.0024401 | 1.0112e-06 | 0.0025006 | 1.0151e-06 | -1.3503 | 0.00043461 | 2 | 1m | 0.001 | 43 | 2 | 3m | 61 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | -1.9849e-08 | 0.00034442 | 1.0002e-07 | 0.0026891 | 1Hz3m3m | -1.1752e-08 | 0.00047154 | 1.285e-07 | 0.000646 | -1.3774e-08 | 0.00047154 | 1.2138e-07 | 0.0011562 | 0.00054685 | 0.00047154 | -0.0024399 | 0.0011562 | 0.99998 | 3.6518e-07 | 9.9995e-05 | 0.0046729 | OK | 1 | 7.7074e-07 | 4.2828e-05 | 0.0005917 | OK | 0.99999 | 7.7074e-07 | 5.7103e-05 | 0.0016108 | 0.0025006 | 2.0302e-06 | -1.3503 | 0.0018304 | 0.000547 | 4.4881e-06 | -0.00244 | 2.2197e-06
10 | m | 200 | 7.9581 | 0.0011898 | 4.4994 | 0.0012253 | m | -0.0017209 | 0.00038707 | 0.00030454 | 0.0018349 | m | 7.96 | 0.0036451 | 4.5 | 0.0046292 | m | OK | 7.96 |  | 4.5 |  | m | 1.1289e-07 | 0.0036451 | 4.0949e-08 | 0.0046292 | m | -0.00018746 |  | -0.00091077 |  | m | 200 | Hz | 0.001 | 0.0079598 | 1.2512e-06 | 0.0044991 | 2.2064e-06 | 0.0091433 | 1.5379e-06 | 0.51446 | 0.0002206 | 3 | 3m | 0.003 | 86 | 3 | 10m | 104 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -1.5438e-07 | 0.0011539 | 2.9983e-07 | 0.00034655 | 200Hz10m10m | 8.3222e-07 | 0.0040858 | 7.3506e-07 | 0.0027163 | 7.1148e-07 | 0.0040858 | 6.818e-07 | 0.0024263 | 0.0079605 | 0.0040858 | 0.0044998 | 0.0024263 | 0.99995 | 2.2817e-06 | 9.9996e-05 | 0.00064432 | OK | 1.0001 | 4.9523e-06 | 7.3361e-05 | 0.00028354 | OK | 1.0001 | 4.9523e-06 | 7.6621e-05 | 0.00032769 | 0.0091439 | 3.0761e-06 | 0.51453 | 0.00054959 | 0.00796 | 3.6451e-06 | 0.0045 | 4.6292e-06
1 | m | 0.1 | -0.39948 | 0.0011921 | 0.50312 | 0.0017135 | m | -0.0014746 | 0.0010992 | 7.6429e-05 | 0.00051183 | m | -0.39801 | 0.0040354 | 0.50299 | 0.0038277 | m | OK | -0.398 |  | 0.503 |  | m | -9.894e-06 | 0.0040354 | -7.8282e-06 | 0.0038277 | m | -5.0001e-06 |  | 4.2695e-05 |  | m | 100 | mHz | 0.001 | -0.00039801 | 1.6215e-06 | 0.00050304 | 1.7883e-06 | 0.00064145 | 1.726e-06 | 2.2401 | 0.0026311 | 1 | 0m | 0 | 4 | 1 | 1m | 22 | 1 | 0m | 1m | 0 | 0.001 | 100mHz1m0m | 100mHz1m1m | 4 | 22 | 100mHz1m0m | 0 | 0.0012943 | 0 | 0.0026332 | 100mHz1m1m | -8.8918e-08 | 0.0027176 | 5.4866e-08 | 0.0027374 | -5.7037e-08 | 0.0027176 | 5.4866e-08 | 0.0027 | -0.00039806 | 0.0027176 | 0.0005031 | 0.0027 | 1 | 2.9341e-06 | 5.4857e-05 | 0.0043058 | OK | 0.99991 | 2.7297e-06 | 5.4857e-05 | 0.0043058 | OK | 0.99994 | 2.7297e-06 | 5.4857e-05 | 0.0043058 | 0.00064141 | 3.452e-06 | 2.2402 | 0.0067992 | -0.00039801 | 4.0354e-06 | 0.00050299 | 3.8277e-06
10 | m | 0.5 | 2.8786 | 0.00033606 | 3.1807 | 0.00022488 | m | -0.0017664 | 0.0012331 | 0.00081277 | 0.0019115 | m | 2.88 | 0.0036912 | 3.18 | 0.0036288 | m | OK | 2.88 |  | 3.18 |  | m | 1.628e-08 | 0.0036912 | -8.0006e-09 | 0.0036288 | m | 0.0003246 |  | -0.00010729 |  | m | 500 | mHz | 0.001 | 0.0028803 | 1.2781e-06 | 0.0031799 | 1.9246e-06 | 0.0042905 | 1.6646e-06 | 0.83479 | 0.00037341 | 3 | 3m | 0.003 | 78 | 3 | 10m | 96 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | -1.1635e-07 | 0.00232 | 3.0001e-07 | 0.0016707 | 500mHz10m10m | -3.3994e-08 | 0.00076531 | -4.7206e-07 | 0.0033821 | -1.0117e-07 | 0.00076531 | 1.5768e-07 | 0.0019862 | 0.0028802 | 0.00076532 | 0.0031801 | 0.0019862 | 0.99996 | 3.3039e-06 | 9.9997e-05 | 0.00055959 | OK | 1 | 3.0413e-06 | -4.7205e-05 | 0.00047381 | OK | 0.99997 | 3.0413e-06 | 7.286e-05 | 0.00054378 | 0.0042903 | 3.3293e-06 | 0.83486 | 0.00092382 | 0.00288 | 3.6912e-06 | 0.00318 | 3.6288e-06
1 | m | 20 | -0.93157 | 0.00042688 | 0.39804 | 0.0010686 | m | 0.0013838 | 0.0011917 | 0.0011426 | 0.0013722 | m | -0.933 | 0.0030696 | 0.397 | 0.0043709 | m | Extrapolated | -0.933 |  | 0.397 |  | m | -4.2702e-07 | 0.0030696 | -9.972e-07 | 0.0043709 | m | 4.3766e-05 |  | -0.00010499 |  | m | 20 | Hz | 0.001 | -0.00093296 | 1.2659e-06 | 0.0003969 | 1.7392e-06 | 0.0010139 | 1.3492e-06 | 2.7394 | 0.0016524 | 2 | 1m | 0.001 | 29 | 2 | 3m | 0 | 1 | 0m | 1m | 0 | 0.001 | 20Hz1m0m | 20Hz1m1m | 11 | 29 | 20Hz1m0m | 0 | 0.00046609 | 0 | 0.0010025 | 20Hz1m1m | 7.918e-08 | 0.0037401 | -7.7293e-08 | 0.0011892 | 8.0278e-08 | 0.0037401 | -7.7293e-08 | 0.0011918 | -0.00093288 | 0.0037401 | 0.00039682 | 0.0011918 | 1 | 1.1055e-06 | -7.7312e-05 | 0.0031213 | OK | 1.0001 | 5.0353e-06 | -7.7312e-05 | 0.0031213 | OK | 1.0001 | 5.0353e-06 | -7.7312e-05 | 0.0031213 | 0.001014 | 2.6985e-06 | 2.7393 | 0.0045458 | -0.000933 | 3.0696e-06 | 0.000397 | 4.3709e-06
10 | m | 0.5 | -7.3883 | 0.0013533 | -5.8695 | 0.0017012 | m | 0.001505 | 0.00060072 | 0.00077667 | 0.00046106 | m | -7.39 | 0.0042973 | -5.87 | 0.0048282 | m | OK | -7.39 |  | -5.87 |  | m | -5.6297e-08 | 0.0042973 | 4.2947e-08 | 0.0048282 | m | 0.00016152 |  | -0.00029804 |  | m | 500 | mHz | 0.001 | -0.0073898 | 1.4807e-06 | -0.0058703 | 1.7626e-06 | 0.0094377 | 1.5957e-06 | -2.4703 | 0.00017581 | 3 | 3m | 0.003 | 78 | 3 | 10m | 96 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | -1.1635e-07 | 0.00232 | 3.0001e-07 | 0.0016707 | 500mHz10m10m | -3.3994e-08 | 0.00076531 | -4.7206e-07 | 0.0033821 | -4.061e-08 | 0.00076531 | -4.1004e-07 | 0.0032446 | -0.0073899 | 0.00076532 | -0.0058707 | 0.0032446 | 0.99996 | 3.3039e-06 | 9.9997e-05 | 0.00055959 | OK | 1 | 3.0413e-06 | -4.7205e-05 | 0.00047381 | OK | 0.99999 | 3.0413e-06 | -3.538e-05 | 0.0004807 | 0.0094376 | 3.1914e-06 | -2.4703 | 0.00059557 | -0.00739 | 4.2973e-06 | -0.00587 | 4.8282e-06
100 | m | 0.2 | 85.1074 | 0.0011587 | -56.8007 | 0.0011057 | m | 0.00021279 | 0.0010256 | 0.00017764 | 0.00046692 | m | 85.1 | 0.0043132 | -56.8 | 0.0055628 | m | Extrapolated | 85.1 |  | -56.8 |  | m | 2.1078e-06 | 0.0043132 | 1.6596e-06 | 0.0055628 | m | 0.007174 |  | -0.00091952 |  | m | 200 | mHz | 0.001 | 0.085107 | 1.5474e-06 | -0.056801 | 1.2002e-06 | 0.10232 | 1.4493e-06 | -0.58851 | 1.2871e-05 | 5 | 100m | 0.1 | 131 | 5 | 0 | 0 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | -2.0189e-07 | 0.0027242 | 9.9997e-07 | 0.0011937 | 200mHz100m100m | -6.2211e-06 | 0.004259 | -2.8108e-06 | 0.0031892 | -6.3763e-06 | 0.004259 | -2.9091e-06 | 0.0032406 | 0.085101 | 0.004259 | -0.056804 | 0.0032406 | 0.99998 | 3.5304e-06 | 0.0001 | 0.00032803 | OK | 0.99994 | 5.2294e-06 | -2.812e-05 | 6.3741e-05 | OK | 0.99994 | 5.2294e-06 | -3.1424e-05 | 5.6925e-05 | 0.10231 | 2.9476e-06 | -0.58854 | 6.2476e-05 | 0.0851 | 4.3132e-06 | -0.0568 | 5.5628e-06
10 | m | 50 | 3.9603 | 0.0019627 | -9.4785 | 0.0009216 | m | 0.00035225 | 0.0016928 | 0.0012451 | 0.0018166 | m | 3.96 | 0.006921 | -9.48 | 0.0048292 | m | Extrapolated | 3.96 |  | -9.48 |  | m | 5.8293e-08 | 0.006921 | 1.1269e-07 | 0.0048292 | m | -9.8886e-05 |  | 0.00028147 |  | m | 50 | Hz | 0.001 | 0.0039599 | 2.5918e-06 | -0.0094797 | 2.037e-06 | 0.010274 | 2.1286e-06 | -1.1751 | 0.00024501 | 4 | 10m | 0.01 | 102 | 4 | 100m | 0 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | -1.6068e-07 | 0.00012487 | 2.9993e-07 | 0.0016747 | 50Hz10m10m | 2.5877e-07 | 0.0037452 | 2.1659e-08 | 0.0033887 | 2.7516e-07 | 0.0037452 | 1.0784e-08 | 0.0034557 | 0.0039602 | 0.0037452 | -0.0094797 | 0.0034557 | 0.99995 | 2.1729e-06 | 9.9999e-05 | 0.0005617 | OK | 1 | 4.7426e-06 | 2.1544e-06 | 0.00051459 | OK | 1 | 4.7426e-06 | -1.6693e-06 | 0.00051274 | 0.010274 | 4.2575e-06 | -1.1751 | 0.00070925 | 0.00396 | 6.921e-06 | -0.00948 | 4.8292e-06
3 | m | 0.02 | 2.0084 | 0.00050664 | -0.4339 | 0.0012361 | m | -0.0017013 | 0.001679 | -0.0007169 | 0.001477 | m | 2.01 | 0.003708 | -0.433 | 0.006415 | m | OK | 2.01 |  | -0.433 |  | m | 4.0021e-09 | 0.003708 | 2.8116e-09 | 0.006415 | m | 5.6084e-05 |  | -0.0001825 |  | m | 20 | mHz | 0.001 | 0.0020101 | 1.7538e-06 | -0.00043318 | 1.926e-06 | 0.0020562 | 1.7618e-06 | -0.21226 | 0.00093312 | 2 | 1m | 0.001 | 38 | 2 | 3m | 56 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | -3.0087e-08 | 0.0039189 | 1.0001e-07 | 0.0032919 | 20mHz3m3m | -1.7701e-07 | 0.00080857 | 1.9219e-07 | 0.0012775 | -1.0768e-07 | 0.00080857 | 1.4869e-07 | 0.0022281 | 0.0020099 | 0.00080858 | -0.00043303 | 0.0022281 | 0.99997 | 4.4502e-06 | 9.9998e-05 | 0.004916 | OK | 0.99994 | 1.8569e-06 | 6.4068e-05 | 0.00045958 | OK | 0.99995 | 1.8569e-06 | 8.1023e-05 | 0.0025626 | 0.0020561 | 3.5236e-06 | -0.21218 | 0.0031701 | 0.00201 | 3.708e-06 | -0.000433 | 6.415e-06
1 | m | 5000 | -0.10387 | 0.00016525 | 0.1794 | 0.00074144 | m | 0.0011282 | 0.00034015 | 0.0014007 | 0.00051606 | m | -0.10501 | 0.0013607 | 0.17799 | 0.001523 | m | OK | -0.105 |  | 0.178 |  | m | -1.435e-05 | 0.0013607 | -8.5057e-06 | 0.001523 | m | 2.8095e-06 |  | 3.5701e-06 |  | m | 5 | kHz | 0.001 | -0.000105 | 3.7816e-07 | 0.000178 | 9.0335e-07 | 0.00020666 | 8.0145e-07 | 2.1037 | 0.0027232 | 1 | 0m | 0 | 18 | 1 | 1m | 36 | 1 | 0m | 1m | 0 | 0.001 | 5kHz1m0m | 5kHz1m1m | 18 | 36 | 5kHz1m0m | 0 | 0.0014999 | 0 | 0.0035373 | 5kHz1m1m | -4.4901e-08 | 0.001159 | 1.0137e-07 | 0.0013002 | -9.2794e-09 | 0.001159 | 1.0137e-07 | 0.0030749 | -0.00010501 | 0.001159 | 0.0001781 | 0.0030749 | 1 | 3.8421e-06 | 0.0001012 | 0.0027993 | OK | 0.99996 | 1.2513e-06 | 0.0001012 | 0.0027993 | OK | 0.99999 | 1.2513e-06 | 0.0001012 | 0.0027993 | 0.00020666 | 1.6029e-06 | 2.1038 | 0.0061237 | -0.00010501 | 1.3607e-06 | 0.00017799 | 1.523e-06
10 | m | 2000 | 8.4818 | 0.00027811 | 0.88421 | 0.0015154 | m | 0.0018211 | 8.9536e-05 | 9.2749e-06 | 0.00010283 | m | 8.48 | 0.00087911 | 0.88399 | 0.0055886 | m | OK | 8.48 |  | 0.884 |  | m | 7.7301e-06 | 0.00087911 | -7.8002e-06 | 0.0055886 | m | -5.1526e-05 |  | 0.0002037 |  | m | 2 | kHz | 0.001 | 0.0084799 | 2.9216e-07 | 0.0008842 | 1.5189e-06 | 0.0085259 | 3.3054e-07 | 0.10389 | 0.00017723 | 3 | 3m | 0.003 | 89 | 3 | 10m | 107 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -2.9449e-07 | 0.0016532 | 2.9509e-07 | 0.00067073 | 2kHz10m10m | 3.2318e-07 | 0.0036551 | -5.8467e-07 | 0.0039341 | 1.9311e-07 | 0.0036551 | -3.9941e-07 | 0.0032469 | 0.0084801 | 0.0036551 | 0.0008838 | 0.0032469 | 0.9999 | 4.0898e-06 | 9.9658e-05 | 0.00042937 | OK | 1 | 4.0238e-06 | -5.8752e-05 | 0.00058918 | OK | 1 | 4.0238e-06 | -2.5394e-05 | 0.00055553 | 0.008526 | 6.6196e-07 | 0.10387 | 0.00065898 | 0.00848 | 8.7911e-07 | 0.00088399 | 5.5886e-06
10 | m | 0.05 | 2.768 | 0.001837 | 0.92891 | 0.00015522 | m | -0.0019815 | 0.0010569 | -0.00076954 | 0.00018466 | m | 2.77 | 0.0038588 | 0.93 | 0.002195 | m | Extrapolated | 2.77 |  | 0.93 |  | m | 2.1308e-10 | 0.0038588 | -3.334e-09 | 0.002195 | m | -1.8962e-05 |  | -0.00031874 |  | m | 50 | mHz | 0.001 | 0.00277 | 2.1193e-06 | 0.00092968 | 2.4123e-07 | 0.0029218 | 2.0106e-06 | 0.32381 | 0.0002437 | 2 | 1m | 0.001 | 0 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 1.2187e-07 | 0.002322 | 3.057e-07 | 0.0012422 | 0.0027701 | 0.002322 | 0.00092999 | 0.0012422 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK | 1 | 3.906e-06 | 0.00010135 | 0.00042058 | 0.002922 | 4.0212e-06 | 0.32392 | 0.00064378 | 0.00277 | 3.8588e-06 | 0.00093 | 2.195e-06
1 | m | 10 | -0.27241 | 2.7277e-05 | 0.057467 | 0.0017305 | m | 0.00059385 | 0.0010326 | -2.9314e-05 | 0.00032881 | m | -0.273 | 0.0022252 | 0.057513 | 0.0034942 | m | OK | -0.273 |  | 0.0575 |  | m | 2.6619e-06 | 0.0022252 | 1.2641e-05 | 0.0034942 | m | -5.2143e-06 |  | -4.011e-06 |  | m | 10 | Hz | 0.001 | -0.00027301 | 1.033e-06 | 5.7496e-05 | 1.7615e-06 | 0.00027899 | 1.074e-06 | 2.934 | 0.006225 | 1 | 0m | 0 | 10 | 1 | 1m | 28 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0001832 | 0 | 0.0021009 | 10Hz1m1m | -5.4917e-08 | 0.00018788 | -6.4216e-08 | 0.0024586 | -1.5321e-08 | 0.00018788 | -6.4216e-08 | 0.0022007 | -0.00027302 | 0.00018789 | 5.7432e-05 | 0.0022007 | 1 | 2.1088e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99995 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99998 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | 0.00027899 | 2.1481e-06 | 2.934 | 0.012696 | -0.000273 | 2.2252e-06 | 5.7513e-05 | 3.4942e-06
3 | m | 0.2 | 1.1688 | 0.00073212 | 0.50657 | 0.00066835 | m | -0.0013166 | 0.00098263 | -0.00037098 | 0.00025726 | m | 1.17 | 0.0023148 | 0.507 | 0.0022931 | m | OK | 1.17 |  | 0.507 |  | m | -4.1413e-09 | 0.0023148 | 7.0939e-09 | 0.0022931 | m | 0.0001153 |  | -6.1646e-05 |  | m | 200 | mHz | 0.001 | 0.0011701 | 1.2254e-06 | 0.00050694 | 7.1615e-07 | 0.0012752 | 1.1599e-06 | 0.40883 | 0.00064146 | 2 | 1m | 0.001 | 41 | 2 | 3m | 59 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | -6.6222e-08 | 0.003636 | 1.0001e-07 | 0.00082951 | 200mHz3m3m | -1.4452e-07 | 0.0040143 | -1.2924e-07 | 0.0036762 | -7.6996e-08 | 0.0040143 | 6.8462e-08 | 0.0012212 | 0.00117 | 0.0040143 | 0.00050701 | 0.0012212 | 0.99993 | 4.4152e-06 | 0.0001 | 0.0011765 | OK | 0.99995 | 5.5669e-06 | -4.3083e-05 | 0.0017443 | OK | 0.99994 | 5.5669e-06 | 8.0312e-05 | 0.0012546 | 0.0012751 | 2.3198e-06 | 0.40891 | 0.0017944 | 0.00117 | 2.3148e-06 | 0.000507 | 2.2931e-06
3 | m | 2000 | -1.1308 | 0.0017268 | -0.49886 | 0.000293 | m | -0.00070447 | 0.0014343 | -0.0019433 | 0.0012235 | m | -1.13 | 0.0041455 | -0.497 | 0.0037603 | m | OK | -1.13 |  | -0.497 |  | m | -7.0871e-08 | 0.0041455 | 7.1858e-07 | 0.0037603 | m | -6.4396e-05 |  | 7.8574e-05 |  | m | 2 | kHz | 0.001 | -0.0011301 | 2.2448e-06 | -0.00049692 | 1.2581e-06 | 0.0012345 | 2.1164e-06 | -2.7273 | 0.0011858 | 2 | 1m | 0.001 | 53 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | -1.7402e-08 | 0.0039544 | 9.9754e-08 | 0.001033 | 2kHz3m3m | -2.1785e-07 | 0.0030069 | -2.3991e-07 | 0.0032239 | -4.0904e-08 | 0.0030069 | 5.9929e-08 | 0.0012899 | -0.0011301 | 0.0030069 | -0.00049686 | 0.0012899 | 0.99998 | 4.1878e-06 | 9.9549e-05 | 0.0018511 | OK | 0.99993 | 4.734e-06 | -7.8203e-05 | 0.0012605 | OK | 0.99998 | 4.734e-06 | 7.8708e-05 | 0.0017819 | 0.0012345 | 4.2327e-06 | -2.7272 | 0.0029664 | -0.00113 | 4.1455e-06 | -0.000497 | 3.7603e-06
3 | m | 0.01 | -0.75881 | 0.00060492 | -1.3984 | 0.0019739 | m | 0.0013531 | 2.0301e-05 | 0.0016425 | 0.0014032 | m | -0.76 | 0.0051279 | -1.4 | 0.0045568 | m | OK | -0.76 |  | -1.4 |  | m | -8.5669e-10 | 0.0051279 | 2.9402e-11 | 0.0045568 | m | -0.00016047 |  | -3.9714e-06 |  | m | 10 | mHz | 0.001 | -0.00076016 | 6.0526e-07 | -0.0014 | 2.4218e-06 | 0.0015931 | 2.1478e-06 | -2.0682 | 0.00079856 | 2 | 1m | 0.001 | 37 | 2 | 3m | 55 | 2 | 1m | 3m | 0.001 | 0.003 | 10mHz3m1m | 10mHz3m3m | 37 | 55 | 10mHz3m1m | -5.5112e-08 | 0.0034877 | 1.0001e-07 | 0.00071514 | 10mHz3m3m | -1.1623e-07 | 0.0016304 | 1.7192e-07 | 0.0032608 | -7.3236e-08 | 0.0016304 | 1.2133e-07 | 0.00147 | -0.00076023 | 0.0016304 | -0.0013999 | 0.00147 | 0.99994 | 4.2004e-06 | 9.9998e-05 | 0.0035946 | OK | 0.99996 | 3.0792e-06 | 5.7307e-05 | 0.0014322 | OK | 0.99995 | 3.0792e-06 | 8.7339e-05 | 0.0029534 | 0.001593 | 4.2957e-06 | -2.0681 | 0.0033576 | -0.00076 | 5.1279e-06 | -0.0014 | 4.5568e-06
100 | m | 50 | -61.6045 | 0.0014319 | 76.3939 | 0.0012851 | m | 0.0002159 | 0.00094823 | 0.00096891 | 0.00081871 | m | -61.6 | 0.0061145 | 76.4 | 0.0052868 | m | OK | -61.6 |  | 76.4 |  | m | 5.2291e-08 | 0.0061145 | 1.627e-07 | 0.0052868 | m | -0.0047019 |  | -0.0071129 |  | m | 50 | Hz | 0.001 | -0.061605 | 1.7174e-06 | 0.076393 | 1.5237e-06 | 0.098138 | 1.6028e-06 | 2.2494 | 1.675e-05 | 4 | 10m | 0.01 | 120 | 4 | 100m | 138 | 4 | 10m | 100m | 0.01 | 0.1 | 50Hz100m10m | 50Hz100m100m | 120 | 138 | 50Hz100m10m | -6.9444e-07 | 0.001543 | 9.9971e-07 | 0.0036907 | 50Hz100m100m | 2.842e-06 | 0.0025742 | -8.6638e-06 | 0.0018828 | 2.7689e-06 | 0.0025742 | -8.4638e-06 | 0.0019203 | -0.061602 | 0.0025742 | 0.076384 | 0.0019203 | 0.99993 | 3.818e-06 | 0.0001 | 0.00040672 | OK | 1 | 3.8454e-06 | -8.6654e-05 | 6.0509e-05 | OK | 1 | 3.8454e-06 | -8.2792e-05 | 6.7673e-05 | 0.09814 | 3.2278e-06 | 2.2494 | 7.5511e-05 | -0.0616 | 6.1145e-06 | 0.0764 | 5.2868e-06
3 | m | 1000 | -0.091195 | 0.00021806 | -1.3198 | 0.00037677 | m | -0.0019957 | 0.0011023 | 0.00014565 | 0.0012484 | m | -0.0891 | 0.005024 | -1.32 | 0.0026225 | m | OK | -0.0891 |  | -1.32 |  | m | -1.8985e-07 | 0.005024 | 1.7841e-07 | 0.0026225 | m | -9.9292e-05 |  | 8.7901e-05 |  | m | 1 | kHz | 0.001 | -8.9199e-05 | 1.1236e-06 | -0.0013199 | 1.304e-06 | 0.0013229 | 1.3032e-06 | -1.6383 | 0.00085002 | 2 | 1m | 0.001 | 52 | 2 | 3m | 70 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 7.748e-08 | 0.0015949 | 1.0078e-07 | 0.00016428 | 1kHz3m3m | -8.3541e-08 | 0.002271 | -8.5512e-08 | 0.0038778 | 5.1481e-08 | 0.002271 | 7.0704e-08 | 0.00076387 | -8.9148e-05 | 0.002271 | -0.0013198 | 0.00076387 | 1.0001 | 3.804e-06 | 9.9887e-05 | 0.0038076 | OK | 0.99997 | 2.2768e-06 | -2.8192e-05 | 0.0013002 | OK | 1.0001 | 2.2768e-06 | 7.9207e-05 | 0.0034027 | 0.001323 | 2.6064e-06 | -1.6382 | 0.0038038 | -8.91e-05 | 5.024e-06 | -0.00132 | 2.6225e-06
10 | m | 0.05 | 3.1919 | 0.0019975 | 2.8885 | 0.00069777 | m | 0.001774 | 0.00043721 | -0.0011326 | 0.00079588 | m | 3.19 | 0.0034354 | 2.89 | 0.0034542 | m | OK | 3.19 |  | 2.89 |  | m | 7.3561e-09 | 0.0034354 | -1.2646e-08 | 0.0034542 | m | 0.00011297 |  | -0.00034743 |  | m | 50 | mHz | 0.001 | 0.0031901 | 2.0448e-06 | 0.0028897 | 1.0584e-06 | 0.0043043 | 1.6738e-06 | 0.73602 | 0.00036733 | 3 | 3m | 0.003 | 75 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 1.166e-07 | 0.002322 | 2.048e-07 | 0.0010972 | 0.0031902 | 0.002322 | 0.0028899 | 0.0010972 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK | 1 | 3.906e-06 | 7.7437e-05 | 0.00036922 | 0.0043044 | 3.3477e-06 | 0.7361 | 0.00082223 | 0.00319 | 3.4354e-06 | 0.00289 | 3.4542e-06
10 | m | 500 | 6.3301 | 0.00081123 | -5.5087 | 0.0014703 | m | 0.00084878 | 0.0011751 | 0.0012721 | 0.0012954 | m | 6.33 | 0.004272 | -5.51 | 0.004526 | m | OK | 6.33 |  | -5.51 |  | m | -2.2031e-07 | 0.004272 | -1.9344e-07 | 0.004526 | m | -0.00078087 |  | -2.1069e-05 |  | m | 500 | Hz | 0.001 | 0.0063292 | 1.4279e-06 | -0.00551 | 1.9596e-06 | 0.0083916 | 1.6779e-06 | -0.71631 | 0.00020857 | 3 | 3m | 0.003 | 87 | 3 | 10m | 105 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 2.1936e-07 | 0.0028245 | 3.0112e-07 | 0.0034461 | 500Hz10m10m | 6.6842e-07 | 0.00013383 | 5.2239e-07 | 0.00095568 | 5.6524e-07 | 0.00013383 | 4.7155e-07 | 0.0015279 | 0.0063298 | 0.00013387 | -0.0055095 | 0.0015279 | 1.0001 | 4.2691e-06 | 9.9988e-05 | 0.0016087 | OK | 1.0001 | 2.8658e-06 | 5.1907e-05 | 0.00012361 | OK | 1.0001 | 2.8658e-06 | 6.2955e-05 | 0.00046482 | 0.0083922 | 3.3559e-06 | -0.71625 | 0.00062456 | 0.00633 | 4.272e-06 | -0.00551 | 4.526e-06
100 | m | 500 | 57.9002 | 0.0021143 | 32.2971 | 0.0020855 | m | 0.00051709 | 0.0018814 | -9.6172e-05 | 0.0012206 | m | 57.9 | 0.0081133 | 32.3 | 0.012046 | m | OK | 57.9 |  | 32.3 |  | m | 3.8553e-06 | 0.0081133 | -2.1286e-07 | 0.012046 | m | -0.0003596 |  | -0.0028107 |  | m | 500 | Hz | 0.001 | 0.0579 | 2.8302e-06 | 0.032297 | 2.4164e-06 | 0.066298 | 2.7377e-06 | 0.50882 | 3.8022e-05 | 4 | 10m | 0.01 | 123 | 4 | 100m | 141 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | -6.7937e-07 | 0.0040166 | 9.9753e-07 | 0.0039442 | 500Hz100m100m | 8.1089e-06 | 0.0018626 | -4.5758e-07 | 0.0028002 | 4.818e-06 | 0.0018626 | 8.7301e-08 | 0.0032286 | 0.057904 | 0.0018626 | 0.032297 | 0.0032286 | 0.99993 | 4.8066e-06 | 9.9993e-05 | 0.00039999 | OK | 1.0001 | 2.5951e-06 | -4.945e-06 | 6.1235e-05 | OK | 1 | 2.5951e-06 | 3.435e-05 | 0.00018808 | 0.0663 | 5.478e-06 | 0.50886 | 0.00020288 | 0.0579 | 8.1133e-06 | 0.0323 | 1.2046e-05
10 | m | 2 | -2.6081 | 0.0010473 | 5.0202 | 0.0018547 | m | 0.0016047 | 0.00094407 | 0.0002303 | 0.0016867 | m | -2.61 | 0.0052064 | 5.02 | 0.0047961 | m | OK | -2.61 |  | 5.02 |  | m | -6.4044e-08 | 0.0052064 | -7.6863e-08 | 0.0047961 | m | 0.00026267 |  | -4.9616e-06 |  | m | 2 | Hz | 0.001 | -0.0026097 | 1.41e-06 | 0.00502 | 2.5069e-06 | 0.0056578 | 2.3174e-06 | 2.0502 | 0.0003011 | 3 | 3m | 0.003 | 80 | 3 | 10m | 98 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | -2.4036e-08 | 0.0035532 | 3e-07 | 0.0027952 | 2Hz10m10m | 7.1523e-07 | 0.00022248 | -5.5883e-07 | 0.0037175 | 2.5666e-07 | 0.00022248 | -2.6088e-08 | 0.0031454 | -0.0026095 | 0.00022249 | 0.00502 | 0.0031454 | 0.99999 | 4.8867e-06 | 9.9999e-05 | 0.00094571 | OK | 1.0001 | 2.9068e-06 | -5.5905e-05 | 0.00037521 | OK | 1 | 2.9068e-06 | 4.0804e-05 | 0.0007291 | 0.005658 | 4.6349e-06 | 2.0503 | 0.00094564 | -0.00261 | 5.2064e-06 | 0.00502 | 4.7961e-06
100 | m | 20 | 66.9016 | 0.002055 | -48.0983 | 2.7403e-05 | m | -0.0019354 | 0.0014401 | 0.00072207 | 0.0012386 | m | 66.9 | 0.00597 | -48.1 | 0.0071703 | m | OK | 66.9 |  | -48.1 |  | m | -1.049e-06 | 0.00597 | -1.819e-06 | 0.0071703 | m | 0.0035472 |  | 0.00099749 |  | m | 20 | Hz | 0.001 | 0.066904 | 2.5094e-06 | -0.048099 | 1.2389e-06 | 0.082399 | 2.162e-06 | -0.62332 | 2.1565e-05 | 4 | 10m | 0.01 | 119 | 4 | 100m | 137 | 4 | 10m | 100m | 0.01 | 0.1 | 20Hz100m10m | 20Hz100m100m | 119 | 137 | 20Hz100m10m | -4.1245e-07 | 0.0030423 | 9.9995e-07 | 0.00018743 | 20Hz100m100m | -2.4634e-06 | 0.0029054 | -6.7817e-06 | 0.0013992 | -2.0623e-06 | 0.0029054 | -5.2599e-06 | 0.0011622 | 0.066901 | 0.0029054 | -0.048104 | 0.0011622 | 0.99996 | 3.6992e-06 | 0.0001 | 0.00024822 | OK | 0.99998 | 3.1351e-06 | -6.7801e-05 | 5.2214e-05 | OK | 0.99997 | 3.1351e-06 | -3.4985e-05 | 9.0545e-05 | 0.082397 | 4.3317e-06 | -0.62335 | 0.00010029 | 0.0669 | 5.97e-06 | -0.0481 | 7.1703e-06
3 | m | 2000 | -2.1905 | 0.00097539 | 2.0588 | 0.001917 | m | -0.00014516 | 0.001668 | -0.0012063 | 0.00055187 | m | -2.19 | 0.0047058 | 2.06 | 0.0047983 | m | Extrapolated | -2.19 |  | 2.06 |  | m | 3.743e-06 | 0.0047058 | 2.0417e-06 | 0.0047983 | m | -0.0003205 |  | -1.8294e-05 |  | m | 2 | kHz | 0.001 | -0.0021903 | 1.9322e-06 | 0.00206 | 1.9949e-06 | 0.0030068 | 1.9619e-06 | 2.3869 | 0.00065375 | 3 | 3m | 0.003 | 71 | 3 | 10m | 0 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | -1.7402e-08 | 0.0039544 | 9.9754e-08 | 0.001033 | 2kHz3m3m | -2.1785e-07 | 0.0030069 | -2.3991e-07 | 0.0032239 | -2.1854e-07 | 0.0030069 | -2.4107e-07 | 0.0032314 | -0.0021905 | 0.0030069 | 0.0020597 | 0.0032314 | 0.99998 | 4.1878e-06 | 9.9549e-05 | 0.0018511 | OK | 0.99993 | 4.734e-06 | -7.8203e-05 | 0.0012605 | OK | 0.99993 | 4.734e-06 | -7.881e-05 | 0.0012585 | 0.0030066 | 3.9238e-06 | 2.3868 | 0.0018148 | -0.00219 | 4.7058e-06 | 0.00206 | 4.7983e-06
3 | m | 5 | -2.4112 | 0.0013852 | 0.91815 | 0.0013935 | m | -0.001177 | 0.0003382 | -0.00079658 | 0.0012261 | m | -2.41 | 0.0036758 | 0.919 | 0.0063959 | m | OK | -2.41 |  | 0.919 |  | m | -3.8109e-08 | 0.0036758 | -2.275e-08 | 0.0063959 | m | -6.59e-05 |  | -5.0296e-05 |  | m | 5 | Hz | 0.001 | -0.0024101 | 1.4259e-06 | 0.00091895 | 1.8561e-06 | 0.0025793 | 1.4874e-06 | 2.7773 | 0.00070065 | 2 | 1m | 0.001 | 45 | 2 | 3m | 63 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | 6.6242e-08 | 0.0024231 | 9.9989e-08 | 0.0030299 | 5Hz3m3m | -1.1715e-07 | 0.0010589 | -1.8369e-07 | 0.003899 | -7.8575e-08 | 0.0010589 | -1.2402e-07 | 0.0037162 | -0.0024101 | 0.0010589 | 0.00091883 | 0.0037162 | 1.0001 | 3.258e-06 | 0.00010001 | 0.0043754 | OK | 0.99996 | 1.3632e-06 | -6.1223e-05 | 0.0016338 | OK | 0.99998 | 1.3632e-06 | -2.731e-05 | 0.0022105 | 0.0025793 | 2.9749e-06 | 2.7773 | 0.0026172 | -0.00241 | 3.6758e-06 | 0.000919 | 6.3959e-06
1 | m | 5 | -0.12052 | 4.225e-05 | -0.26474 | 0.00027568 | m | 0.0014762 | 0.0018357 | -0.0017433 | 0.0005964 | m | -0.12199 | 0.0031945 | -0.263 | 0.0022771 | m | OK | -0.122 |  | -0.263 |  | m | 7.7219e-06 | 0.0031945 | -3.5818e-06 | 0.0022771 | m | -1.121e-06 |  | 5.8429e-06 |  | m | 5 | Hz | 0.001 | -0.000122 | 1.8362e-06 | -0.00026299 | 6.5703e-07 | 0.00028991 | 9.7586e-07 | -2.0051 | 0.005824 | 1 | 0m | 0 | 9 | 1 | 1m | 27 | 1 | 0m | 1m | 0 | 0.001 | 5Hz1m0m | 5Hz1m1m | 9 | 27 | 5Hz1m0m | 0 | 0.0036843 | 0 | 0.00015903 | 5Hz1m1m | 5.7444e-08 | 0.0009626 | 4.1341e-08 | 0.0014432 | 1.6654e-08 | 0.0009626 | 4.1341e-08 | 0.00053134 | -0.00012198 | 0.00096261 | -0.00026295 | 0.00053134 | 1 | 3.6878e-06 | 4.1349e-05 | 0.0014488 | OK | 1.0001 | 1.3021e-06 | 4.1349e-05 | 0.0014488 | OK | 1 | 1.3021e-06 | 4.1349e-05 | 0.0014488 | 0.00028992 | 1.9517e-06 | -2.0051 | 0.011738 | -0.00012199 | 3.1945e-06 | -0.000263 | 2.2771e-06
3 | m | 1000 | -0.70396 | 0.001727 | 1.638 | 0.00088279 | m | 0.0019328 | 0.00062136 | -0.0019854 | 0.0013046 | m | -0.706 | 0.0058285 | 1.64 | 0.0038514 | m | OK | -0.706 |  | 1.64 |  | m | 4.6115e-07 | 0.0058285 | -1.3195e-07 | 0.0038514 | m | 0.00010756 |  | -2.5015e-05 |  | m | 1 | kHz | 0.001 | -0.00070589 | 1.8354e-06 | 0.00164 | 1.5752e-06 | 0.0017854 | 1.6186e-06 | 1.9773 | 0.0010066 | 2 | 1m | 0.001 | 52 | 2 | 3m | 70 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 7.748e-08 | 0.0015949 | 1.0078e-07 | 0.00016428 | 1kHz3m3m | -8.3541e-08 | 0.002271 | -8.5512e-08 | 0.0038778 | 1.4244e-08 | 0.002271 | 2.7622e-08 | 0.0016226 | -0.00070588 | 0.002271 | 0.00164 | 0.0016226 | 1.0001 | 3.804e-06 | 9.9887e-05 | 0.0038076 | OK | 0.99997 | 2.2768e-06 | -2.8192e-05 | 0.0013002 | OK | 1 | 2.2768e-06 | 4.9588e-05 | 0.0028229 | 0.0017855 | 3.2373e-06 | 1.9773 | 0.0034672 | -0.000706 | 5.8285e-06 | 0.00164 | 3.8514e-06
100 | m | 200 | 51.0018 | 0.0018814 | 51.4023 | 0.0016969 | m | 0.00049262 | 0.00043414 | 0.0012545 | 0.00125 | m | 51 | 0.0073856 | 51.4 | 0.0073455 | m | OK | 51 |  | 51.4 |  | m | 9.9373e-07 | 0.0073856 | -2.9421e-07 | 0.0073455 | m | 0.0013255 |  | 0.0010359 |  | m | 200 | Hz | 0.001 | 0.051001 | 1.9308e-06 | 0.051401 | 2.1076e-06 | 0.07241 | 2.0219e-06 | 0.7893 | 2.7903e-05 | 4 | 10m | 0.01 | 122 | 4 | 100m | 140 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | -6.8486e-07 | 0.0012358 | 9.989e-07 | 0.0023593 | 200Hz100m100m | -2.9553e-07 | 0.0039756 | -4.0027e-06 | 0.0032774 | -4.1488e-07 | 0.0039756 | -2.4694e-06 | 0.0029959 | 0.051001 | 0.0039756 | 0.051399 | 0.0029959 | 0.99993 | 2.465e-06 | 9.9999e-05 | 0.00025373 | OK | 1 | 3.9771e-06 | -4.0017e-05 | 6.1157e-05 | OK | 0.99998 | 3.9771e-06 | 2.9059e-06 | 0.00012019 | 0.072408 | 4.0539e-06 | 0.7893 | 0.00013252 | 0.051 | 7.3856e-06 | 0.0514 | 7.3455e-06
100 | m | 1000 | -5.668 | 0.0015314 | 90.3076 | 0.0019957 | m | 0.0015617 | 0.0018498 | 0.0018202 | 0.0008356 | m | -5.67 | 0.011235 | 90.3 | 0.0043889 | m | OK | -5.67 |  | 90.3 |  | m | 1.6629e-05 | 0.011235 | -1.6248e-06 | 0.0043889 | m | 0.00040209 |  | 0.0057688 |  | m | 1 | kHz | 0.001 | -0.0056696 | 2.4014e-06 | 0.090306 | 2.1636e-06 | 0.090484 | 2.1646e-06 | 1.6335 | 2.653e-05 | 4 | 10m | 0.01 | 124 | 4 | 100m | 142 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | -4.9574e-07 | 0.00072607 | 9.9715e-07 | 0.00037823 | 1kHz100m100m | -6.5227e-06 | 0.0025697 | -3.2015e-07 | 0.002343 | -5.8854e-06 | 0.0025697 | -1.8086e-07 | 0.0021353 | -0.0056755 | 0.0025697 | 0.090306 | 0.0021353 | 0.99995 | 3.0702e-06 | 9.9979e-05 | 0.00018724 | OK | 0.99993 | 3.4372e-06 | -2.5982e-06 | 0.00010366 | OK | 0.99994 | 3.4372e-06 | 8.2481e-06 | 0.0001125 | 0.090478 | 4.3403e-06 | 1.6335 | 0.00012439 | -0.00567 | 1.1235e-05 | 0.0903 | 4.3889e-06
1 | m | 1000 | 0.019589 | 0.0014255 | 0.0073393 | 0.00073466 | m | -0.001211 | 0.0010732 | -0.0017707 | 0.0014363 | m | 0.0208 | 0.0034796 | 0.0091104 | 0.0033231 | m | OK | 0.0208 |  | 0.00911 |  | m | -1.65e-07 | 0.0034796 | 3.7662e-07 | 0.0033231 | m | 2.4115e-08 |  | 1.5086e-10 |  | m | 1 | kHz | 0.001 | 2.08e-05 | 1.7843e-06 | 9.11e-06 | 1.6133e-06 | 2.2708e-05 | 1.7579e-06 | 0.41281 | 0.072311 | 1 | 0m | 0 | 16 | 1 | 1m | 34 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 0 | 0.0039248 | 0 | 0.0017447 | 1kHz1m1m | -4.3281e-08 | 0.0021583 | 1.8158e-08 | 0.0031997 | -9.8279e-10 | 0.0021583 | 1.8158e-08 | 0.0017777 | 2.0799e-05 | 0.0021583 | 9.1282e-06 | 0.0017777 | 1 | 4.2951e-06 | 1.8528e-05 | 0.003196 | OK | 0.99996 | 2.262e-06 | 1.8528e-05 | 0.003196 | OK | 1 | 2.262e-06 | 1.8528e-05 | 0.003196 | 2.2708e-05 | 3.5158e-06 | 0.41283 | 0.14466 | 2.08e-05 | 3.4796e-06 | 9.1104e-06 | 3.3231e-06
3 | m | 2 | 2.5417 | 0.00070997 | -0.44345 | 0.0012435 | m | 0.0016855 | 0.0012615 | 0.0014629 | 0.001014 | m | 2.54 | 0.0029618 | -0.445 | 0.0043925 | m | OK | 2.54 |  | -0.445 |  | m | -1.1425e-09 | 0.0029618 | -6.6043e-09 | 0.0043925 | m | -1.4472e-05 |  | 9.0976e-05 |  | m | 2 | Hz | 0.001 | 0.00254 | 1.4476e-06 | -0.00044491 | 1.6045e-06 | 0.0025787 | 1.4525e-06 | -0.1734 | 0.0006205 | 2 | 1m | 0.001 | 44 | 2 | 3m | 62 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | 1.1326e-08 | 0.0026983 | 9.9999e-08 | 0.0014561 | 2Hz3m3m | 3.5078e-08 | 0.0038536 | -2.0848e-07 | 0.0018577 | 3.0074e-08 | 0.0038536 | -1.4349e-07 | 0.0017731 | 0.00254 | 0.0038536 | -0.00044505 | 0.0017731 | 1 | 2.7064e-06 | 0.0001 | 0.0028826 | OK | 1 | 5.491e-06 | -6.9494e-05 | 0.00073582 | OK | 1 | 5.491e-06 | -3.3785e-05 | 0.0011881 | 0.0025787 | 2.905e-06 | -0.17344 | 0.001718 | 0.00254 | 2.9618e-06 | -0.000445 | 4.3925e-06
100 | m | 5000 | -5.7587 | 0.00025551 | 16.3022 | 0.00023822 | m | 3.9735e-05 | 0.0010027 | 0.0011635 | 0.00047404 | m | -5.76 | 0.0085516 | 16.3 | 0.0032298 | m | OK | -5.76 |  | 16.3 |  | m | 1.0816e-05 | 0.0085516 | 3.458e-06 | 0.0032298 | m | 0.001212 |  | 0.0010742 |  | m | 5 | kHz | 0.001 | -0.0057588 | 1.0347e-06 | 0.016301 | 5.3053e-07 | 0.017288 | 6.0748e-07 | 1.9104 | 5.7352e-05 | 4 | 10m | 0.01 | 126 | 4 | 100m | 144 | 4 | 10m | 100m | 0.01 | 0.1 | 5kHz100m10m | 5kHz100m100m | 126 | 144 | 5kHz100m10m | -4.0201e-07 | 0.0036149 | 9.9975e-07 | 0.00012206 | 5kHz100m100m | -3.7561e-06 | 0.001926 | -6.999e-06 | 0.0021307 | -6.7363e-07 | 0.001926 | 3.5199e-07 | 0.00028472 | -0.0057595 | 0.001926 | 0.016301 | 0.00028473 | 0.99997 | 5.2627e-06 | 9.962e-05 | 0.00051319 | OK | 0.99996 | 5.5463e-06 | -6.6801e-05 | 0.00049029 | OK | 0.99996 | 5.5463e-06 | 8.6143e-05 | 0.00051134 | 0.017288 | 1.2187e-06 | 1.9105 | 0.00052405 | -0.00576 | 8.5516e-06 | 0.0163 | 3.2298e-06
3 | m | 20 | -0.89426 | 9.1813e-05 | 1.8299 | 0.0014185 | m | 0.00061657 | 0.0019593 | -0.00030647 | 3.8272e-05 | m | -0.895 | 0.0055801 | 1.83 | 0.0038279 | m | OK | -0.895 |  | 1.83 |  | m | 8.6943e-09 | 0.0055801 | -1.0033e-08 | 0.0038279 | m | 0.00012584 |  | 0.00017882 |  | m | 20 | Hz | 0.001 | -0.00089487 | 1.9615e-06 | 0.0018302 | 1.419e-06 | 0.0020372 | 1.5387e-06 | 2.0256 | 0.00091748 | 2 | 1m | 0.001 | 47 | 2 | 3m | 65 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | -2.7333e-08 | 0.0024512 | 1e-07 | 0.0025754 | 20Hz3m3m | -2.232e-07 | 0.0014029 | 2.6563e-07 | 0.00078242 | -1.2891e-07 | 0.0014029 | 1.859e-07 | 0.0016455 | -0.000895 | 0.0014029 | 0.0018304 | 0.0016455 | 0.99997 | 4.1268e-06 | 0.0001 | 0.0044716 | OK | 0.99993 | 1.6072e-06 | 8.8531e-05 | 0.00032403 | OK | 0.99995 | 1.6072e-06 | 9.4052e-05 | 0.0023206 | 0.0020371 | 3.0773e-06 | 2.0257 | 0.0029584 | -0.000895 | 5.5801e-06 | 0.00183 | 3.8279e-06
1 | m | 0.05 | -0.074209 | 0.00044813 | 0.17967 | 0.0019745 | m | -0.00040606 | 0.001143 | 0.00066866 | 0.0015472 | m | -0.073787 | 0.0032937 | 0.17901 | 0.0045197 | m | OK | -0.0738 |  | 0.179 |  | m | 1.3099e-05 | 0.0032937 | 5.4e-06 | 0.0045197 | m | -2.7853e-06 |  | -2.1701e-06 |  | m | 50 | mHz | 0.001 | -7.3803e-05 | 1.2277e-06 | 0.000179 | 2.5085e-06 | 0.00019362 | 2.3658e-06 | 1.9619 | 0.0076653 | 1 | 0m | 0 | 3 | 1 | 1m | 21 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.0036718 | 0 | 0.0029355 | 50mHz1m1m | 2.5191e-08 | 0.0031843 | -9.0742e-08 | 0.0010893 | 4.8774e-09 | 0.0031843 | -9.0742e-08 | 0.002578 | -7.3798e-05 | 0.0031843 | 0.00017891 | 0.002578 | 1 | 4.7009e-06 | -9.075e-05 | 0.0014316 | OK | 1 | 3.6855e-06 | -9.075e-05 | 0.0014316 | OK | 1 | 3.6855e-06 | -9.075e-05 | 0.0014316 | 0.00019362 | 4.7317e-06 | 1.9618 | 0.015397 | -7.3787e-05 | 3.2937e-06 | 0.00017901 | 4.5197e-06
3 | m | 10 | 1.921 | 0.00041736 | 1.3084 | 4.8226e-05 | m | 0.0010629 | 0.0013059 | -0.001585 | 0.00072778 | m | 1.92 | 0.0034566 | 1.31 | 0.004359 | m | OK | 1.92 |  | 1.31 |  | m | 7.5325e-08 | 0.0034566 | -2.6017e-08 | 0.004359 | m | -4.9268e-05 |  | -6.3044e-05 |  | m | 10 | Hz | 0.001 | 0.00192 | 1.3709e-06 | 0.0013099 | 7.2937e-07 | 0.0023243 | 1.2048e-06 | 0.59873 | 0.00042155 | 2 | 1m | 0.001 | 46 | 2 | 3m | 64 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | -2.7706e-08 | 0.0029754 | 9.9999e-08 | 0.0034496 | 10Hz3m3m | 1.9116e-07 | 0.0014313 | -1.0584e-07 | 0.0032405 | 1.1721e-07 | 0.0014313 | -3.6292e-08 | 0.0033112 | 0.0019201 | 0.0014313 | 0.0013099 | 0.0033112 | 0.99997 | 2.994e-06 | 0.00010001 | 0.0036882 | OK | 1.0001 | 3.285e-06 | -3.5278e-05 | 0.001117 | OK | 1 | 3.285e-06 | 1.0432e-05 | 0.0019857 | 0.0023243 | 2.4095e-06 | 0.59874 | 0.0021573 | 0.00192 | 3.4566e-06 | 0.00131 | 4.359e-06
3 | m | 2 | 1.882 | 0.0018869 | -1.1419 | 0.0015522 | m | 0.0019929 | 0.00021269 | -0.0018663 | 7.7913e-05 | m | 1.88 | 0.0039826 | -1.14 | 0.0045317 | m | OK | 1.88 |  | -1.14 |  | m | -1.3027e-09 | 0.0039826 | -2.1588e-09 | 0.0045317 | m | -1.9911e-05 |  | 1.6126e-05 |  | m | 2 | Hz | 0.001 | 0.00188 | 1.8988e-06 | -0.00114 | 1.5542e-06 | 0.0021986 | 1.8126e-06 | -0.5451 | 0.00075225 | 2 | 1m | 0.001 | 44 | 2 | 3m | 62 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | 1.1326e-08 | 0.0026983 | 9.9999e-08 | 0.0014561 | 2Hz3m3m | 3.5078e-08 | 0.0038536 | -2.0848e-07 | 0.0018577 | 2.5561e-08 | 0.0038536 | -8.4876e-08 | 0.0016968 | 0.00188 | 0.0038536 | -0.0011401 | 0.0016968 | 1 | 2.7064e-06 | 0.0001 | 0.0028826 | OK | 1 | 5.491e-06 | -6.9494e-05 | 0.00073582 | OK | 1 | 5.491e-06 | -1.577e-06 | 0.001596 | 0.0021986 | 3.6252e-06 | -0.5451 | 0.0021933 | 0.00188 | 3.9826e-06 | -0.00114 | 4.5317e-06
100 | m | 0.05 | 31.4043 | 0.0021058 | 50.7968 | 0.0013774 | m | 0.00034346 | 0.0012294 | 0.0015469 | 0.00054615 | m | 31.4 | 0.011193 | 50.8 | 0.0074777 | m | OK | 31.4 |  | 50.8 |  | m | 2.7346e-08 | 0.011193 | 5.0133e-08 | 0.0074777 | m | 0.0040061 |  | -0.0047685 |  | m | 50 | mHz | 0.001 | 0.031404 | 2.4384e-06 | 0.050795 | 1.4817e-06 | 0.059719 | 1.7979e-06 | 1.0171 | 3.71e-05 | 4 | 10m | 0.01 | 111 | 4 | 100m | 129 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | 4.7128e-07 | 0.0013032 | 9.9997e-07 | 0.0033159 | 50mHz100m100m | 2.0948e-06 | 0.0032386 | 9.8259e-06 | 0.0035729 | 1.3682e-06 | 0.0032386 | 5.8757e-06 | 0.0034579 | 0.031405 | 0.0032386 | 0.050801 | 0.0034579 | 1 | 3.0828e-06 | 0.0001 | 0.00037978 | OK | 1 | 5.0094e-06 | 9.827e-05 | 6.1747e-05 | OK | 1 | 5.0094e-06 | 9.9044e-05 | 0.00020409 | 0.059721 | 3.6083e-06 | 1.0172 | 0.00021716 | 0.0314 | 1.1193e-05 | 0.0508 | 7.4777e-06
3 | m | 100 | 2.4183 | 0.001826 | -0.55534 | 0.00083419 | m | -0.0016367 | 7.8273e-05 | 0.00086325 | 0.0014383 | m | 2.42 | 0.0037511 | -0.556 | 0.005372 | m | OK | 2.42 |  | -0.556 |  | m | -9.1583e-09 | 0.0037511 | -1.3324e-09 | 0.005372 | m | -4.7293e-05 |  | -0.00020677 |  | m | 100 | Hz | 0.001 | 0.00242 | 1.8277e-06 | -0.00055621 | 1.6627e-06 | 0.002483 | 1.8198e-06 | -0.22592 | 0.0006731 | 2 | 1m | 0.001 | 49 | 2 | 3m | 67 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | 4.5086e-08 | 0.0005826 | 1.0004e-07 | 0.0026998 | 100Hz3m3m | -4.7835e-08 | 0.0016958 | 2.41e-07 | 0.003416 | -2.3817e-08 | 0.0016958 | 2.0456e-07 | 0.0032309 | 0.0024199 | 0.0016958 | -0.000556 | 0.0032309 | 1 | 1.7597e-06 | 0.0001 | 0.0027087 | OK | 0.99998 | 2.2974e-06 | 8.0342e-05 | 0.0013924 | OK | 1 | 2.2974e-06 | 8.5423e-05 | 0.0017326 | 0.002483 | 3.6396e-06 | -0.22583 | 0.0021941 | 0.00242 | 3.7511e-06 | -0.000556 | 5.372e-06
100 | m | 0.05 | -57.6987 | 0.0013328 | -12.1922 | 7.056e-05 | m | 0.00059371 | 0.00045926 | 0.0017027 | 0.0013242 | m | -57.7 | 0.003784 | -12.2 | 0.012222 | m | OK | -57.7 |  | -12.2 |  | m | -5.3643e-08 | 0.003784 | -1.441e-08 | 0.012222 | m | 0.00068787 |  | 0.0061165 |  | m | 50 | mHz | 0.001 | -0.057699 | 1.4098e-06 | -0.012194 | 1.3261e-06 | 0.058974 | 1.4063e-06 | -2.9333 | 2.2548e-05 | 4 | 10m | 0.01 | 111 | 4 | 100m | 129 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | 4.7128e-07 | 0.0013032 | 9.9997e-07 | 0.0033159 | 50mHz100m100m | 2.0948e-06 | 0.0032386 | 9.8259e-06 | 0.0035729 | 1.3548e-06 | 0.0032386 | 5.8026e-06 | 0.0034557 | -0.057698 | 0.0032386 | -0.012188 | 0.0034557 | 1 | 3.0828e-06 | 0.0001 | 0.00037978 | OK | 1 | 5.0094e-06 | 9.827e-05 | 6.1747e-05 | OK | 1 | 5.0094e-06 | 9.9059e-05 | 0.00020672 | 0.058976 | 2.828e-06 | -2.9332 | 0.00021158 | -0.0577 | 3.784e-06 | -0.0122 | 1.2222e-05
10 | m | 0.5 | 4.5698 | 0.0014832 | 0.23116 | 0.0012358 | m | -0.00038736 | 0.00094006 | -0.00054027 | 0.0013896 | m | 4.57 | 0.0035154 | 0.232 | 0.0044633 | m | OK | 4.57 |  | 0.232 |  | m | 6.6415e-09 | 0.0035154 | -1.91e-08 | 0.0044633 | m | 0.00015631 |  | -0.00029841 |  | m | 500 | mHz | 0.001 | 0.0045702 | 1.756e-06 | 0.0002317 | 1.8596e-06 | 0.004576 | 1.7563e-06 | 0.050655 | 0.00040633 | 3 | 3m | 0.003 | 78 | 3 | 10m | 96 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | -1.1635e-07 | 0.00232 | 3.0001e-07 | 0.0016707 | 500mHz10m10m | -3.3994e-08 | 0.00076531 | -4.7206e-07 | 0.0033821 | -9.7805e-08 | 0.00076531 | 1.2618e-07 | 0.002056 | 0.0045701 | 0.00076532 | 0.00023183 | 0.002056 | 0.99996 | 3.3039e-06 | 9.9997e-05 | 0.00055959 | OK | 1 | 3.0413e-06 | -4.7205e-05 | 0.00047381 | OK | 0.99997 | 3.0413e-06 | 6.6855e-05 | 0.00054028 | 0.0045759 | 3.5126e-06 | 0.050722 | 0.00097587 | 0.00457 | 3.5154e-06 | 0.000232 | 4.4633e-06
100 | m | 0.1 | -19.6 | 0.001007 | -48.2014 | 0.0015385 | m | 0.001365 | 0.00051735 | -0.00020819 | 0.00078004 | m | -19.6 | 0.010298 | -48.2 | 0.005171 | m | OK | -19.6 |  | -48.2 |  | m | -4.3272e-07 | 0.010298 | -1.0291e-07 | 0.005171 | m | -0.0013585 |  | -0.0011811 |  | m | 100 | mHz | 0.001 | -0.019601 | 1.1321e-06 | -0.048201 | 1.725e-06 | 0.052034 | 1.6538e-06 | -1.957 | 2.371e-05 | 4 | 10m | 0.01 | 112 | 4 | 100m | 130 | 4 | 10m | 100m | 0.01 | 0.1 | 100mHz100m10m | 100mHz100m100m | 112 | 130 | 100mHz100m10m | -6.9929e-07 | 0.0027568 | 1.0001e-06 | 0.0015827 | 100mHz100m100m | 1.3731e-06 | 0.0055328 | -8.0654e-06 | 0.0031956 | 2.6861e-07 | 0.0055328 | -3.2339e-06 | 0.002336 | -0.019601 | 0.0055328 | -0.048204 | 0.002336 | 0.99993 | 4.0692e-06 | 9.9999e-05 | 0.00033583 | OK | 1 | 5.6654e-06 | -8.0649e-05 | 5.9337e-05 | OK | 0.99997 | 5.6654e-06 | 1.5628e-05 | 0.0002067 | 0.052033 | 3.3208e-06 | -1.957 | 0.00021207 | -0.0196 | 1.0298e-05 | -0.0482 | 5.171e-06
1 | m | 10 | 0.00014678 | 0.00063858 | 0.32771 | 0.00068212 | m | -0.0019064 | 0.0019005 | 0.0016998 | 0.0017909 | m | 0.0020741 | 0.004091 | 0.326 | 0.0038329 | m | OK | 0.00206 |  | 0.326 |  | m | 1.411e-05 | 0.004091 | -8.8467e-08 | 0.0038329 | m | -6.7897e-06 |  | 5.8813e-06 |  | m | 10 | Hz | 0.001 | 2.0532e-06 | 2.005e-06 | 0.00032601 | 1.9164e-06 | 0.00032601 | 1.9164e-06 | 1.5645 | 0.0061499 | 1 | 0m | 0 | 10 | 1 | 1m | 28 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0001832 | 0 | 0.0021009 | 10Hz1m1m | -5.4917e-08 | 0.00018788 | -6.4216e-08 | 0.0024586 | -1.7904e-08 | 0.00018788 | -6.4216e-08 | 0.0022175 | 2.0353e-06 | 0.00018792 | 0.00032594 | 0.0022175 | 1 | 2.1088e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99995 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99998 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | 0.00032601 | 3.8329e-06 | 1.5644 | 0.012549 | 2.0741e-06 | 4.091e-06 | 0.000326 | 3.8329e-06
3 | m | 1 | -1.18 | 0.0011018 | 0.90946 | 0.0019336 | m | -1.9277e-05 | 0.00072773 | 0.0013487 | 0.0013164 | m | -1.18 | 0.0050077 | 0.908 | 0.0058074 | m | OK | -1.18 |  | 0.908 |  | m | 3.2373e-09 | 0.0050077 | 8.1909e-09 | 0.0058074 | m | 5.9284e-05 |  | 0.000116 |  | m | 1 | Hz | 0.001 | -0.0011799 | 1.3205e-06 | 0.00090812 | 2.3392e-06 | 0.0014889 | 1.7693e-06 | 2.4856 | 0.0013574 | 2 | 1m | 0.001 | 43 | 2 | 3m | 61 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | -1.9849e-08 | 0.00034442 | 1.0002e-07 | 0.0026891 | 1Hz3m3m | -1.1752e-08 | 0.00047154 | 1.285e-07 | 0.000646 | -1.7869e-08 | 0.00047154 | 1.0698e-07 | 0.0021896 | -0.00118 | 0.00047154 | 0.00090822 | 0.0021896 | 0.99998 | 3.6518e-07 | 9.9995e-05 | 0.0046729 | OK | 1 | 7.7074e-07 | 4.2828e-05 | 0.0005917 | OK | 0.99998 | 7.7074e-07 | 8.6019e-05 | 0.0036751 | 0.0014889 | 3.5386e-06 | 2.4857 | 0.0045691 | -0.00118 | 5.0077e-06 | 0.000908 | 5.8074e-06
3 | m | 0.01 | -0.72105 | 0.00024078 | -0.99125 | 0.00045915 | m | 0.0010823 | 0.0014178 | -0.0012688 | 0.00098749 | m | -0.722 | 0.0042045 | -0.99 | 0.0034924 | m | OK | -0.722 |  | -0.99 |  | m | -1.5385e-09 | 0.0042045 | 3.8437e-10 | 0.0034924 | m | -0.0001327 |  | 1.6003e-05 |  | m | 10 | mHz | 0.001 | -0.00072213 | 1.4381e-06 | -0.00098998 | 1.089e-06 | 0.0012254 | 1.2216e-06 | -2.201 | 0.0010832 | 2 | 1m | 0.001 | 37 | 2 | 3m | 55 | 2 | 1m | 3m | 0.001 | 0.003 | 10mHz3m1m | 10mHz3m3m | 37 | 55 | 10mHz3m1m | -5.5112e-08 | 0.0034877 | 1.0001e-07 | 0.00071514 | 10mHz3m3m | -1.1623e-07 | 0.0016304 | 1.7192e-07 | 0.0032608 | -6.1999e-08 | 0.0016304 | 1.0811e-07 | 0.001002 | -0.00072219 | 0.0016304 | -0.00098988 | 0.001002 | 0.99994 | 4.2004e-06 | 9.9998e-05 | 0.0035946 | OK | 0.99996 | 3.0792e-06 | 5.7307e-05 | 0.0014322 | OK | 0.99995 | 3.0792e-06 | 9.5188e-05 | 0.0033509 | 0.0012253 | 2.4432e-06 | -2.2009 | 0.0039902 | -0.000722 | 4.2045e-06 | -0.00099 | 3.4924e-06
3 | m | 0.1 | 0.38908 | 0.00018174 | 1.209 | 1.9944e-05 | m | -0.0019985 | 0.0019826 | -0.0009486 | 0.0019464 | m | 0.391 | 0.0054182 | 1.21 | 0.0040863 | m | OK | 0.391 |  | 1.21 |  | m | 2.2537e-09 | 0.0054182 | 2.1323e-12 | 0.0040863 | m | 7.4772e-05 |  | -8.3234e-05 |  | m | 100 | mHz | 0.001 | 0.00039107 | 1.9909e-06 | 0.0012099 | 1.9465e-06 | 0.0012715 | 1.9507e-06 | 1.2582 | 0.0015625 | 2 | 1m | 0.001 | 40 | 2 | 3m | 58 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 6.0542e-08 | 0.0019825 | 9.9992e-08 | 0.0031027 | 100mHz3m3m | -1.7924e-07 | 0.0040053 | -2.2851e-07 | 2.3224e-05 | 2.7985e-08 | 0.0040053 | 5.539e-08 | 0.0026846 | 0.0003911 | 0.0040053 | 0.00121 | 0.0026846 | 1.0001 | 1.9826e-06 | 9.9998e-05 | 0.0034787 | OK | 0.99994 | 4.8158e-06 | -7.6159e-05 | 0.00032186 | OK | 1 | 4.8158e-06 | 7.6081e-05 | 0.0030501 | 0.0012716 | 3.9015e-06 | 1.2582 | 0.0043667 | 0.000391 | 5.4182e-06 | 0.00121 | 4.0863e-06
1 | m | 200 | -0.050963 | 0.00023671 | 0.013023 | 0.0013854 | m | -0.00056311 | 0.00055 | -0.0010776 | 0.00094148 | m | -0.050401 | 0.0016598 | 0.014098 | 0.0031533 | m | OK | -0.0504 |  | 0.0141 |  | m | -5.9414e-07 | 0.0016598 | -2.1239e-06 | 0.0031533 | m | -1.1926e-07 |  | 1.5983e-07 |  | m | 200 | Hz | 0.001 | -5.04e-05 | 5.9878e-07 | 1.41e-05 | 1.675e-06 | 5.2335e-05 | 7.3224e-07 | 2.8688 | 0.030976 | 1 | 0m | 0 | 14 | 1 | 1m | 32 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 0.00093895 | 0 | 0.0021931 | 200Hz1m1m | -5.7742e-08 | 0.0028813 | 4.4363e-08 | 0.0018199 | -3.0219e-09 | 0.0028813 | 4.4363e-08 | 0.0021735 | -5.0403e-05 | 0.0028813 | 1.4145e-05 | 0.0021735 | 1 | 2.3856e-06 | 4.4468e-05 | 0.0038907 | OK | 0.99994 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | OK | 1 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | 5.2335e-05 | 1.4645e-06 | 2.8688 | 0.062074 | -5.0401e-05 | 1.6598e-06 | 1.4098e-05 | 3.1533e-06
3 | m | 2000 | 0.13901 | 0.00022034 | -1.1204 | 0.00013465 | m | 0.00010944 | 0.0002271 | -0.00040237 | 0.00081565 | m | 0.139 | 0.0021434 | -1.12 | 0.0016515 | m | OK | 0.139 |  | -1.12 |  | m | -5.9283e-07 | 0.0021434 | 1.1573e-07 | 0.0016515 | m | -9.6615e-05 |  | -3.3662e-05 |  | m | 2 | kHz | 0.001 | 0.0001389 | 3.1643e-07 | -0.00112 | 8.2669e-07 | 0.0011286 | 8.2133e-07 | -1.4474 | 0.00029248 | 2 | 1m | 0.001 | 53 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | -1.7402e-08 | 0.0039544 | 9.9754e-08 | 0.001033 | 2kHz3m3m | -2.1785e-07 | 0.0030069 | -2.3991e-07 | 0.0032239 | -3.0292e-08 | 0.0030069 | 7.7911e-08 | 0.0011739 | 0.00013887 | 0.0030069 | -0.00112 | 0.0011739 | 0.99998 | 4.1878e-06 | 9.9549e-05 | 0.0018511 | OK | 0.99993 | 4.734e-06 | -7.8203e-05 | 0.0012605 | OK | 0.99998 | 4.734e-06 | 8.8118e-05 | 0.0018132 | 0.0011286 | 1.6427e-06 | -1.4473 | 0.0019052 | 0.000139 | 2.1434e-06 | -0.00112 | 1.6515e-06
10 | m | 0.01 | -7.7383 | 0.00042334 | 4.8108 | 0.00078795 | m | 0.0016384 | 0.001763 | -0.00021378 | 0.00081799 | m | -7.74 | 0.0038805 | 4.81 | 0.0046507 | m | OK | -7.74 |  | 4.81 |  | m | 4.0063e-08 | 0.0038805 | -2.3374e-08 | 0.0046507 | m | 0.00010248 |  | 0.00098258 |  | m | 10 | mHz | 0.001 | -0.0077399 | 1.8131e-06 | 0.004811 | 1.1358e-06 | 0.0091133 | 1.6525e-06 | 2.5855 | 0.00014911 | 3 | 3m | 0.003 | 73 | 3 | 10m | 91 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | 2.6834e-07 | 0.0037136 | 2.9997e-07 | 0.00072218 | 10mHz10m10m | -6.7217e-07 | 0.0028332 | 9.7161e-07 | 0.0037958 | -5.5303e-07 | 0.0028332 | 8.8653e-07 | 0.0034064 | -0.0077405 | 0.0028332 | 0.0048119 | 0.0034064 | 1.0001 | 3.723e-06 | 0.0001 | 0.00097788 | OK | 0.99993 | 3.8128e-06 | 9.715e-05 | 0.00039681 | OK | 0.99995 | 3.8128e-06 | 9.7511e-05 | 0.00047042 | 0.0091128 | 3.3052e-06 | 2.5856 | 0.00055699 | -0.00774 | 3.8805e-06 | 0.00481 | 4.6507e-06
1 | m | 10 | 0.68756 | 3.1069e-05 | -0.77573 | 0.00021542 | m | 0.00047183 | 0.0014762 | 0.00026908 | 0.0012008 | m | 0.687 | 0.0033291 | -0.776 | 0.0031995 | m | Extrapolated | 0.687 |  | -0.776 |  | m | 1.8296e-06 | 0.0033291 | 1.6093e-06 | 0.0031995 | m | 9.0777e-05 |  | 1.5582e-06 |  | m | 10 | Hz | 0.001 | 0.00068709 | 1.4765e-06 | -0.000776 | 1.22e-06 | 0.0010365 | 1.3388e-06 | -0.84609 | 0.0013215 | 2 | 1m | 0.001 | 28 | 2 | 3m | 0 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0001832 | 0 | 0.0021009 | 10Hz1m1m | -5.4917e-08 | 0.00018788 | -6.4216e-08 | 0.0024586 | -5.692e-08 | 0.00018788 | -6.4216e-08 | 0.0024716 | 0.00068703 | 0.0001879 | -0.00077606 | 0.0024716 | 1 | 2.1088e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99995 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | OK | 0.99994 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | 0.0010364 | 2.6776e-06 | -0.84615 | 0.0036295 | 0.000687 | 3.3291e-06 | -0.000776 | 3.1995e-06
1 | m | 1000 | 0.76825 | 0.0014279 | -0.49515 | 0.0013777 | m | -0.00076807 | 0.00063564 | 0.00088507 | 0.00077859 | m | 0.769 | 0.0035194 | -0.496 | 0.0039944 | m | OK | 0.769 |  | -0.496 |  | m | 7.4143e-07 | 0.0035194 | 1.2669e-06 | 0.0039944 | m | 2.1835e-05 |  | -3.2514e-05 |  | m | 1 | kHz | 0.001 | 0.00076902 | 1.563e-06 | -0.00049603 | 1.5825e-06 | 0.00091512 | 1.5687e-06 | -0.57286 | 0.001723 | 1 | 0m | 0 | 16 | 1 | 1m | 34 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 0 | 0.0039248 | 0 | 0.0017447 | 1kHz1m1m | -4.3281e-08 | 0.0021583 | 1.8158e-08 | 0.0031997 | -3.9607e-08 | 0.0021583 | 1.8158e-08 | 0.0030762 | 0.00076898 | 0.0021583 | -0.00049601 | 0.0030762 | 1 | 4.2951e-06 | 1.8528e-05 | 0.003196 | OK | 0.99996 | 2.262e-06 | 1.8528e-05 | 0.003196 | OK | 0.99996 | 2.262e-06 | 1.8528e-05 | 0.003196 | 0.00091508 | 3.1375e-06 | -0.57285 | 0.0047 | 0.000769 | 3.5194e-06 | -0.000496 | 3.9944e-06
3 | m | 0.5 | -0.38493 | 6.2499e-05 | -1.2016 | 0.001026 | m | -0.0018022 | 0.00020433 | -0.0015539 | 0.0014568 | m | -0.383 | 0.0044213 | -1.2 | 0.0035153 | m | OK | -0.383 |  | -1.2 |  | m | -4.0508e-09 | 0.0044213 | 1.3297e-09 | 0.0035153 | m | -0.00012444 |  | -5.8845e-05 |  | m | 500 | mHz | 0.001 | -0.00038312 | 2.1367e-07 | -0.0012001 | 1.7818e-06 | 0.0012597 | 1.6986e-06 | -1.8798 | 0.00045952 | 2 | 1m | 0.001 | 42 | 2 | 3m | 60 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | -7.4389e-08 | 0.00055263 | 9.9998e-08 | 0.00059317 | 500mHz3m3m | -2.2698e-07 | 0.0017093 | -1.643e-07 | 0.00091673 | -9.4205e-08 | 0.0017093 | 6.5675e-08 | 0.00063518 | -0.00038322 | 0.0017093 | -0.0012 | 0.00063519 | 0.99993 | 7.8698e-07 | 9.9998e-05 | 0.0038777 | OK | 0.99992 | 1.8266e-06 | -5.4754e-05 | 0.00068081 | OK | 0.99993 | 1.8266e-06 | 7.9901e-05 | 0.0034626 | 0.0012596 | 3.3973e-06 | -1.8797 | 0.0035825 | -0.000383 | 4.4213e-06 | -0.0012 | 3.5153e-06
3 | m | 0.1 | 1.421 | 0.00091461 | 2.0898 | 0.001882 | m | 0.0010608 | 0.0016324 | -0.00027292 | 5.6992e-05 | m | 1.42 | 0.0043668 | 2.09 | 0.0040506 | m | OK | 1.42 |  | 2.09 |  | m | 6.0082e-09 | 0.0043668 | -3.7579e-10 | 0.0040506 | m | -2.7626e-05 |  | 0.00011438 |  | m | 100 | mHz | 0.001 | 0.00142 | 1.8711e-06 | 0.0020901 | 1.8828e-06 | 0.0025268 | 1.8791e-06 | 0.97405 | 0.00074197 | 2 | 1m | 0.001 | 40 | 2 | 3m | 58 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 6.0542e-08 | 0.0019825 | 9.9992e-08 | 0.0031027 | 100mHz3m3m | -1.7924e-07 | 0.0040053 | -2.2851e-07 | 2.3224e-05 | -1.2251e-07 | 0.0040053 | -1.5079e-07 | 0.00075178 | 0.0014198 | 0.0040053 | 0.00209 | 0.00075179 | 1.0001 | 1.9826e-06 | 9.9998e-05 | 0.0034787 | OK | 0.99994 | 4.8158e-06 | -7.6159e-05 | 0.00032186 | OK | 0.99997 | 4.8158e-06 | -3.4483e-05 | 0.0010687 | 0.0025268 | 3.7583e-06 | 0.97401 | 0.0018287 | 0.00142 | 4.3668e-06 | 0.00209 | 4.0506e-06
1 | m | 200 | -0.42839 | 0.00023757 | 0.044165 | 0.0017459 | m | -0.0003766 | 0.0019809 | 0.00045553 | 0.00076459 | m | -0.428 | 0.0039902 | 0.043689 | 0.0041632 | m | OK | -0.428 |  | 0.0437 |  | m | -1.107e-06 | 0.0039902 | -1.0844e-05 | 0.0041632 | m | -9.7805e-06 |  | 9.2722e-06 |  | m | 200 | Hz | 0.001 | -0.00042801 | 1.9951e-06 | 4.3709e-05 | 1.906e-06 | 0.00043024 | 1.9942e-06 | 3.0398 | 0.0044323 | 1 | 0m | 0 | 14 | 1 | 1m | 32 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 0.00093895 | 0 | 0.0021931 | 200Hz1m1m | -5.7742e-08 | 0.0028813 | 4.4363e-08 | 0.0018199 | -2.4843e-08 | 0.0028813 | 4.4363e-08 | 0.0020325 | -0.00042803 | 0.0028813 | 4.3754e-05 | 0.0020325 | 1 | 2.3856e-06 | 4.4468e-05 | 0.0038907 | OK | 0.99994 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | OK | 0.99998 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | 0.00043023 | 3.9884e-06 | 3.0399 | 0.0096809 | -0.000428 | 3.9902e-06 | 4.3689e-05 | 4.1632e-06
3 | m | 2 | 1.0702 | 0.00028043 | -2.4018 | 0.0014285 | m | 0.00013074 | 0.0017187 | -0.001849 | 0.0015013 | m | 1.07 | 0.0045774 | -2.4 | 0.0041543 | m | OK | 1.07 |  | -2.4 |  | m | -6.6813e-09 | 0.0045774 | -2.9959e-09 | 0.0041543 | m | 7.8611e-05 |  | 6.8493e-05 |  | m | 2 | Hz | 0.001 | 0.0010701 | 1.7414e-06 | -0.0023999 | 2.0723e-06 | 0.0026277 | 2.0212e-06 | -1.1514 | 0.0006852 | 2 | 1m | 0.001 | 44 | 2 | 3m | 62 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | 1.1326e-08 | 0.0026983 | 9.9999e-08 | 0.0014561 | 2Hz3m3m | 3.5078e-08 | 0.0038536 | -2.0848e-07 | 0.0018577 | 3.0656e-08 | 0.0038536 | -1.5106e-07 | 0.001783 | 0.0010701 | 0.0038536 | -0.0024001 | 0.001783 | 1 | 2.7064e-06 | 0.0001 | 0.0028826 | OK | 1 | 5.491e-06 | -6.9494e-05 | 0.00073582 | OK | 1 | 5.491e-06 | -3.7941e-05 | 0.0011354 | 0.0026277 | 4.0424e-06 | -1.1514 | 0.0017797 | 0.00107 | 4.5774e-06 | -0.0024 | 4.1543e-06
100 | m | 2000 | -45.7008 | -0.00018204 | -84.6907 | -0.00036978 | m | 0.00057785 | 0.0011706 | 0.0010892 | 0.00029013 | m | -45.7 | 0.018222 | -84.7 | 0.0099077 | m | OK | -45.7 |  | -84.7 |  | m | -4.1147e-05 | 0.018222 | -2.1383e-05 | 0.0099077 | m | -0.0013302 |  | 0.0082346 |  | m | 2 | kHz | 0.001 | -0.045701 | 1.1846e-06 | -0.084692 | 4.7001e-07 | 0.096236 | 6.9827e-07 | -2.0656 | 1.1079e-05 | 4 | 10m | 0.01 | 125 | 4 | 100m | 143 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | -2.5323e-08 | 0.00057942 | 1.0047e-06 | 0.00010185 | 2kHz100m100m | 7.1609e-06 | 0.0019925 | 5.2033e-06 | 0.0027879 | 6.8603e-06 | 0.0019925 | 5.0277e-06 | 0.0026755 | -0.045694 | 0.0019925 | -0.084687 | 0.0026755 | 1 | 1.8411e-06 | 9.9971e-05 | 0.00043125 | OK | 1.0001 | 4.2684e-06 | 5.0453e-05 | 0.00020434 | OK | 1.0001 | 4.2684e-06 | 5.2524e-05 | 0.00021383 | 0.096242 | 1.4557e-06 | -2.0656 | 0.00021498 | -0.0457 | 1.8222e-05 | -0.0847 | 9.9077e-06
100 | m | 5000 | -46.8031 | 0.0015387 | -54.9046 | 0.00070038 | m | -0.0019809 | 0.00096247 | -0.0019554 | 0.00022876 | m | -46.8001 | 0.02745 | -54.8999 | 0.023441 | m | OK | -46.8 |  | -54.9 |  | m | -0.00014643 | 0.02745 | 0.00013268 | 0.023441 | m | -0.0011004 |  | -0.0026869 |  | m | 5 | kHz | 0.001 | -0.046801 | 1.8149e-06 | -0.054903 | 7.3679e-07 | 0.072143 | 1.3041e-06 | -2.2767 | 2.0259e-05 | 4 | 10m | 0.01 | 126 | 4 | 100m | 144 | 4 | 10m | 100m | 0.01 | 0.1 | 5kHz100m10m | 5kHz100m100m | 126 | 144 | 5kHz100m10m | -4.0201e-07 | 0.0036149 | 9.9975e-07 | 0.00012206 | 5kHz100m100m | -3.7561e-06 | 0.001926 | -6.999e-06 | 0.0021307 | -2.7179e-06 | 0.001926 | -4.5232e-06 | 0.001509 | -0.046804 | 0.001926 | -0.054907 | 0.001509 | 0.99997 | 5.2627e-06 | 9.962e-05 | 0.00051319 | OK | 0.99996 | 5.5463e-06 | -6.6801e-05 | 0.00049029 | OK | 0.99996 | 5.5463e-06 | -1.5291e-05 | 0.00049738 | 0.07214 | 2.6387e-06 | -2.2767 | 0.00049903 | -0.0468 | 2.745e-05 | -0.0549 | 2.3441e-05
10 | m | 200 | 5.6696 | 0.001631 | 5.2082 | 0.0014387 | m | -0.00060284 | 5.3829e-05 | -0.0011192 | 0.00037668 | m | 5.67 | 0.0037611 | 5.21 | 0.0038655 | m | OK | 5.67 |  | 5.21 |  | m | 1.1322e-07 | 0.0037611 | 7.6075e-08 | 0.0038655 | m | 0.00020607 |  | -0.00066932 |  | m | 200 | Hz | 0.001 | 0.0056702 | 1.6319e-06 | 0.0052093 | 1.4872e-06 | 0.0076999 | 1.5673e-06 | 0.74306 | 0.00020196 | 3 | 3m | 0.003 | 86 | 3 | 10m | 104 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -1.5438e-07 | 0.0011539 | 2.9983e-07 | 0.00034655 | 200Hz10m10m | 8.3222e-07 | 0.0040858 | 7.3506e-07 | 0.0027163 | 5.0804e-07 | 0.0040858 | 5.9205e-07 | 0.0019376 | 0.0056707 | 0.0040858 | 0.0052099 | 0.0019376 | 0.99995 | 2.2817e-06 | 9.9996e-05 | 0.00064432 | OK | 1.0001 | 4.9523e-06 | 7.3361e-05 | 0.00028354 | OK | 1 | 4.9523e-06 | 8.2113e-05 | 0.00040209 | 0.0077002 | 3.1349e-06 | 0.74314 | 0.00056994 | 0.00567 | 3.7611e-06 | 0.00521 | 3.8655e-06
10 | m | 5000 | -2.1011 | 0.0015003 | -4.9006 | 0.001185 | m | -0.0011398 | 0.0018471 | -0.0011013 | 0.00080387 | m | -2.1 | 0.0067516 | -4.9 | 0.0041108 | m | OK | -2.1 |  | -4.9 |  | m | -1.5305e-05 | 0.0067516 | 6.9171e-06 | 0.0041108 | m | 2.9825e-05 |  | 0.00053675 |  | m | 5 | kHz | 0.001 | -0.0021 | 2.3796e-06 | -0.0048995 | 1.4319e-06 | 0.0053305 | 1.6159e-06 | -1.9757 | 0.00042374 | 3 | 3m | 0.003 | 90 | 3 | 10m | 108 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | 2.8001e-07 | 0.00028801 | 3.1399e-07 | 0.003471 | 5kHz10m10m | 9.8702e-07 | 0.0018932 | -1.0028e-06 | 0.0009333 | 5.154e-07 | 0.0018932 | -1.2443e-07 | 0.0026261 | -0.0020995 | 0.0018932 | -0.0048996 | 0.0026261 | 1.0001 | 1.0917e-06 | 9.822e-05 | 0.0012855 | OK | 1.0001 | 2.7282e-06 | -0.00010256 | 0.00059291 | OK | 1.0001 | 2.7282e-06 | 3.1372e-05 | 0.0010549 | 0.005331 | 3.2318e-06 | -1.9757 | 0.0013532 | -0.0021 | 6.7516e-06 | -0.0049 | 4.1108e-06
1 | m | 200 | -0.054848 | 5.2195e-05 | -0.13808 | 0.0011676 | m | -0.00014658 | 0.0017484 | 0.0009252 | 0.0017629 | m | -0.054695 | 0.0037204 | -0.139 | 0.0040765 | m | OK | -0.0547 |  | -0.139 |  | m | 5.2578e-06 | 0.0037204 | -2.069e-06 | 0.0040765 | m | -1.3944e-06 |  | -8.3378e-07 |  | m | 200 | Hz | 0.001 | -5.4701e-05 | 1.7491e-06 | -0.000139 | 2.1145e-06 | 0.00014938 | 2.0692e-06 | -1.9457 | 0.012066 | 1 | 0m | 0 | 14 | 1 | 1m | 32 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 0.00093895 | 0 | 0.0021931 | 200Hz1m1m | -5.7742e-08 | 0.0028813 | 4.4363e-08 | 0.0018199 | -8.6253e-09 | 0.0028813 | 4.4363e-08 | 0.0021373 | -5.471e-05 | 0.0028813 | -0.00013896 | 0.0021373 | 1 | 2.3856e-06 | 4.4468e-05 | 0.0038907 | OK | 0.99994 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | OK | 0.99999 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | 0.00014938 | 4.1384e-06 | -1.9457 | 0.024444 | -5.4695e-05 | 3.7204e-06 | -0.000139 | 4.0765e-06
100 | m | 2000 | -8.191 | 0.0014632 | 5.1014 | 0.0014765 | m | -0.0015145 | 0.00022533 | 0.00054374 | 0.00014155 | m | -8.19 | 0.0036931 | 5.1 | 0.0046166 | m | Extrapolated | -8.19 |  | 5.1 |  | m | -1.9241e-07 | 0.0036931 | 4.3486e-07 | 0.0046166 | m | 0.00050293 |  | 0.00082559 |  | m | 2 | kHz | 0.001 | -0.0081895 | 1.4805e-06 | 0.0051008 | 1.4833e-06 | 0.0096481 | 1.4812e-06 | 2.5845 | 0.00015365 | 3 | 3m | 0.003 | 0 | 4 | 100m | 143 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | -2.5323e-08 | 0.00057942 | 1.0047e-06 | 0.00010185 | 2kHz100m100m | 7.1609e-06 | 0.0019925 | 5.2033e-06 | 0.0027879 | -5.3419e-08 | 0.0019925 | 9.8824e-07 | 9.1348e-05 | -0.0081896 | 0.0019925 | 0.0051018 | 9.1396e-05 | 1 | 1.8411e-06 | 9.9971e-05 | 0.00043125 | OK | 1.0001 | 4.2684e-06 | 5.0453e-05 | 0.00020434 | OK | 1 | 4.2684e-06 | 0.00010016 | 0.00043214 | 0.0096481 | 2.9628e-06 | 2.5846 | 0.00053027 | -0.00819 | 3.6931e-06 | 0.0051 | 4.6166e-06
1 | m | 500 | -0.44905 | 0.001446 | -0.56907 | 0.00053236 | m | -5.2256e-05 | 0.00023819 | 0.0019943 | 0.00056548 | m | -0.44901 | 0.0032055 | -0.57099 | 0.0028592 | m | OK | -0.449 |  | -0.571 |  | m | -1.3367e-05 | 0.0032055 | 1.0556e-05 | 0.0028592 | m | 5.0945e-06 |  | -6.6501e-05 |  | m | 500 | Hz | 0.001 | -0.00044899 | 1.4654e-06 | -0.00057107 | 7.7664e-07 | 0.00072644 | 1.0923e-06 | -2.2371 | 0.001718 | 1 | 0m | 0 | 15 | 1 | 1m | 33 | 1 | 0m | 1m | 0 | 0.001 | 500Hz1m0m | 500Hz1m1m | 15 | 33 | 500Hz1m0m | 0 | 0.00075646 | 0 | 0.0021488 | 500Hz1m1m | -9.2849e-08 | 0.00052854 | -8.6039e-08 | 0.00070921 | -6.7449e-08 | 0.00052854 | -8.6039e-08 | 0.001103 | -0.00044906 | 0.00052855 | -0.00057115 | 0.001103 | 1 | 2.278e-06 | -8.5547e-05 | 0.0037572 | OK | 0.99991 | 2.5207e-06 | -8.5547e-05 | 0.0037572 | OK | 0.99993 | 2.5207e-06 | -8.5547e-05 | 0.0037572 | 0.00072639 | 2.1846e-06 | -2.2372 | 0.0050914 | -0.00044901 | 3.2055e-06 | -0.00057099 | 2.8592e-06
3 | m | 200 | -2.6201 | 0.001086 | -0.82899 | 0.0012903 | m | 5.1119e-05 | 0.0017167 | 0.0011282 | 0.0015878 | m | -2.62 | 0.0041924 | -0.83 | 0.0051937 | m | OK | -2.62 |  | -0.83 |  | m | 9.4353e-08 | 0.0041924 | 2.2959e-07 | 0.0051937 | m | -0.00016939 |  | -0.00011792 |  | m | 200 | Hz | 0.001 | -0.0026202 | 2.0314e-06 | -0.00083012 | 2.046e-06 | 0.0027485 | 2.0327e-06 | -2.8348 | 0.0007439 | 2 | 1m | 0.001 | 50 | 2 | 3m | 68 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | 3.3235e-08 | 0.0005621 | 1.0006e-07 | 0.0013627 | 200Hz3m3m | -2.6057e-07 | 0.00043955 | -1.204e-07 | 0.0027986 | -2.2363e-07 | 0.00043955 | -9.2682e-08 | 0.002618 | -0.0026204 | 0.00043957 | -0.00083021 | 0.002618 | 1 | 2.5752e-06 | 9.9997e-05 | 0.0015283 | OK | 0.99991 | 6.2866e-07 | -3.9954e-05 | 0.0011794 | OK | 0.99993 | 6.2866e-07 | -2.2357e-05 | 0.0012233 | 0.0027483 | 4.0654e-06 | -2.8348 | 0.0019261 | -0.00262 | 4.1924e-06 | -0.00083 | 5.1937e-06
10 | m | 0.02 | -4.4319 | 0.00082232 | -1.368 | 0.0013716 | m | -0.0018997 | 0.0017022 | 0.0016865 | 0.0013059 | m | -4.43 | 0.0039196 | -1.37 | 0.0050423 | m | OK | -4.43 |  | -1.37 |  | m | -1.7747e-09 | 0.0039196 | 1.0368e-08 | 0.0050423 | m | -4.8111e-05 |  | 0.00030687 |  | m | 20 | mHz | 0.001 | -0.00443 | 1.8904e-06 | -0.0013697 | 1.8939e-06 | 0.004637 | 1.8907e-06 | -2.8417 | 0.00040837 | 3 | 3m | 0.003 | 74 | 3 | 10m | 92 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | 7.0524e-08 | 0.0022397 | 2.9998e-07 | 0.002096 | 20mHz10m10m | -3.5801e-07 | 0.0034885 | -4.4193e-07 | 0.0018726 | -2.969e-08 | 0.0034885 | 1.2649e-07 | 0.0020437 | -0.0044301 | 0.0034885 | -0.0013696 | 0.0020437 | 1 | 4.262e-06 | 0.00010001 | 0.00092181 | OK | 0.99996 | 3.7418e-06 | -4.418e-05 | 0.00019379 | OK | 1 | 3.7418e-06 | 6.6288e-05 | 0.00075156 | 0.004637 | 3.7815e-06 | -2.8417 | 0.0011099 | -0.00443 | 3.9196e-06 | -0.00137 | 5.0423e-06
1 | m | 500 | -0.86755 | 0.00054372 | 0.0033558 | 0.0010232 | m | 0.0015234 | 0.0012021 | 0.00070015 | 0.00014603 | m | -0.869 | 0.0026387 | 0.0027298 | 0.0038642 | m | OK | -0.869 |  | 0.00272 |  | m | 8.2156e-08 | 0.0026387 | 9.781e-06 | 0.0038642 | m | -7.0513e-05 |  | -6.4345e-05 |  | m | 500 | Hz | 0.001 | -0.00086907 | 1.3193e-06 | 2.6557e-06 | 1.0336e-06 | 0.00086907 | 1.3193e-06 | 3.1385 | 0.0011893 | 1 | 0m | 0 | 15 | 1 | 1m | 33 | 1 | 0m | 1m | 0 | 0.001 | 500Hz1m0m | 500Hz1m1m | 15 | 33 | 500Hz1m0m | 0 | 0.00075646 | 0 | 0.0021488 | 500Hz1m1m | -9.2849e-08 | 0.00052854 | -8.6039e-08 | 0.00070921 | -8.0693e-08 | 0.00052854 | -8.6039e-08 | 0.00089768 | -0.00086915 | 0.00052855 | 2.5696e-06 | 0.00089768 | 1 | 2.278e-06 | -8.5547e-05 | 0.0037572 | OK | 0.99991 | 2.5207e-06 | -8.5547e-05 | 0.0037572 | OK | 0.99992 | 2.5207e-06 | -8.5547e-05 | 0.0037572 | 0.000869 | 2.6386e-06 | 3.1385 | 0.0044468 | -0.000869 | 2.6387e-06 | 2.7298e-06 | 3.8642e-06
10 | m | 0.05 | 9.1315 | 0.0019107 | 3.7992 | 0.0020187 | m | 0.001685 | 0.0010896 | -0.00094657 | 0.00027002 | m | 9.13 | 0.0043631 | 3.8 | 0.0044123 | m | OK | 9.13 |  | 3.8 |  | m | 1.8087e-08 | 0.0043631 | -1.2414e-07 | 0.0044123 | m | -0.00016396 |  | 0.00013696 |  | m | 50 | mHz | 0.001 | 0.0091298 | 2.1996e-06 | 0.0038001 | 2.0367e-06 | 0.0098891 | 2.1763e-06 | 0.39442 | 0.00020847 | 3 | 3m | 0.003 | 75 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 9.5305e-08 | 0.002322 | -2.0281e-07 | 0.00051156 | 0.0091299 | 0.002322 | 0.0037999 | 0.00051157 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK | 1 | 3.906e-06 | -1.917e-05 | 0.00016173 | 0.0098892 | 4.3528e-06 | 0.3944 | 0.0004472 | 0.00913 | 4.3631e-06 | 0.0038 | 4.4123e-06
3 | m | 1000 | 2.1591 | 0.0012028 | 1.501 | 0.00031625 | m | -0.00089962 | 0.0015752 | 0.0010158 | 0.00044886 | m | 2.16 | 0.0040432 | 1.5 | 0.0047008 | m | OK | 2.16 |  | 1.5 |  | m | -8.7162e-08 | 0.0040432 | -1.0524e-06 | 0.0047008 | m | 1.1275e-05 |  | 2.119e-05 |  | m | 1 | kHz | 0.001 | 0.00216 | 1.9819e-06 | 0.0015 | 5.4908e-07 | 0.0026298 | 1.6577e-06 | 0.60699 | 0.00046283 | 2 | 1m | 0.001 | 52 | 2 | 3m | 70 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 7.748e-08 | 0.0015949 | 1.0078e-07 | 0.00016428 | 1kHz3m3m | -8.3541e-08 | 0.002271 | -8.5512e-08 | 0.0038778 | -5.3734e-08 | 0.002271 | -5.1026e-08 | 0.0031904 | 0.00216 | 0.002271 | 0.0015 | 0.0031904 | 1.0001 | 3.804e-06 | 9.9887e-05 | 0.0038076 | OK | 0.99997 | 2.2768e-06 | -2.8192e-05 | 0.0013002 | OK | 0.99999 | 2.2768e-06 | -4.4826e-06 | 0.0017643 | 0.0026298 | 3.3155e-06 | 0.60699 | 0.0019924 | 0.00216 | 4.0432e-06 | 0.0015 | 4.7008e-06
100 | m | 0.02 | -4.1418 | 1.2546e-05 | -16.7023 | 0.00068954 | m | 9.8208e-05 | 0.00099308 | -0.0015281 | 0.00051391 | m | -4.14 | 0.0070502 | -16.7 | 0.002427 | m | OK | -4.14 |  | -16.7 |  | m | -1.1429e-08 | 0.0070502 | -2.3367e-08 | 0.002427 | m | -0.0018966 |  | -0.00072546 |  | m | 20 | mHz | 0.001 | -0.0041419 | 9.9316e-07 | -0.016701 | 8.5998e-07 | 0.017207 | 8.6826e-07 | -1.8139 | 5.73e-05 | 4 | 10m | 0.01 | 110 | 4 | 100m | 128 | 4 | 10m | 100m | 0.01 | 0.1 | 20mHz100m10m | 20mHz100m100m | 110 | 128 | 20mHz100m10m | -7.5251e-07 | 0.0030346 | 1.0001e-06 | 0.0035718 | 20mHz100m100m | 2.19e-06 | 0.004003 | 6.056e-06 | 0.0038788 | -5.1689e-07 | 0.004003 | 1.405e-06 | 0.0035964 | -0.0041424 | 0.004003 | -0.016699 | 0.0035964 | 0.99992 | 3.0673e-06 | 9.9999e-05 | 0.00043485 | OK | 1 | 4.3997e-06 | 6.0567e-05 | 6.8822e-05 | OK | 0.99993 | 4.3997e-06 | 9.6842e-05 | 0.00040554 | 0.017206 | 1.7382e-06 | -1.8138 | 0.00042142 | -0.00414 | 7.0502e-06 | -0.0167 | 2.427e-06
3 | m | 0.5 | -2.0204 | 0.00012336 | -2.2023 | 0.00031725 | m | -0.00038989 | 0.0012263 | -0.0019834 | 0.0013862 | m | -2.02 | 0.0030754 | -2.2 | 0.0030164 | m | OK | -2.02 |  | -2.2 |  | m | -2.6044e-08 | 0.0030754 | 2.4314e-08 | 0.0030164 | m | -3.4601e-05 |  | -0.00027498 |  | m | 500 | mHz | 0.001 | -0.00202 | 1.2325e-06 | -0.0022003 | 1.4221e-06 | 0.0029869 | 1.3387e-06 | -2.3135 | 0.00044279 | 2 | 1m | 0.001 | 42 | 2 | 3m | 60 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | -7.4389e-08 | 0.00055263 | 9.9998e-08 | 0.00059317 | 500mHz3m3m | -2.2698e-07 | 0.0017093 | -1.643e-07 | 0.00091673 | -2.2598e-07 | 0.0017093 | -1.6257e-07 | 0.00091461 | -0.0020203 | 0.0017093 | -0.0022004 | 0.00091462 | 0.99993 | 7.8698e-07 | 9.9998e-05 | 0.0038777 | OK | 0.99992 | 1.8266e-06 | -5.4754e-05 | 0.00068081 | OK | 0.99992 | 1.8266e-06 | -5.3742e-05 | 0.0007017 | 0.0029867 | 2.6774e-06 | -2.3136 | 0.0011299 | -0.00202 | 3.0754e-06 | -0.0022 | 3.0164e-06
100 | m | 200 | -17.6998 | -0.00010287 | -59.7026 | 0.00053037 | m | 0.0018905 | 0.0015157 | -0.0011583 | 0.0011486 | m | -17.7 | 0.008976 | -59.7 | 0.0036313 | m | OK | -17.7 |  | -59.7 |  | m | -7.4419e-07 | 0.008976 | -1.7017e-07 | 0.0036313 | m | -0.001654 |  | -0.0014842 |  | m | 200 | Hz | 0.001 | -0.017702 | 1.5192e-06 | -0.059701 | 1.2652e-06 | 0.062271 | 1.2875e-06 | -1.859 | 2.4092e-05 | 4 | 10m | 0.01 | 122 | 4 | 100m | 140 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | -6.8486e-07 | 0.0012358 | 9.989e-07 | 0.0023593 | 200Hz100m100m | -2.9553e-07 | 0.0039756 | -4.0027e-06 | 0.0032774 | -4.5874e-07 | 0.0039756 | -1.906e-06 | 0.0028925 | -0.017702 | 0.0039756 | -0.059703 | 0.0028925 | 0.99993 | 2.465e-06 | 9.9999e-05 | 0.00025373 | OK | 1 | 3.9771e-06 | -4.0017e-05 | 6.1157e-05 | OK | 0.99997 | 3.9771e-06 | 1.868e-05 | 0.00014189 | 0.062269 | 2.587e-06 | -1.859 | 0.00014985 | -0.0177 | 8.976e-06 | -0.0597 | 3.6313e-06

## Cal Data
Calibration data: measured values of reference standards.
[5 empty] | Measured calibration Rs-Xs |  |  |  |  | Measured residual SHORT Rs-Xs | [5 empty] | Processing stuff | [12 empty] | Reference Z (auto search from Ref Z sheet) | [9 empty] | Absolute correction |  |  |  |  | Aux stuff to generate zero Z gain correction |  |  | Gain-phase correction
Range |  | f | Nominal Z |  | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Status | Range Z | f |  | Nom Z | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | U(Zr) | phi | U(phir) | ref tag | ref ID | Rs | U(Rs) | Xs | U(Xs) | Zr | U(Zr) | phir | U(phir) | tag | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | tag f-rng | tag f-rng' | zero id' | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi))
Ohm |  | Hz | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | - |  | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | Ohm | Ohm | Ohm | Ohm | - | - | - | - | - | rad | rad
1 | m | 0.01 | 0 | m |  | 0.0018728 |  | 0.00068112 | m |  |  |  |  | m | OK | 0.001 | 10 | mHz | 0 | 0.001 | 0 | 1.8728e-06 | 0 | 6.8112e-07 | 0 | 1.9928e-06 | #DIV/0! | 0 | 10mHz0m | 1 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10mHz1m0m | 0 | 0.0037455 | 0 | 0.0013622 | 10mHz1m |  | 19 | 1 | 3.9856e-06 | -7.4603e-05 | 0.0024081 | OK
1 | m | 0.02 | 0 | m |  | 0.0001063 |  | 0.0017413 | m |  |  |  |  | m | OK | 0.001 | 20 | mHz | 0 | 0.001 | 0 | 1.063e-07 | 0 | 1.7413e-06 | 0 | 1.7446e-06 | #DIV/0! | 0 | 20mHz0m | 2 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20mHz1m0m | 0 | 0.0002126 | 0 | 0.0034827 | 20mHz1m |  | 20 | 1 | 3.4891e-06 | 1.4477e-05 | 0.0043398 | OK
1 | m | 0.05 | 0 | m |  | 0.0018359 |  | 0.0014677 | m |  |  |  |  | m | OK | 0.001 | 50 | mHz | 0 | 0.001 | 0 | 1.8359e-06 | 0 | 1.4677e-06 | 0 | 2.3505e-06 | #DIV/0! | 0 | 50mHz0m | 3 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50mHz1m0m | 0 | 0.0036718 | 0 | 0.0029355 | 50mHz1m |  | 21 | 1 | 4.7009e-06 | -9.075e-05 | 0.0014316 | OK
1 | m | 0.1 | 0 | m |  | 0.00064717 |  | 0.0013166 | m |  |  |  |  | m | OK | 0.001 | 100 | mHz | 0 | 0.001 | 0 | 6.4717e-07 | 0 | 1.3166e-06 | 0 | 1.467e-06 | #DIV/0! | 0 | 100mHz0m | 4 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100mHz1m0m | 0 | 0.0012943 | 0 | 0.0026332 | 100mHz1m |  | 22 | 1 | 2.9341e-06 | 5.4857e-05 | 0.0043058 | OK
1 | m | 0.2 | 0 | m |  | 0.00086635 |  | 0.0007648 | m |  |  |  |  | m | OK | 0.001 | 200 | mHz | 0 | 0.001 | 0 | 8.6635e-07 | 0 | 7.648e-07 | 0 | 1.1556e-06 | #DIV/0! | 0 | 200mHz0m | 5 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200mHz1m0m | 0 | 0.0017327 | 0 | 0.0015296 | 200mHz1m |  | 23 | 1 | 2.3113e-06 | -6.5952e-05 | 0.001886 | OK
1 | m | 0.5 | 0 | m |  | 0.00065295 |  | 0.00068229 | m |  |  |  |  | m | OK | 0.001 | 500 | mHz | 0 | 0.001 | 0 | 6.5295e-07 | 0 | 6.8229e-07 | 0 | 9.4438e-07 | #DIV/0! | 0 | 500mHz0m | 6 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500mHz1m0m | 0 | 0.0013059 | 0 | 0.0013646 | 500mHz1m |  | 24 | 1 | 1.8888e-06 | -8.99e-05 | 0.0018112 | OK
1 | m | 1 | 0 | m |  | 0.0019244 |  | 0.00014513 | m |  |  |  |  | m | OK | 0.001 | 1 | Hz | 0 | 0.001 | 0 | 1.9244e-06 | 0 | 1.4513e-07 | 0 | 1.9299e-06 | #DIV/0! | 0 | 1Hz0m | 7 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1Hz1m0m | 0 | 0.0038489 | 0 | 0.00029027 | 1Hz1m |  | 25 | 1 | 3.8598e-06 | -9.556e-05 | 0.0036582 | OK
1 | m | 2 | 0 | m |  | 0.00011149 |  | 0.00015725 | m |  |  |  |  | m | OK | 0.001 | 2 | Hz | 0 | 0.001 | 0 | 1.1149e-07 | 0 | 1.5725e-07 | 0 | 1.9276e-07 | #DIV/0! | 0 | 2Hz0m | 8 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2Hz1m0m | 0 | 0.00022298 | 0 | 0.0003145 | 2Hz1m |  | 26 | 1 | 3.8553e-07 | 1.7136e-06 | 0.0015793 | OK
1 | m | 5 | 0 | m |  | 0.0018422 |  | 7.9515e-05 | m |  |  |  |  | m | OK | 0.001 | 5 | Hz | 0 | 0.001 | 0 | 1.8422e-06 | 0 | 7.9515e-08 | 0 | 1.8439e-06 | #DIV/0! | 0 | 5Hz0m | 9 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5Hz1m0m | 0 | 0.0036843 | 0 | 0.00015903 | 5Hz1m |  | 27 | 1 | 3.6878e-06 | 4.1349e-05 | 0.0014488 | OK
1 | m | 10 | 0 | m |  | 9.16e-05 |  | 0.0010504 | m |  |  |  |  | m | OK | 0.001 | 10 | Hz | 0 | 0.001 | 0 | 9.16e-08 | 0 | 1.0504e-06 | 0 | 1.0544e-06 | #DIV/0! | 0 | 10Hz0m | 10 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10Hz1m0m | 0 | 0.0001832 | 0 | 0.0021009 | 10Hz1m |  | 28 | 1 | 2.1088e-06 | -6.4222e-05 | 0.0024875 | OK
1 | m | 20 | 0 | m |  | 0.00023305 |  | 0.00050124 | m |  |  |  |  | m | OK | 0.001 | 20 | Hz | 0 | 0.001 | 0 | 2.3305e-07 | 0 | 5.0124e-07 | 0 | 5.5277e-07 | #DIV/0! | 0 | 20Hz0m | 11 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20Hz1m0m | 0 | 0.00046609 | 0 | 0.0010025 | 20Hz1m |  | 29 | 1 | 1.1055e-06 | -7.7312e-05 | 0.0031213 | OK
1 | m | 50 | 0 | m |  | 0.00049164 |  | 0.0010278 | m |  |  |  |  | m | OK | 0.001 | 50 | Hz | 0 | 0.001 | 0 | 4.9164e-07 | 0 | 1.0278e-06 | 0 | 1.1393e-06 | #DIV/0! | 0 | 50Hz0m | 12 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50Hz1m0m | 0 | 0.00098328 | 0 | 0.0020555 | 50Hz1m |  | 30 | 1 | 2.2786e-06 | -4.5663e-05 | 0.0015461 | OK
1 | m | 100 | 0 | m |  | 0.00066734 |  | 0.00016473 | m |  |  |  |  | m | OK | 0.001 | 100 | Hz | 0 | 0.001 | 0 | 6.6734e-07 | 0 | 1.6473e-07 | 0 | 6.8737e-07 | #DIV/0! | 0 | 100Hz0m | 13 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100Hz1m0m | 0 | 0.0013347 | 0 | 0.00032947 | 100Hz1m |  | 31 | 1 | 1.3747e-06 | 2.9182e-05 | 0.0030423 | OK
1 | m | 200 | 0 | m |  | 0.00046948 |  | 0.0010965 | m |  |  |  |  | m | OK | 0.001 | 200 | Hz | 0 | 0.001 | 0 | 4.6948e-07 | 0 | 1.0965e-06 | 0 | 1.1928e-06 | #DIV/0! | 0 | 200Hz0m | 14 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200Hz1m0m | 0 | 0.00093895 | 0 | 0.0021931 | 200Hz1m |  | 32 | 1 | 2.3856e-06 | 4.4468e-05 | 0.0038907 | OK
1 | m | 500 | 0 | m |  | 0.00037823 |  | 0.0010744 | m |  |  |  |  | m | OK | 0.001 | 500 | Hz | 0 | 0.001 | 0 | 3.7823e-07 | 0 | 1.0744e-06 | 0 | 1.139e-06 | #DIV/0! | 0 | 500Hz0m | 15 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500Hz1m0m | 0 | 0.00075646 | 0 | 0.0021488 | 500Hz1m |  | 33 | 1 | 2.278e-06 | -8.5547e-05 | 0.0037572 | OK
1 | m | 1000 | 0 | m |  | 0.0019624 |  | 0.00087233 | m |  |  |  |  | m | OK | 0.001 | 1 | kHz | 0 | 0.001 | 0 | 1.9624e-06 | 0 | 8.7233e-07 | 0 | 2.1476e-06 | #DIV/0! | 0 | 1kHz0m | 16 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1kHz1m0m | 0 | 0.0039248 | 0 | 0.0017447 | 1kHz1m |  | 34 | 1 | 4.2951e-06 | 1.8528e-05 | 0.003196 | OK
1 | m | 2000 | 0 | m |  | 0.00082659 |  | 0.00035666 | m |  |  |  |  | m | OK | 0.001 | 2 | kHz | 0 | 0.001 | 0 | 8.2659e-07 | 0 | 3.5666e-07 | 0 | 9.0025e-07 | #DIV/0! | 0 | 2kHz0m | 17 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2kHz1m0m | 0 | 0.0016532 | 0 | 0.00071332 | 2kHz1m |  | 35 | 1 | 1.8005e-06 | 8.0822e-05 | 0.00085427 | OK
1 | m | 5000 | 0 | m |  | 0.00074993 |  | 0.0017686 | m |  |  |  |  | m | OK | 0.001 | 5 | kHz | 0 | 0.001 | 0 | 7.4993e-07 | 0 | 1.7686e-06 | 0 | 1.9211e-06 | #DIV/0! | 0 | 5kHz0m | 18 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5kHz1m0m | 0 | 0.0014999 | 0 | 0.0035373 | 5kHz1m |  | 36 | 1 | 3.8421e-06 | 0.0001012 | 0.0027993 | OK
1 | m | 0.01 | 1 | m | 1.0016 | 0.0017339 | 0.0019387 | 0.00019407 | m | 0.0015039 | 0.0011662 | 0.0018133 | 0.0011882 | m | OK | 0.001 | 10 | mHz | 0.001 | 0.001 | 0.0010001 | 2.0896e-06 | 1.2539e-07 | 1.204e-06 | 0.0010001 | 2.0896e-06 | 0.00012537 | 0.0012038 | 10mHz1m | 19 | 0.0010001 | 1e-08 | 5.0774e-08 | 5e-08 | 0.0010001 | 1e-08 | 5.077e-05 | 4.9996e-05 | 10mHz1m1m | -6.2031e-08 | 0.0034678 | -7.4616e-08 | 0.00038814 | 10mHz1m | 10mHz1m | 19 | 0.99994 | 4.1792e-06 | -7.4603e-05 | 0.0024081 | OK
1 | m | 0.02 | 1 | m | 1.0009 | 0.00080398 | 0.0014657 | 0.0019425 | m | 0.00078813 | 0.00091158 | 0.0014166 | 0.00096717 | m | OK | 0.001 | 20 | mHz | 0.001 | 0.001 | 0.0010001 | 1.2155e-06 | 4.9069e-08 | 2.17e-06 | 0.0010001 | 1.2155e-06 | 4.9064e-05 | 0.0021698 | 20mHz1m | 20 | 0.0010001 | 1e-08 | 6.3549e-08 | 5e-08 | 0.0010001 | 1e-08 | 6.3541e-05 | 4.9993e-05 | 20mHz1m1m | 2.7351e-08 | 0.001608 | 1.4481e-08 | 0.003885 | 20mHz1m | 20mHz1m | 20 | 1 | 2.4309e-06 | 1.4477e-05 | 0.0043398 | OK
1 | m | 0.05 | 1 | m | 1.0011 | 0.0015922 | -0.0014625 | 0.00054466 | m | 0.001192 | 0.00092779 | -0.0014764 | 0.00046366 | m | OK | 0.001 | 50 | mHz | 0.001 | 0.001 | 0.00099989 | 1.8428e-06 | 1.3843e-08 | 7.1529e-07 | 0.00099989 | 1.8428e-06 | 1.3845e-05 | 0.00071537 | 50mHz1m | 21 | 0.00099991 | 1e-08 | -7.6899e-08 | 5e-08 | 0.00099991 | 1e-08 | -7.6905e-05 | 5.0004e-05 | 50mHz1m1m | 2.5191e-08 | 0.0031843 | -9.0742e-08 | 0.0010893 | 50mHz1m | 50mHz1m | 21 | 1 | 3.6855e-06 | -9.075e-05 | 0.0014316 | OK
1 | m | 0.1 | 1 | m | 1.0008 | 0.0013588 | 0.0015782 | 0.0013687 | m | 0.00065322 | 0.00012823 | 0.0016716 | 0.0016619 | m | OK | 0.001 | 100 | mHz | 0.001 | 0.001 | 0.0010001 | 1.3649e-06 | -9.34e-08 | 2.153e-06 | 0.0010001 | 1.3649e-06 | -9.339e-05 | 0.0021527 | 100mHz1m | 22 | 0.001 | 1e-08 | -3.8534e-08 | 5e-08 | 0.001 | 1e-08 | -3.8533e-05 | 4.9999e-05 | 100mHz1m1m | -8.8918e-08 | 0.0027176 | 5.4866e-08 | 0.0027374 | 100mHz1m | 100mHz1m | 22 | 0.99991 | 2.7297e-06 | 5.4857e-05 | 0.0043058 | OK
1 | m | 0.2 | 1 | m | 0.99852 | 0.0017964 | 2.0269e-05 | 0.00075518 | m | -0.0013729 | 0.00069274 | -2.4998e-05 | 0.00056406 | m | OK | 0.001 | 200 | mHz | 0.001 | 0.001 | 0.00099989 | 1.9253e-06 | 4.5267e-08 | 9.4258e-07 | 0.00099989 | 1.9253e-06 | 4.5272e-05 | 0.00094268 | 200mHz1m | 23 | 0.00099998 | 1e-08 | -2.0679e-08 | 5e-08 | 0.00099998 | 1e-08 | -2.068e-05 | 5.0001e-05 | 200mHz1m1m | 8.5826e-08 | 0.0035927 | -6.5946e-08 | 0.0015104 | 200mHz1m | 200mHz1m | 23 | 1.0001 | 3.8506e-06 | -6.5952e-05 | 0.001886 | OK
1 | m | 0.5 | 1 | m | 0.99933 | 0.0014562 | 0.00074151 | 0.0008396 | m | -0.00070871 | 0.001835 | 0.00052003 | 0.00033857 | m | OK | 0.001 | 500 | mHz | 0.001 | 0.001 | 0.001 | 2.3426e-06 | 2.2148e-07 | 9.0529e-07 | 0.001 | 2.3426e-06 | 0.00022147 | 0.00090526 | 500mHz1m | 24 | 0.001 | 1e-08 | 1.3157e-07 | 5e-08 | 0.001 | 1e-08 | 0.00013157 | 4.9999e-05 | 500mHz1m1m | -1.2908e-08 | 0.0029123 | -8.9905e-08 | 0.0016792 | 500mHz1m | 500mHz1m | 24 | 0.99999 | 4.6852e-06 | -8.99e-05 | 0.0018112 | OK
1 | m | 1 | 1 | m | 1.0011 | 0.0013321 | 0.0018024 | 0.0012392 | m | 0.00091178 | 0.0018108 | 0.0017446 | 0.0013455 | m | OK | 0.001 | 1 | Hz | 0.001 | 0.001 | 0.0010002 | 2.248e-06 | 5.7864e-08 | 1.8293e-06 | 0.0010002 | 2.248e-06 | 5.7853e-05 | 0.0018289 | 1Hz1m | 25 | 0.0010002 | 1e-08 | -3.7714e-08 | 5e-08 | 0.0010002 | 1e-08 | -3.7707e-05 | 4.9991e-05 | 1Hz1m1m | -4.1944e-09 | 0.0026643 | -9.5577e-08 | 0.0024785 | 1Hz1m | 1Hz1m | 25 | 1 | 4.496e-06 | -9.556e-05 | 0.0036582 | OK
1 | m | 2 | 1 | m | 1.0007 | 0.0018703 | 0.0017915 | 0.00041526 | m | 0.0008269 | 0.00028021 | 0.0017106 | 0.00067109 | m | OK | 0.001 | 2 | Hz | 0.001 | 0.001 | 0.00099991 | 1.8912e-06 | 8.0926e-08 | 7.8918e-07 | 0.00099991 | 1.8912e-06 | 8.0934e-05 | 0.00078925 | 2Hz1m | 26 | 0.00099996 | 1e-08 | 8.2644e-08 | 5e-08 | 0.00099996 | 1e-08 | 8.2647e-05 | 5.0002e-05 | 2Hz1m1m | 5.023e-08 | 0.0037407 | 1.7176e-09 | 0.00083052 | 2Hz1m | 2Hz1m | 26 | 1.0001 | 3.7824e-06 | 1.7136e-06 | 0.0015793 | OK
1 | m | 5 | 1 | m | 0.99951 | 0.0004813 | -0.00094537 | 0.00072162 | m | -0.00035907 | 0.0004384 | -0.00085093 | 5.6786e-05 | m | OK | 0.001 | 5 | Hz | 0.001 | 0.001 | 0.00099987 | 6.5104e-07 | -9.4444e-08 | 7.2385e-07 | 0.00099987 | 6.5104e-07 | -9.4455e-05 | 0.00072395 | 5Hz1m | 27 | 0.00099993 | 1e-08 | -5.3103e-08 | 5e-08 | 0.00099993 | 1e-08 | -5.3106e-05 | 5.0003e-05 | 5Hz1m1m | 5.7444e-08 | 0.0009626 | 4.1341e-08 | 0.0014432 | 5Hz1m | 5Hz1m | 27 | 1.0001 | 1.3021e-06 | 4.1349e-05 | 0.0014488 | OK
1 | m | 10 | 1 | m | 1.0017 | 9.3939e-05 | -0.0017337 | 0.0012293 | m | 0.0017857 | 0.00069323 | -0.0017752 | 0.00018687 | m | OK | 0.001 | 10 | Hz | 0.001 | 0.001 | 0.00099992 | 6.9957e-07 | 4.1515e-08 | 1.2434e-06 | 0.00099992 | 6.9957e-07 | 4.1519e-05 | 0.0012435 | 10Hz1m | 28 | 0.00099987 | 1e-08 | -2.27e-08 | 5e-08 | 0.00099987 | 1e-08 | -2.2704e-05 | 5.0007e-05 | 10Hz1m1m | -5.4917e-08 | 0.00018788 | -6.4216e-08 | 0.0024586 | 10Hz1m | 10Hz1m | 28 | 0.99995 | 1.3991e-06 | -6.4222e-05 | 0.0024875 | OK
1 | m | 20 | 1 | m | 0.99907 | 0.00187 | -0.00056219 | 0.0005946 | m | -0.00087984 | 0.0016857 | -0.00083091 | 0.0014426 | m | OK | 0.001 | 20 | Hz | 0.001 | 0.001 | 0.00099995 | 2.5177e-06 | 2.6872e-07 | 1.5604e-06 | 0.00099995 | 2.5177e-06 | 0.00026873 | 0.0015604 | 20Hz1m | 29 | 0.001 | 1e-08 | 1.9142e-07 | 5e-08 | 0.001 | 1e-08 | 0.00019142 | 4.9998e-05 | 20Hz1m1m | 7.918e-08 | 0.0037401 | -7.7293e-08 | 0.0011892 | 20Hz1m | 20Hz1m | 29 | 1.0001 | 5.0353e-06 | -7.7312e-05 | 0.0031213 | OK
1 | m | 50 | 1 | m | 0.99859 | 0.0015497 | -0.0011073 | 0.00051844 | m | -0.0013472 | 0.00053025 | -0.0014555 | 0.00057284 | m | OK | 0.001 | 50 | Hz | 0.001 | 0.001 | 0.00099994 | 1.6379e-06 | 3.4824e-07 | 7.7261e-07 | 0.00099994 | 1.6379e-06 | 0.00034826 | 0.00077265 | 50Hz1m | 30 | 0.001 | 1e-08 | 3.026e-07 | 5e-08 | 0.001 | 1e-08 | 0.0003026 | 5e-05 | 50Hz1m1m | 5.573e-08 | 0.0030993 | -4.5644e-08 | 0.0010369 | 50Hz1m | 50Hz1m | 30 | 1.0001 | 3.2757e-06 | -4.5663e-05 | 0.0015461 | OK
1 | m | 100 | 1 | m | 1.0017 | 0.00015639 | 0.0024745 | 0.00097606 | m | 0.001672 | 0.0017861 | 0.0015301 | 0.0011665 | m | OK | 0.001 | 100 | Hz | 0.001 | 0.001 | 0.001 | 1.7929e-06 | 9.443e-07 | 1.521e-06 | 0.001 | 1.7929e-06 | 0.00094426 | 0.0015209 | 100Hz1m | 31 | 0.001 | 1e-08 | 9.7347e-07 | 5e-08 | 0.001 | 1e-08 | 0.00097344 | 4.9998e-05 | 100Hz1m1m | -1.1112e-08 | 0.00031278 | 2.9172e-08 | 0.0019521 | 100Hz1m | 100Hz1m | 31 | 0.99999 | 3.5858e-06 | 2.9182e-05 | 0.0030423 | OK
1 | m | 200 | 1 | m | 0.99923 | 0.0014407 | 0.001694 | 0.00090996 | m | -0.0008661 | 3.4465e-06 | -0.0001455 | 0.0017195 | m | OK | 0.001 | 200 | Hz | 0.001 | 0.001 | 0.0010001 | 1.4407e-06 | 1.8395e-06 | 1.9454e-06 | 0.0010001 | 1.4407e-06 | 0.0018394 | 0.0019452 | 200Hz1m | 32 | 0.001 | 1e-08 | 1.8839e-06 | 5e-08 | 0.001 | 1e-08 | 0.0018838 | 4.9998e-05 | 200Hz1m1m | -5.7742e-08 | 0.0028813 | 4.4363e-08 | 0.0018199 | 200Hz1m | 200Hz1m | 32 | 0.99994 | 2.8814e-06 | 4.4468e-05 | 0.0038907 | OK
1 | m | 500 | 1 | m | 1.0004 | 0.00026427 | 0.0044275 | 0.0003546 | m | -0.00042909 | 0.0012323 | -0.00018145 | 0.0018463 | m | OK | 0.001 | 500 | Hz | 0.001 | 0.001 | 0.0010008 | 1.2603e-06 | 4.6089e-06 | 1.88e-06 | 0.0010008 | 1.2604e-06 | 0.0046051 | 0.0018784 | 500Hz1m | 33 | 0.0010007 | 1.5811e-08 | 4.5229e-06 | 5e-08 | 0.0010007 | 1.5813e-08 | 0.0045195 | 4.9962e-05 | 500Hz1m1m | -9.2849e-08 | 0.00052854 | -8.6039e-08 | 0.00070921 | 500Hz1m | 500Hz1m | 33 | 0.99991 | 2.5207e-06 | -8.5547e-05 | 0.0037572 | OK
1 | m | 1000 | 1 | m | 1.0001 | 0.0010791 | 0.008583 | 0.0015999 | m | -0.0016753 | 0.00033842 | -0.00075594 | 3.1819e-05 | m | OK | 0.001 | 1 | kHz | 0.001 | 0.001 | 0.0010018 | 1.131e-06 | 9.3389e-06 | 1.6002e-06 | 0.0010018 | 1.131e-06 | 0.0093221 | 0.0015972 | 1kHz1m | 34 | 0.0010017 | 4.4721e-08 | 9.3571e-06 | 1e-07 | 0.0010018 | 4.4729e-08 | 0.0093407 | 9.982e-05 | 1kHz1m1m | -4.3281e-08 | 0.0021583 | 1.8158e-08 | 0.0031997 | 1kHz1m | 1kHz1m | 34 | 0.99996 | 2.262e-06 | 1.8528e-05 | 0.003196 | OK
1 | m | 2000 | 1 | m | 1.0032 | 0.001147 | 0.017548 | 4.4198e-06 | m | -0.0016877 | 0.00064274 | -0.00094136 | 0.00041684 | m | OK | 0.001 | 2 | kHz | 0.001 | 0.001 | 0.0010049 | 1.3148e-06 | 1.849e-05 | 4.1687e-07 | 0.0010051 | 1.3146e-06 | 0.018397 | 0.00041539 | 2kHz1m | 35 | 0.0010049 | 1.2649e-07 | 1.8571e-05 | 2e-07 | 0.0010051 | 1.2652e-07 | 0.018478 | 0.00019896 | 2kHz1m1m | 2.3678e-08 | 0.0022941 | 8.1684e-08 | 8.8418e-06 | 2kHz1m | 2kHz1m | 35 | 1 | 2.6293e-06 | 8.0822e-05 | 0.00085427 | OK
1 | m | 5000 | 1 | m | 1.019 | 0.0005795 | 0.047748 | 0.00065012 | m | -0.00090366 | 0.00022863 | 0.0010186 | 0.0012491 | m | OK | 0.001 | 5 | kHz | 0.001 | 0.001 | 0.0010199 | 6.2297e-07 | 4.6729e-05 | 1.4081e-06 | 0.001021 | 6.2565e-07 | 0.045786 | 0.0013781 | 5kHz1m | 36 | 0.0010199 | 5e-07 | 4.6831e-05 | 5e-07 | 0.0010209 | 5e-07 | 0.045887 | 0.00048975 | 5kHz1m1m | -4.4901e-08 | 0.001159 | 1.0137e-07 | 0.0013002 | 5kHz1m | 5kHz1m | 36 | 0.99996 | 1.2513e-06 | 0.0001012 | 0.0027993 | OK
3 | m | 0.01 | 1 | m | 1.001 | 0.0017439 | -0.00095545 | 0.00035757 | m | 0.00083078 | 0.0011704 | -0.00090621 | 0.0017615 | m | OK | 0.003 | 10 | mHz | 0.001 | 0.001 | 0.0010001 | 2.1002e-06 | -4.9234e-08 | 1.7974e-06 | 0.0010001 | 2.1002e-06 | -4.9228e-05 | 0.0017971 | 10mHz1m | 19 | 0.0010001 | 1e-08 | 5.0774e-08 | 5e-08 | 0.0010001 | 1e-08 | 5.077e-05 | 4.9996e-05 | 10mHz3m1m | -5.5112e-08 | 0.0034877 | 1.0001e-07 | 0.00071514 | 10mHz3m | 10mHz3m | 37 | 0.99994 | 4.2004e-06 | 9.9998e-05 | 0.0035946 | OK
3 | m | 0.02 | 1 | m | 1.0018 | 0.0019594 | 1.6811e-05 | 0.001646 | m | 0.0016867 | 0.0010544 | 5.3274e-05 | 0.0018259 | m | OK | 0.003 | 20 | mHz | 0.001 | 0.001 | 0.0010002 | 2.2251e-06 | -3.6462e-08 | 2.4583e-06 | 0.0010002 | 2.2251e-06 | -3.6456e-05 | 0.0024579 | 20mHz1m | 20 | 0.0010001 | 1e-08 | 6.3549e-08 | 5e-08 | 0.0010001 | 1e-08 | 6.3541e-05 | 4.9993e-05 | 20mHz3m1m | -3.0087e-08 | 0.0039189 | 1.0001e-07 | 0.0032919 | 20mHz3m | 20mHz3m | 38 | 0.99997 | 4.4502e-06 | 9.9998e-05 | 0.004916 | OK
3 | m | 0.05 | 1 | m | 0.99972 | 0.0010196 | 0.0015938 | 0.0016648 | m | -0.00025585 | 0.00059341 | 0.0017707 | 0.0014308 | m | OK | 0.003 | 50 | mHz | 0.001 | 0.001 | 0.00099998 | 1.1797e-06 | -1.7691e-07 | 2.1952e-06 | 0.00099998 | 1.1797e-06 | -0.00017691 | 0.0021953 | 50mHz1m | 21 | 0.00099991 | 1e-08 | -7.6899e-08 | 5e-08 | 0.00099991 | 1e-08 | -7.6905e-05 | 5.0004e-05 | 50mHz3m1m | -6.7316e-08 | 0.0020391 | 1.0001e-07 | 0.0033297 | 50mHz3m | 50mHz3m | 39 | 0.99993 | 2.3594e-06 | 0.00010001 | 0.0043908 | OK
3 | m | 0.1 | 1 | m | 0.99981 | 0.00099125 | 0.0017443 | 0.0015514 | m | -0.00014947 | 8.8975e-06 | 0.0018828 | 0.00078593 | m | OK | 0.003 | 100 | mHz | 0.001 | 0.001 | 0.00099996 | 9.9129e-07 | -1.3853e-07 | 1.7391e-06 | 0.00099996 | 9.9129e-07 | -0.00013853 | 0.0017392 | 100mHz1m | 22 | 0.001 | 1e-08 | -3.8534e-08 | 5e-08 | 0.001 | 1e-08 | -3.8533e-05 | 4.9999e-05 | 100mHz3m1m | 6.0542e-08 | 0.0019825 | 9.9992e-08 | 0.0031027 | 100mHz3m | 100mHz3m | 40 | 1.0001 | 1.9826e-06 | 9.9998e-05 | 0.0034787 | OK
3 | m | 0.2 | 1 | m | 0.99854 | 0.001818 | 0.00097112 | 0.00041476 | m | -0.001507 | 0.0012523 | 0.0010918 | 0.00041641 | m | OK | 0.003 | 200 | mHz | 0.001 | 0.001 | 0.001 | 2.2076e-06 | -1.2069e-07 | 5.8773e-07 | 0.001 | 2.2076e-06 | -0.00012068 | 0.0005877 | 200mHz1m | 23 | 0.00099998 | 1e-08 | -2.0679e-08 | 5e-08 | 0.00099998 | 1e-08 | -2.068e-05 | 5.0001e-05 | 200mHz3m1m | -6.6222e-08 | 0.003636 | 1.0001e-07 | 0.00082951 | 200mHz3m | 200mHz3m | 41 | 0.99993 | 4.4152e-06 | 0.0001 | 0.0011765 | OK
3 | m | 0.5 | 1 | m | 1.0014 | 0.00027631 | -0.00038418 | 0.00029658 | m | 0.0012674 | 0.00028015 | -0.00041575 | 0.0019161 | m | OK | 0.003 | 500 | mHz | 0.001 | 0.001 | 0.0010001 | 3.9349e-07 | 3.1573e-08 | 1.9389e-06 | 0.0010001 | 3.9349e-07 | 3.157e-05 | 0.0019387 | 500mHz1m | 24 | 0.001 | 1e-08 | 1.3157e-07 | 5e-08 | 0.001 | 1e-08 | 0.00013157 | 4.9999e-05 | 500mHz3m1m | -7.4389e-08 | 0.00055263 | 9.9998e-08 | 0.00059317 | 500mHz3m | 500mHz3m | 42 | 0.99993 | 7.8698e-07 | 9.9998e-05 | 0.0038777 | OK
3 | m | 1 | 1 | m | 0.99911 | 0.00017221 | 0.0016985 | 0.0013445 | m | -0.0010907 | 6.0692e-05 | 0.0018362 | 0.0019112 | m | OK | 0.003 | 1 | Hz | 0.001 | 0.001 | 0.0010002 | 1.8259e-07 | -1.3773e-07 | 2.3368e-06 | 0.0010002 | 1.8259e-07 | -0.0001377 | 0.0023363 | 1Hz1m | 25 | 0.0010002 | 1e-08 | -3.7714e-08 | 5e-08 | 0.0010002 | 1e-08 | -3.7707e-05 | 4.9991e-05 | 1Hz3m1m | -1.9849e-08 | 0.00034442 | 1.0002e-07 | 0.0026891 | 1Hz3m | 1Hz3m | 43 | 0.99998 | 3.6518e-07 | 9.9995e-05 | 0.0046729 | OK
3 | m | 2 | 1 | m | 1.0002 | 0.0013491 | -0.00093578 | 0.00072805 | m | 0.00024976 | 0.00010509 | -0.00091842 | 0.0012435 | m | OK | 0.003 | 2 | Hz | 0.001 | 0.001 | 0.00099995 | 1.3532e-06 | -1.7355e-08 | 1.441e-06 | 0.00099995 | 1.3532e-06 | -1.7356e-05 | 0.0014411 | 2Hz1m | 26 | 0.00099996 | 1e-08 | 8.2644e-08 | 5e-08 | 0.00099996 | 1e-08 | 8.2647e-05 | 5.0002e-05 | 2Hz3m1m | 1.1326e-08 | 0.0026983 | 9.9999e-08 | 0.0014561 | 2Hz3m | 2Hz3m | 44 | 1 | 2.7064e-06 | 0.0001 | 0.0028826 | OK
3 | m | 5 | 1 | m | 0.99898 | 0.0012115 | -0.0017711 | 0.0015149 | m | -0.00088299 | 0.001089 | -0.001618 | 0.0015777 | m | OK | 0.003 | 5 | Hz | 0.001 | 0.001 | 0.00099987 | 1.629e-06 | -1.5309e-07 | 2.1873e-06 | 0.00099987 | 1.629e-06 | -0.00015311 | 0.0021876 | 5Hz1m | 27 | 0.00099993 | 1e-08 | -5.3103e-08 | 5e-08 | 0.00099993 | 1e-08 | -5.3106e-05 | 5.0003e-05 | 5Hz3m1m | 6.6242e-08 | 0.0024231 | 9.9989e-08 | 0.0030299 | 5Hz3m | 5Hz3m | 45 | 1.0001 | 3.258e-06 | 0.00010001 | 0.0043754 | OK
3 | m | 10 | 1 | m | 1.0001 | 0.0014877 | 0.0011092 | 0.0017248 | m | 0.00023114 | 0.00016674 | 0.0012319 | 0.00065148 | m | OK | 0.003 | 10 | Hz | 0.001 | 0.001 | 0.00099989 | 1.497e-06 | -1.227e-07 | 1.8437e-06 | 0.00099989 | 1.497e-06 | -0.00012271 | 0.0018439 | 10Hz1m | 28 | 0.00099987 | 1e-08 | -2.27e-08 | 5e-08 | 0.00099987 | 1e-08 | -2.2704e-05 | 5.0007e-05 | 10Hz3m1m | -2.7706e-08 | 0.0029754 | 9.9999e-08 | 0.0034496 | 10Hz3m | 10Hz3m | 46 | 0.99997 | 2.994e-06 | 0.00010001 | 0.0036882 | OK
3 | m | 20 | 1 | m | 1.0011 | 0.0012256 | 0.0017826 | 0.0012877 | m | 0.0010133 | 0.00166 | 0.0016911 | 0.0018277 | m | OK | 0.003 | 20 | Hz | 0.001 | 0.001 | 0.0010001 | 2.0634e-06 | 9.1424e-08 | 2.2358e-06 | 0.0010001 | 2.0634e-06 | 9.1419e-05 | 0.0022356 | 20Hz1m | 29 | 0.001 | 1e-08 | 1.9142e-07 | 5e-08 | 0.001 | 1e-08 | 0.00019142 | 4.9998e-05 | 20Hz3m1m | -2.7333e-08 | 0.0024512 | 1e-07 | 0.0025754 | 20Hz3m | 20Hz3m | 47 | 0.99997 | 4.1268e-06 | 0.0001 | 0.0044716 | OK
3 | m | 50 | 1 | m | 0.99915 | 0.00014824 | 0.001104 | 0.0011706 | m | -0.00083203 | 0.0015309 | 0.00090145 | 0.0012731 | m | OK | 0.003 | 50 | Hz | 0.001 | 0.001 | 0.00099998 | 1.538e-06 | 2.0259e-07 | 1.7295e-06 | 0.00099998 | 1.538e-06 | 0.0002026 | 0.0017295 | 50Hz1m | 30 | 0.001 | 1e-08 | 3.026e-07 | 5e-08 | 0.001 | 1e-08 | 0.0003026 | 5e-05 | 50Hz3m1m | 1.4893e-08 | 0.00029649 | 1e-07 | 0.0023413 | 50Hz3m | 50Hz3m | 48 | 1 | 3.0761e-06 | 0.0001 | 0.0034594 | OK
3 | m | 100 | 1 | m | 1.0004 | 0.0002913 | 0.0004255 | 0.0013499 | m | 0.0003972 | 0.00083021 | -0.00044793 | 0.00010677 | m | OK | 0.003 | 100 | Hz | 0.001 | 0.001 | 0.00099999 | 8.7983e-07 | 8.7343e-07 | 1.3541e-06 | 0.00099999 | 8.7983e-07 | 0.00087344 | 0.0013541 | 100Hz1m | 31 | 0.001 | 1e-08 | 9.7347e-07 | 5e-08 | 0.001 | 1e-08 | 0.00097344 | 4.9998e-05 | 100Hz3m1m | 4.5086e-08 | 0.0005826 | 1.0004e-07 | 0.0026998 | 100Hz3m | 100Hz3m | 49 | 1 | 1.7597e-06 | 0.0001 | 0.0027087 | OK
3 | m | 200 | 1 | m | 0.99977 | 0.00028105 | 0.00090784 | 0.00068136 | m | -0.00023262 | 0.0012566 | -0.00087599 | 0.000345 | m | OK | 0.003 | 200 | Hz | 0.001 | 0.001 | 0.001 | 1.2876e-06 | 1.7838e-06 | 7.6372e-07 | 0.001 | 1.2876e-06 | 0.0017838 | 0.00076372 | 200Hz1m | 32 | 0.001 | 1e-08 | 1.8839e-06 | 5e-08 | 0.001 | 1e-08 | 0.0018838 | 4.9998e-05 | 200Hz3m1m | 3.3235e-08 | 0.0005621 | 1.0006e-07 | 0.0013627 | 200Hz3m | 200Hz3m | 50 | 1 | 2.5752e-06 | 9.9997e-05 | 0.0015283 | OK
3 | m | 500 | 1 | m | 1.0021 | 0.00014005 | 0.0047788 | 0.0015973 | m | 0.0012427 | 6.7287e-05 | 0.00035558 | 0.00019702 | m | OK | 0.003 | 500 | Hz | 0.001 | 0.001 | 0.0010008 | 1.5537e-07 | 4.4232e-06 | 1.6094e-06 | 0.0010008 | 1.5553e-07 | 0.0044196 | 0.0016081 | 500Hz1m | 33 | 0.0010007 | 1.5811e-08 | 4.5229e-06 | 5e-08 | 0.0010007 | 1.5813e-08 | 0.0045195 | 4.9962e-05 | 500Hz3m1m | -8.6726e-08 | 0.00028009 | 9.9642e-08 | 0.0031946 | 500Hz3m | 500Hz3m | 51 | 0.99991 | 3.1107e-07 | 9.995e-05 | 0.0032165 | OK
3 | m | 1000 | 1 | m | 1.0022 | 0.00079747 | 0.009465 | 8.214e-05 | m | 0.00057894 | 0.0017267 | 0.00020867 | 0.0019046 | m | OK | 0.003 | 1 | kHz | 0.001 | 0.001 | 0.0010017 | 1.902e-06 | 9.2563e-06 | 1.9064e-06 | 0.0010017 | 1.902e-06 | 0.0092408 | 0.0019031 | 1kHz1m | 34 | 0.0010017 | 4.4721e-08 | 9.3571e-06 | 1e-07 | 0.0010018 | 4.4729e-08 | 0.0093407 | 9.982e-05 | 1kHz3m1m | 7.748e-08 | 0.0015949 | 1.0078e-07 | 0.00016428 | 1kHz3m | 1kHz3m | 52 | 1.0001 | 3.804e-06 | 9.9887e-05 | 0.0038076 | OK
3 | m | 2000 | 1 | m | 1.0041 | 0.0019772 | 0.017023 | 0.00051651 | m | -0.00085189 | 0.00069022 | -0.0014487 | 0.00076648 | m | OK | 0.003 | 2 | kHz | 0.001 | 0.001 | 0.0010049 | 2.0942e-06 | 1.8471e-05 | 9.2427e-07 | 0.0010051 | 2.0939e-06 | 0.018378 | 0.00092021 | 2kHz1m | 35 | 0.0010049 | 1.2649e-07 | 1.8571e-05 | 2e-07 | 0.0010051 | 1.2652e-07 | 0.018478 | 0.00019896 | 2kHz3m1m | -1.7402e-08 | 0.0039544 | 9.9754e-08 | 0.001033 | 2kHz3m | 2kHz3m | 53 | 0.99998 | 4.1878e-06 | 9.9549e-05 | 0.0018511 | OK
3 | m | 5000 | 1 | m | 1.0193 | 0.0019129 | 0.048095 | 0.00038918 | m | -0.00059733 | 0.0010383 | 0.0013634 | 0.00065893 | m | OK | 0.003 | 5 | kHz | 0.001 | 0.001 | 0.0010199 | 2.1765e-06 | 4.6732e-05 | 7.6528e-07 | 0.001021 | 2.1745e-06 | 0.045787 | 0.0007551 | 5kHz1m | 36 | 0.0010199 | 5e-07 | 4.6831e-05 | 5e-07 | 0.0010209 | 5e-07 | 0.045887 | 0.00048975 | 5kHz3m1m | -4.8442e-08 | 0.0038257 | 9.9272e-08 | 0.00077836 | 5kHz3m | 5kHz3m | 54 | 0.99996 | 4.3489e-06 | 9.9306e-05 | 0.0015876 | OK
3 | m | 0.01 | 3 | m | 2.9998 | 0.00081522 | 0.0014546 | 0.0016304 | m | -0.00017887 | 0.0013061 | 0.0015533 | 0.0013968 | m | OK | 0.003 | 10 | mHz | 0.003 | 0.001 | 0.003 | 1.5396e-06 | -9.8618e-08 | 2.1469e-06 | 0.003 | 1.5396e-06 | -3.2873e-05 | 0.00071564 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 7.3301e-08 | 1.5e-07 | 0.0029999 | 3e-08 | 2.4435e-05 | 5.0002e-05 | 10mHz3m3m | -1.1623e-07 | 0.0016304 | 1.7192e-07 | 0.0032608 | 10mHz3m | 10mHz3m | 37 | 0.99996 | 3.0792e-06 | 5.7307e-05 | 0.0014322 | OK
3 | m | 0.02 | 3 | m | 2.9998 | 0.00040429 | 0.00094912 | 0.00063875 | m | -6.2178e-05 | 0.00083582 | 0.0010994 | 0.00024811 | m | OK | 0.003 | 20 | mHz | 0.003 | 0.001 | 0.0029999 | 9.2846e-07 | -1.5031e-07 | 6.8524e-07 | 0.0029999 | 9.2846e-07 | -5.0105e-05 | 0.00022842 | 20mHz3m | 38 | 0.0029997 | 3e-08 | 4.1886e-08 | 1.5e-07 | 0.0029997 | 3e-08 | 1.3963e-05 | 5.0005e-05 | 20mHz3m3m | -1.7701e-07 | 0.00080857 | 1.9219e-07 | 0.0012775 | 20mHz3m | 20mHz3m | 38 | 0.99994 | 1.8569e-06 | 6.4068e-05 | 0.00045958 | OK
3 | m | 0.05 | 3 | m | 2.9984 | 0.0015498 | -0.0015269 | 0.0010126 | m | -0.0017507 | 6.6358e-05 | -0.0017396 | 0.0011265 | m | OK | 0.003 | 50 | mHz | 0.003 | 0.001 | 0.0030002 | 1.5512e-06 | 2.1274e-07 | 1.5147e-06 | 0.0030002 | 1.5512e-06 | 7.091e-05 | 0.00050489 | 50mHz3m | 39 | 0.0030001 | 3e-08 | 9.8063e-08 | 1.5e-07 | 0.0030001 | 3e-08 | 3.2687e-05 | 4.9999e-05 | 50mHz3m3m | -7.9542e-08 | 0.0030995 | -1.1468e-07 | 0.0020253 | 50mHz3m | 50mHz3m | 39 | 0.99997 | 3.1024e-06 | -3.8223e-05 | 0.001011 | OK
3 | m | 0.1 | 3 | m | 2.9997 | 0.0020026 | 0.0011127 | 1.1612e-05 | m | -0.00047351 | 0.0013369 | 0.00059232 | 0.00047683 | m | OK | 0.003 | 100 | mHz | 0.003 | 0.001 | 0.0030002 | 2.4079e-06 | 5.2042e-07 | 4.7697e-07 | 0.0030002 | 2.4079e-06 | 0.00017346 | 0.00015898 | 100mHz3m | 40 | 0.003 | 3e-08 | 2.9191e-07 | 1.5e-07 | 0.003 | 3e-08 | 9.7303e-05 | 5e-05 | 100mHz3m3m | -1.7924e-07 | 0.0040053 | -2.2851e-07 | 2.3224e-05 | 100mHz3m | 100mHz3m | 40 | 0.99994 | 4.8158e-06 | -7.6159e-05 | 0.00032186 | OK
3 | m | 0.2 | 3 | m | 2.9993 | 0.0020071 | -0.00049634 | 0.0018381 | m | -0.00034465 | 0.0019285 | -0.0007516 | 0.0018601 | m | OK | 0.003 | 200 | mHz | 0.003 | 0.001 | 0.0029996 | 2.7835e-06 | 2.5526e-07 | 2.6151e-06 | 0.0029996 | 2.7835e-06 | 8.5098e-05 | 0.00087182 | 200mHz3m | 41 | 0.0029995 | 3e-08 | 1.2602e-07 | 1.5e-07 | 0.0029995 | 3e-08 | 4.2015e-05 | 5.0009e-05 | 200mHz3m3m | -1.4452e-07 | 0.0040143 | -1.2924e-07 | 0.0036762 | 200mHz3m | 200mHz3m | 41 | 0.99995 | 5.5669e-06 | -4.3083e-05 | 0.0017443 | OK
3 | m | 0.5 | 3 | m | 3.0016 | 0.00085465 | 0.0021875 | 0.00045836 | m | 0.0011937 | 0.00032199 | 0.0017601 | 0.00090962 | m | OK | 0.003 | 500 | mHz | 0.003 | 0.001 | 0.0030004 | 9.1329e-07 | 4.274e-07 | 1.0186e-06 | 0.0030004 | 9.1329e-07 | 0.00014245 | 0.00033948 | 500mHz3m | 42 | 0.0030002 | 3e-08 | 2.6309e-07 | 1.5e-07 | 0.0030002 | 3e-08 | 8.7694e-05 | 4.9997e-05 | 500mHz3m3m | -2.2698e-07 | 0.0017093 | -1.643e-07 | 0.00091673 | 500mHz3m | 500mHz3m | 42 | 0.99992 | 1.8266e-06 | -5.4754e-05 | 0.00068081 | OK
3 | m | 1 | 3 | m | 3.0018 | 0.00023577 | 0.0016481 | 0.000323 | m | 0.0014903 | 0.00030483 | 0.0018558 | 0.00082337 | m | OK | 0.003 | 1 | Hz | 0.003 | 0.001 | 0.0030003 | 3.8537e-07 | -2.0776e-07 | 8.8445e-07 | 0.0030003 | 3.8537e-07 | -6.9247e-05 | 0.00029479 | 1Hz3m | 43 | 0.0030002 | 3e-08 | -7.9262e-08 | 1.5e-07 | 0.0030002 | 3e-08 | -2.6418e-05 | 4.9996e-05 | 1Hz3m3m | -1.1752e-08 | 0.00047154 | 1.285e-07 | 0.000646 | 1Hz3m | 1Hz3m | 43 | 1 | 7.7074e-07 | 4.2828e-05 | 0.0005917 | OK
3 | m | 2 | 3 | m | 3 | 0.0019268 | -0.00097242 | 0.00092887 | m | -2.857e-05 | 0.0019558 | -0.0010998 | 0.00059144 | m | OK | 0.003 | 2 | Hz | 0.003 | 0.001 | 0.003 | 2.7455e-06 | 1.2737e-07 | 1.1012e-06 | 0.003 | 2.7455e-06 | 4.2455e-05 | 0.00036706 | 2Hz3m | 44 | 0.003 | 3e-08 | -8.1116e-08 | 1.5e-07 | 0.003 | 3e-08 | -2.7038e-05 | 4.9999e-05 | 2Hz3m3m | 3.5078e-08 | 0.0038536 | -2.0848e-07 | 0.0018577 | 2Hz3m | 2Hz3m | 44 | 1 | 5.491e-06 | -6.9494e-05 | 0.00073582 | OK
3 | m | 5 | 3 | m | 3.0011 | 0.00052946 | 0.0020797 | 0.0019495 | m | 0.0012204 | 0.00042924 | 0.0013156 | 0.001483 | m | OK | 0.003 | 5 | Hz | 0.003 | 0.001 | 0.0029999 | 6.816e-07 | 7.6408e-07 | 2.4495e-06 | 0.0029999 | 6.816e-07 | 0.0002547 | 0.00081652 | 5Hz3m | 45 | 0.0029998 | 3e-08 | 5.804e-07 | 1.5e-07 | 0.0029998 | 3e-08 | 0.00019348 | 5.0003e-05 | 5Hz3m3m | -1.1715e-07 | 0.0010589 | -1.8369e-07 | 0.003899 | 5Hz3m | 5Hz3m | 45 | 0.99996 | 1.3632e-06 | -6.1223e-05 | 0.0016338 | OK
3 | m | 10 | 3 | m | 3.0004 | 0.00071565 | 0.0007331 | 0.0016203 | m | -0.00011519 | 0.0014784 | 0.00042792 | 0.00042105 | m | OK | 0.003 | 10 | Hz | 0.003 | 0.001 | 0.0030005 | 1.6425e-06 | 3.0518e-07 | 1.6741e-06 | 0.0030005 | 1.6425e-06 | 0.00010171 | 0.00055793 | 10Hz3m | 46 | 0.0030007 | 3e-08 | 1.9934e-07 | 1.5e-07 | 0.0030007 | 3e-08 | 6.6431e-05 | 4.9989e-05 | 10Hz3m3m | 1.9116e-07 | 0.0014313 | -1.0584e-07 | 0.0032405 | 10Hz3m | 10Hz3m | 46 | 1.0001 | 3.285e-06 | -3.5278e-05 | 0.001117 | OK
3 | m | 20 | 3 | m | 3.0012 | 0.00070145 | 0.0020395 | 0.00039121 | m | 0.00039719 | 0.00039213 | 0.0019297 | 0.00027872 | m | OK | 0.003 | 20 | Hz | 0.003 | 0.001 | 0.0030008 | 8.0362e-07 | 1.0974e-07 | 4.8035e-07 | 0.0030008 | 8.0362e-07 | 3.6569e-05 | 0.00016008 | 20Hz3m | 47 | 0.0030005 | 3e-08 | 3.7537e-07 | 1.5e-07 | 0.0030005 | 3e-08 | 0.0001251 | 4.9991e-05 | 20Hz3m3m | -2.232e-07 | 0.0014029 | 2.6563e-07 | 0.00078242 | 20Hz3m | 20Hz3m | 47 | 0.99993 | 1.6072e-06 | 8.8531e-05 | 0.00032403 | OK
3 | m | 50 | 3 | m | 2.9996 | 0.00098894 | 0.0020863 | 0.00088264 | m | -0.00016004 | 0.00069314 | 0.00032678 | 0.0010108 | m | OK | 0.003 | 50 | Hz | 0.003 | 0.001 | 0.0029998 | 1.2077e-06 | 1.7595e-06 | 1.342e-06 | 0.0029998 | 1.2077e-06 | 0.00058655 | 0.00044735 | 50Hz3m | 48 | 0.003 | 3e-08 | 1.7363e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00057877 | 4.9999e-05 | 50Hz3m3m | 2.6126e-07 | 0.0019779 | -2.3182e-08 | 0.0017653 | 50Hz3m | 50Hz3m | 48 | 1.0001 | 2.4153e-06 | -7.7782e-06 | 0.0008961 | OK
3 | m | 100 | 3 | m | 3.0022 | 0.00084788 | 0.0011973 | 0.001708 | m | 0.0019999 | 0.000775 | -0.0014097 | 0.0011999 | m | OK | 0.003 | 100 | Hz | 0.003 | 0.001 | 0.0030002 | 1.1487e-06 | 2.607e-06 | 2.0874e-06 | 0.0030002 | 1.1487e-06 | 0.00086894 | 0.00069573 | 100Hz3m | 49 | 0.0030002 | 3e-08 | 2.848e-06 | 1.5e-07 | 0.0030002 | 3e-08 | 0.00094928 | 4.9997e-05 | 100Hz3m3m | -4.7835e-08 | 0.0016958 | 2.41e-07 | 0.003416 | 100Hz3m | 100Hz3m | 49 | 0.99998 | 2.2974e-06 | 8.0342e-05 | 0.0013924 | OK
3 | m | 200 | 3 | m | 3.0005 | 0.00021978 | 0.007449 | 0.0013993 | m | -0.00089013 | 0.0002247 | 0.001733 | 0.0010813 | m | OK | 0.003 | 200 | Hz | 0.003 | 0.001 | 0.0030014 | 3.1431e-07 | 5.7159e-06 | 1.7684e-06 | 0.0030014 | 3.1433e-07 | 0.0019044 | 0.00058918 | 200Hz3m | 50 | 0.0030011 | 3e-08 | 5.5955e-06 | 1.5e-07 | 0.0030011 | 3.0001e-08 | 0.0018645 | 4.9981e-05 | 200Hz3m3m | -2.6057e-07 | 0.00043955 | -1.204e-07 | 0.0027986 | 200Hz3m | 200Hz3m | 50 | 0.99991 | 6.2866e-07 | -3.9954e-05 | 0.0011794 | OK
3 | m | 500 | 3 | m | 3.0026 | 0.0011875 | 0.01503 | 0.00033549 | m | 0.0010865 | 0.00022962 | 0.0012795 | 0.00049924 | m | OK | 0.003 | 500 | Hz | 0.003 | 0.001 | 0.0030015 | 1.2095e-06 | 1.3751e-05 | 6.0149e-07 | 0.0030015 | 1.2095e-06 | 0.0045812 | 0.0002004 | 500Hz3m | 51 | 0.0030017 | 4.7434e-08 | 1.3645e-05 | 1.5e-07 | 0.0030018 | 4.7439e-08 | 0.0045458 | 4.997e-05 | 500Hz3m3m | 2.2928e-07 | 0.002375 | -1.0533e-07 | 0.00067098 | 500Hz3m | 500Hz3m | 51 | 1.0001 | 2.419e-06 | -3.544e-05 | 0.0004039 | OK
3 | m | 1000 | 3 | m | 3.0036 | 0.0011355 | 0.028924 | 0.0019389 | m | -0.0019373 | 8.0084e-05 | 0.00097324 | 0.00019016 | m | OK | 0.003 | 1 | kHz | 0.003 | 0.001 | 0.0030055 | 1.1383e-06 | 2.7951e-05 | 1.9482e-06 | 0.0030056 | 1.1384e-06 | 0.0092997 | 0.00064816 | 1kHz3m | 52 | 0.0030054 | 1.3416e-07 | 2.7866e-05 | 3e-07 | 0.0030055 | 1.3419e-07 | 0.0092715 | 9.9812e-05 | 1kHz3m3m | -8.3541e-08 | 0.002271 | -8.5512e-08 | 0.0038778 | 1kHz3m | 1kHz3m | 52 | 0.99997 | 2.2768e-06 | -2.8192e-05 | 0.0013002 | OK
3 | m | 2000 | 3 | m | 3.0155 | 0.0015035 | 0.055796 | 0.001612 | m | 0.00057504 | 0.0018284 | -0.00051364 | 0.00096068 | m | OK | 0.003 | 2 | kHz | 0.003 | 0.001 | 0.0030149 | 2.3672e-06 | 5.6309e-05 | 1.8765e-06 | 0.0030154 | 2.367e-06 | 0.018675 | 0.00062237 | 2kHz3m | 53 | 0.0030147 | 3.7947e-07 | 5.6069e-05 | 6e-07 | 0.0030152 | 3.7957e-07 | 0.018597 | 0.00019897 | 2kHz3m3m | -2.1785e-07 | 0.0030069 | -2.3991e-07 | 0.0032239 | 2kHz3m | 2kHz3m | 53 | 0.99993 | 4.734e-06 | -7.8203e-05 | 0.0012605 | OK
3 | m | 5000 | 3 | m | 3.0615 | 0.00085411 | 0.14067 | 0.00041101 | m | 0.0012333 | 8.7391e-05 | 0.00067882 | 0.00043361 | m | OK | 0.003 | 5 | kHz | 0.003 | 0.001 | 0.0030602 | 8.5857e-07 | 0.00014 | 5.9745e-07 | 0.0030634 | 8.5811e-07 | 0.045715 | 0.00019524 | 5kHz3m | 54 | 0.0030601 | 1.5e-06 | 0.00014009 | 1.5e-06 | 0.0030633 | 1.5e-06 | 0.045747 | 0.00048966 | 5kHz3m3m | -8.919e-08 | 0.0017082 | 9.2925e-08 | 0.00082202 | 5kHz3m | 5kHz3m | 54 | 0.99997 | 1.7162e-06 | 3.1633e-05 | 0.0006263 | OK
10 | m | 0.01 | 3 | m | 2.9986 | 0.0018568 | 0.00070275 | 0.00036109 | m | -0.0010538 | 0.00013243 | 0.00092942 | 0.0014195 | m | OK | 0.01 | 10 | mHz | 0.003 | 0.001 | 0.0029996 | 1.8615e-06 | -2.2667e-07 | 1.4647e-06 | 0.0029996 | 1.8615e-06 | -7.5565e-05 | 0.0004883 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 7.3301e-08 | 1.5e-07 | 0.0029999 | 3e-08 | 2.4435e-05 | 5.0002e-05 | 10mHz10m3m | 2.6834e-07 | 0.0037136 | 2.9997e-07 | 0.00072218 | 10mHz10m | 10mHz10m | 73 | 1.0001 | 3.723e-06 | 0.0001 | 0.00097788 | OK
10 | m | 0.02 | 3 | m | 3.0009 | 0.0011198 | 0.0017183 | 0.001048 | m | 0.0012235 | 0.001813 | 0.0019764 | 0.00089863 | m | OK | 0.01 | 20 | mHz | 0.003 | 0.001 | 0.0029996 | 2.131e-06 | -2.581e-07 | 1.3805e-06 | 0.0029996 | 2.131e-06 | -8.6043e-05 | 0.00046022 | 20mHz3m | 38 | 0.0029997 | 3e-08 | 4.1886e-08 | 1.5e-07 | 0.0029997 | 3e-08 | 1.3963e-05 | 5.0005e-05 | 20mHz10m3m | 7.0524e-08 | 0.0022397 | 2.9998e-07 | 0.002096 | 20mHz10m | 20mHz10m | 74 | 1 | 4.262e-06 | 0.00010001 | 0.00092181 | OK
10 | m | 0.05 | 3 | m | 2.9987 | 0.00042255 | 0.00068515 | 0.000617 | m | -0.0012846 | 0.0018742 | 0.00088708 | 7.8727e-05 | m | OK | 0.01 | 50 | mHz | 0.003 | 0.001 | 0.0029999 | 1.9212e-06 | -2.0193e-07 | 6.2201e-07 | 0.0029999 | 1.9212e-06 | -6.7312e-05 | 0.00020734 | 50mHz3m | 39 | 0.0030001 | 3e-08 | 9.8063e-08 | 1.5e-07 | 0.0030001 | 3e-08 | 3.2687e-05 | 4.9999e-05 | 50mHz10m3m | 1.2157e-07 | 0.0008451 | 3e-07 | 0.001234 | 50mHz10m | 50mHz10m | 75 | 1 | 3.8425e-06 | 9.9999e-05 | 0.00041768 | OK
10 | m | 0.1 | 3 | m | 2.9999 | 8.8784e-05 | 0.001813 | 0.00051933 | m | -0.00019936 | 0.00051655 | 0.0018211 | 0.00019366 | m | OK | 0.01 | 100 | mHz | 0.003 | 0.001 | 0.0030001 | 5.2412e-07 | -8.0905e-09 | 5.5426e-07 | 0.0030001 | 5.2412e-07 | -2.6967e-06 | 0.00018475 | 100mHz3m | 40 | 0.003 | 3e-08 | 2.9191e-07 | 1.5e-07 | 0.003 | 3e-08 | 9.7303e-05 | 5e-05 | 100mHz10m3m | -7.9009e-08 | 0.00017757 | 3e-07 | 0.0010387 | 100mHz10m | 100mHz10m | 76 | 0.99997 | 1.0482e-06 | 0.0001 | 0.00037286 | OK
10 | m | 0.2 | 3 | m | 2.9992 | 0.00087989 | 0.0017674 | 0.00076837 | m | -0.00014681 | 0.0017672 | 0.0019413 | 0.00068831 | m | OK | 0.01 | 200 | mHz | 0.003 | 0.001 | 0.0029994 | 1.9741e-06 | -1.7394e-07 | 1.0316e-06 | 0.0029994 | 1.9741e-06 | -5.7991e-05 | 0.00034393 | 200mHz3m | 41 | 0.0029995 | 3e-08 | 1.2602e-07 | 1.5e-07 | 0.0029995 | 3e-08 | 4.2015e-05 | 5.0009e-05 | 200mHz10m3m | 1.0921e-07 | 0.0017598 | 2.9996e-07 | 0.0015367 | 200mHz10m | 200mHz10m | 77 | 1 | 3.9482e-06 | 0.00010001 | 0.00068968 | OK
10 | m | 0.5 | 3 | m | 3.0013 | 0.00116 | 0.0013266 | 0.00083535 | m | 0.0010278 | 0.0011762 | 0.0013635 | 3.5617e-05 | m | OK | 0.01 | 500 | mHz | 0.003 | 0.001 | 0.0030003 | 1.652e-06 | -3.6913e-08 | 8.361e-07 | 0.0030003 | 1.652e-06 | -1.2303e-05 | 0.00027868 | 500mHz3m | 42 | 0.0030002 | 3e-08 | 2.6309e-07 | 1.5e-07 | 0.0030002 | 3e-08 | 8.7694e-05 | 4.9997e-05 | 500mHz10m3m | -1.1635e-07 | 0.00232 | 3.0001e-07 | 0.0016707 | 500mHz10m | 500mHz10m | 78 | 0.99996 | 3.3039e-06 | 9.9997e-05 | 0.00055959 | OK
10 | m | 1 | 3 | m | 2.9986 | 0.001041 | 0.00063684 | 0.0018269 | m | -0.0018928 | 0.0011147 | 0.0010161 | 0.0012408 | m | OK | 0.01 | 1 | Hz | 0.003 | 0.001 | 0.0030004 | 1.5252e-06 | -3.793e-07 | 2.2085e-06 | 0.0030004 | 1.5252e-06 | -0.00012641 | 0.00073604 | 1Hz3m | 43 | 0.0030002 | 3e-08 | -7.9262e-08 | 1.5e-07 | 0.0030002 | 3e-08 | -2.6418e-05 | 4.9996e-05 | 1Hz10m3m | -2.004e-07 | 0.0020819 | 3.0004e-07 | 0.0036539 | 1Hz10m | 1Hz10m | 79 | 0.99993 | 3.0503e-06 | 9.9996e-05 | 0.0014729 | OK
10 | m | 2 | 3 | m | 2.9991 | 0.0017766 | -0.00075634 | 0.0013976 | m | -0.000977 | 0.0016774 | -0.00037522 | 0.00023129 | m | OK | 0.01 | 2 | Hz | 0.003 | 0.001 | 0.0030001 | 2.4433e-06 | -3.8112e-07 | 1.4166e-06 | 0.0030001 | 2.4433e-06 | -0.00012704 | 0.00047219 | 2Hz3m | 44 | 0.003 | 3e-08 | -8.1116e-08 | 1.5e-07 | 0.003 | 3e-08 | -2.7038e-05 | 4.9999e-05 | 2Hz10m3m | -2.4036e-08 | 0.0035532 | 3e-07 | 0.0027952 | 2Hz10m | 2Hz10m | 80 | 0.99999 | 4.8867e-06 | 9.9999e-05 | 0.00094571 | OK
10 | m | 5 | 3 | m | 2.9983 | 0.00047634 | 0.0011224 | 0.0010045 | m | -0.0015925 | 0.0014507 | 0.00084194 | 0.00078586 | m | OK | 0.01 | 5 | Hz | 0.003 | 0.001 | 0.0029999 | 1.5269e-06 | 2.8043e-07 | 1.2754e-06 | 0.0029999 | 1.5269e-06 | 9.3478e-05 | 0.00042514 | 5Hz3m | 45 | 0.0029998 | 3e-08 | 5.804e-07 | 1.5e-07 | 0.0029998 | 3e-08 | 0.00019348 | 5.0003e-05 | 5Hz10m3m | -1.2341e-07 | 0.00095268 | 2.9997e-07 | 0.002009 | 5Hz10m | 5Hz10m | 81 | 0.99996 | 3.0538e-06 | 0.0001 | 0.00085175 | OK
10 | m | 10 | 3 | m | 2.9995 | 0.0017057 | 2.6863e-05 | 0.0014303 | m | -0.0012033 | 0.0012148 | 0.00012756 | 0.0013805 | m | OK | 0.01 | 10 | Hz | 0.003 | 0.001 | 0.0030007 | 2.0941e-06 | -1.007e-07 | 1.9878e-06 | 0.0030007 | 2.0941e-06 | -3.3559e-05 | 0.00066246 | 10Hz3m | 46 | 0.0030007 | 3e-08 | 1.9934e-07 | 1.5e-07 | 0.0030007 | 3e-08 | 6.6431e-05 | 4.9989e-05 | 10Hz10m3m | -7.5776e-09 | 0.0034114 | 3.0004e-07 | 0.0028606 | 10Hz10m | 10Hz10m | 82 | 1 | 4.1881e-06 | 9.999e-05 | 0.0013259 | OK
10 | m | 20 | 3 | m | 2.9991 | 0.0019122 | 0.0012331 | 0.001021 | m | -0.0011674 | 0.0015453 | 0.0011577 | 0.00079398 | m | OK | 0.01 | 20 | Hz | 0.003 | 0.001 | 0.0030002 | 2.4586e-06 | 7.5351e-08 | 1.2934e-06 | 0.0030002 | 2.4586e-06 | 2.5115e-05 | 0.0004311 | 20Hz3m | 47 | 0.0030005 | 3e-08 | 3.7537e-07 | 1.5e-07 | 0.0030005 | 3e-08 | 0.0001251 | 4.9991e-05 | 20Hz10m3m | 2.8614e-07 | 0.0038244 | 3.0002e-07 | 0.002042 | 20Hz10m | 20Hz10m | 83 | 1.0001 | 4.9171e-06 | 9.9985e-05 | 0.00086364 | OK
10 | m | 50 | 3 | m | 3.0006 | 6.2433e-05 | 0.0029959 | 0.00083737 | m | 0.00034747 | 0.0010847 | 0.0015594 | 5.6439e-05 | m | OK | 0.01 | 50 | Hz | 0.003 | 0.001 | 0.0030002 | 1.0865e-06 | 1.4364e-06 | 8.3927e-07 | 0.0030002 | 1.0865e-06 | 0.00047877 | 0.00027974 | 50Hz3m | 48 | 0.003 | 3e-08 | 1.7363e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00057877 | 4.9999e-05 | 50Hz10m3m | -1.6068e-07 | 0.00012487 | 2.9993e-07 | 0.0016747 | 50Hz10m | 50Hz10m | 84 | 0.99995 | 2.1729e-06 | 9.9999e-05 | 0.0005617 | OK
10 | m | 100 | 3 | m | 3.0012 | 0.0014008 | 0.0021574 | 0.0015992 | m | 0.00077748 | 0.0011875 | -0.00039071 | 0.001219 | m | OK | 0.01 | 100 | Hz | 0.003 | 0.001 | 0.0030004 | 1.8364e-06 | 2.5482e-06 | 2.0108e-06 | 0.0030004 | 1.8364e-06 | 0.00084928 | 0.00067018 | 100Hz3m | 49 | 0.0030002 | 3e-08 | 2.848e-06 | 1.5e-07 | 0.0030002 | 3e-08 | 0.00094928 | 4.9997e-05 | 100Hz10m3m | -2.1019e-07 | 0.0028016 | 2.9984e-07 | 0.0031983 | 100Hz10m | 100Hz10m | 85 | 0.99993 | 3.6728e-06 | 9.9999e-05 | 0.0013413 | OK
10 | m | 200 | 3 | m | 3 | 0.00057697 | 0.0053556 | 0.00017328 | m | -0.0012958 | 0.00098422 | 5.9853e-05 | 0.00094828 | m | OK | 0.01 | 200 | Hz | 0.003 | 0.001 | 0.0030013 | 1.1409e-06 | 5.2957e-06 | 9.6398e-07 | 0.0030013 | 1.1409e-06 | 0.0017645 | 0.00032119 | 200Hz3m | 50 | 0.0030011 | 3e-08 | 5.5955e-06 | 1.5e-07 | 0.0030011 | 3.0001e-08 | 0.0018645 | 4.9981e-05 | 200Hz10m3m | -1.5438e-07 | 0.0011539 | 2.9983e-07 | 0.00034655 | 200Hz10m | 200Hz10m | 86 | 0.99995 | 2.2817e-06 | 9.9996e-05 | 0.00064432 | OK
10 | m | 500 | 3 | m | 3.0003 | 0.0014123 | 0.014093 | 0.0017231 | m | -0.001208 | 0.0016005 | 0.00074875 | 0.0016894 | m | OK | 0.01 | 500 | Hz | 0.003 | 0.001 | 0.0030015 | 2.1345e-06 | 1.3344e-05 | 2.4131e-06 | 0.0030016 | 2.1345e-06 | 0.0044458 | 0.00080394 | 500Hz3m | 51 | 0.0030017 | 4.7434e-08 | 1.3645e-05 | 1.5e-07 | 0.0030018 | 4.7439e-08 | 0.0045458 | 4.997e-05 | 500Hz10m3m | 2.1936e-07 | 0.0028245 | 3.0112e-07 | 0.0034461 | 500Hz10m | 500Hz10m | 87 | 1.0001 | 4.2691e-06 | 9.9988e-05 | 0.0016087 | OK
10 | m | 1000 | 3 | m | 3.0053 | 0.00076354 | 0.028911 | 0.0013918 | m | -6.1184e-05 | 0.00078264 | 0.0013458 | 0.00060784 | m | OK | 0.01 | 1 | kHz | 0.003 | 0.001 | 0.0030054 | 1.0934e-06 | 2.7565e-05 | 1.5187e-06 | 0.0030055 | 1.0934e-06 | 0.0091716 | 0.0005053 | 1kHz3m | 52 | 0.0030054 | 1.3416e-07 | 2.7866e-05 | 3e-07 | 0.0030055 | 1.3419e-07 | 0.0092715 | 9.9812e-05 | 1kHz10m3m | 3.7743e-08 | 0.0015271 | 3.0077e-07 | 0.0027835 | 1kHz10m | 1kHz10m | 88 | 1 | 2.1869e-06 | 9.9951e-05 | 0.0010155 | OK
10 | m | 2000 | 3 | m | 3.0144 | 0.0008266 | 0.056474 | 0.00033537 | m | -0.00062128 | 0.0018708 | 0.00069937 | 0.00046402 | m | OK | 0.01 | 2 | kHz | 0.003 | 0.001 | 0.003015 | 2.0452e-06 | 5.5774e-05 | 5.7252e-07 | 0.0030155 | 2.0449e-06 | 0.018497 | 0.00019024 | 2kHz3m | 53 | 0.0030147 | 3.7947e-07 | 5.6069e-05 | 6e-07 | 0.0030152 | 3.7957e-07 | 0.018597 | 0.00019897 | 2kHz10m3m | -2.9449e-07 | 0.0016532 | 2.9509e-07 | 0.00067073 | 2kHz10m | 2kHz10m | 89 | 0.9999 | 4.0898e-06 | 9.9658e-05 | 0.00042937 | OK
10 | m | 5000 | 3 | m | 3.0588 | 0.000144 | 0.13969 | 0.0017355 | m | -0.0010439 | 0.00052051 | -8.6643e-05 | 0.00055476 | m | OK | 0.01 | 5 | kHz | 0.003 | 0.001 | 0.0030599 | 5.4006e-07 | 0.00013978 | 1.822e-06 | 0.0030631 | 5.4587e-07 | 0.045648 | 0.00059428 | 5kHz3m | 54 | 0.0030601 | 1.5e-06 | 0.00014009 | 1.5e-06 | 0.0030633 | 1.5e-06 | 0.045747 | 0.00048966 | 5kHz10m3m | 2.8001e-07 | 0.00028801 | 3.1399e-07 | 0.003471 | 5kHz10m | 5kHz10m | 90 | 1.0001 | 1.0917e-06 | 9.822e-05 | 0.0012855 | OK
10 | m | 0.01 | 10 | m | 10.0004 | 0.0014166 | -0.0029688 | 0.0018979 | m | -0.000735 | 0.0012758 | -0.0019418 | 0.00052235 | m | OK | 0.01 | 10 | mHz | 0.01 | 0.001 | 0.010001 | 1.9064e-06 | -1.027e-06 | 1.9685e-06 | 0.010001 | 1.9064e-06 | -0.00010269 | 0.00019683 | 10mHz10m | 55 | 0.01 | 1e-07 | -5.5374e-08 | 5e-07 | 0.01 | 1e-07 | -5.5372e-06 | 4.9998e-05 | 10mHz10m10m | -6.7217e-07 | 0.0028332 | 9.7161e-07 | 0.0037958 | 10mHz10m | 10mHz10m | 73 | 0.99993 | 3.8128e-06 | 9.715e-05 | 0.00039681 | OK
10 | m | 0.02 | 10 | m | 10.0032 | 0.0017443 | 0.0024477 | 0.00093629 | m | 0.0013668 | 0.00067662 | 0.00067719 | 1.0166e-05 | m | OK | 0.01 | 20 | mHz | 0.01 | 0.001 | 0.010002 | 1.8709e-06 | 1.7705e-06 | 9.3634e-07 | 0.010002 | 1.8709e-06 | 0.00017702 | 9.3617e-05 | 20mHz10m | 56 | 0.010001 | 1e-07 | 1.3286e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00013284 | 4.9993e-05 | 20mHz10m10m | -3.5801e-07 | 0.0034885 | -4.4193e-07 | 0.0018726 | 20mHz10m | 20mHz10m | 74 | 0.99996 | 3.7418e-06 | -4.418e-05 | 0.00019379 | OK
10 | m | 0.05 | 10 | m | 9.9998 | 0.001161 | -0.00067734 | 0.00024997 | m | -0.00077851 | 0.0015704 | 0.00025661 | 0.00070437 | m | OK | 0.01 | 50 | mHz | 0.01 | 0.001 | 0.010001 | 1.953e-06 | -9.3396e-07 | 7.4741e-07 | 0.010001 | 1.953e-06 | -9.339e-05 | 7.4737e-05 | 50mHz10m | 57 | 0.010001 | 1e-07 | -1.1449e-06 | 5e-07 | 0.010001 | 1e-07 | -0.00011448 | 4.9996e-05 | 50mHz10m10m | 9.4882e-08 | 0.002322 | -2.109e-07 | 0.00049993 | 50mHz10m | 50mHz10m | 75 | 1 | 3.906e-06 | -2.1088e-05 | 0.00015761 | OK
10 | m | 0.1 | 10 | m | 9.9989 | 0.0020821 | -0.00066073 | 0.0013845 | m | -0.00053421 | 0.00040589 | -0.00055215 | 7.0143e-05 | m | OK | 0.01 | 100 | mHz | 0.01 | 0.001 | 0.0099994 | 2.1213e-06 | -1.0858e-07 | 1.3863e-06 | 0.0099994 | 2.1213e-06 | -1.0859e-05 | 0.00013863 | 100mHz10m | 58 | 0.01 | 1e-07 | -1.6302e-07 | 5e-07 | 0.01 | 1e-07 | -1.6301e-05 | 4.9998e-05 | 100mHz10m10m | 9.8254e-07 | 0.0041642 | -5.4433e-08 | 0.002769 | 100mHz10m | 100mHz10m | 76 | 1.0001 | 4.2426e-06 | -5.442e-06 | 0.00028174 | OK
10 | m | 0.2 | 10 | m | 9.9978 | 0.00078812 | 0.00021078 | 0.0019457 | m | -0.00091473 | 0.00049459 | -0.00044805 | 0.0017649 | m | OK | 0.01 | 200 | mHz | 0.01 | 0.001 | 0.0099987 | 9.3046e-07 | 6.5883e-07 | 2.627e-06 | 0.0099987 | 9.3046e-07 | 6.5892e-05 | 0.00026273 | 200mHz10m | 59 | 0.0099996 | 1e-07 | 7.4735e-07 | 5e-07 | 0.0099996 | 1e-07 | 7.4739e-05 | 5.0002e-05 | 200mHz10m10m | 8.6912e-07 | 0.0015762 | 8.8521e-08 | 0.0038915 | 200mHz10m | 200mHz10m | 77 | 1.0001 | 1.8609e-06 | 8.8467e-06 | 0.00052783 | OK
10 | m | 0.5 | 10 | m | 9.9984 | 0.00038266 | 0.001148 | 0.001691 | m | -0.0019599 | 0.0014717 | 0.0011358 | 0.0016404 | m | OK | 0.01 | 500 | mHz | 0.01 | 0.001 | 0.01 | 1.5206e-06 | 1.2199e-08 | 2.3559e-06 | 0.01 | 1.5206e-06 | 1.2198e-06 | 0.00023558 | 500mHz10m | 60 | 0.01 | 1e-07 | -4.5986e-07 | 5e-07 | 0.01 | 1e-07 | -4.5985e-05 | 4.9998e-05 | 500mHz10m10m | -3.3994e-08 | 0.00076531 | -4.7206e-07 | 0.0033821 | 500mHz10m | 500mHz10m | 78 | 1 | 3.0413e-06 | -4.7205e-05 | 0.00047381 | OK
10 | m | 1 | 10 | m | 9.9975 | 0.0019249 | 0.0011768 | 0.0017926 | m | -0.0013801 | 0.0008994 | 0.0016533 | 0.0013932 | m | OK | 0.01 | 1 | Hz | 0.01 | 0.001 | 0.0099989 | 2.1246e-06 | -4.7648e-07 | 2.2703e-06 | 0.0099989 | 2.1246e-06 | -4.7653e-05 | 0.00022706 | 1Hz10m | 61 | 0.009999 | 1e-07 | -4.9945e-07 | 5e-07 | 0.009999 | 1e-07 | -4.995e-05 | 5.0005e-05 | 1Hz10m10m | 1.5495e-07 | 0.0038498 | -2.2976e-08 | 0.0035852 | 1Hz10m | 1Hz10m | 79 | 1 | 4.2493e-06 | -2.2971e-06 | 0.00045686 | OK
10 | m | 2 | 10 | m | 9.9971 | 0.00011124 | 0.0030534 | 0.0018588 | m | -0.00019017 | 0.0014492 | 0.0015503 | 1.0929e-05 | m | OK | 0.01 | 2 | Hz | 0.01 | 0.001 | 0.0099973 | 1.4534e-06 | 1.5031e-06 | 1.8588e-06 | 0.0099973 | 1.4534e-06 | 0.00015036 | 0.00018593 | 2Hz10m | 62 | 0.009998 | 1e-07 | 9.4431e-07 | 5e-07 | 0.009998 | 1e-07 | 9.445e-05 | 5.001e-05 | 2Hz10m10m | 7.1523e-07 | 0.00022248 | -5.5883e-07 | 0.0037175 | 2Hz10m | 2Hz10m | 80 | 1.0001 | 2.9068e-06 | -5.5905e-05 | 0.00037521 | OK
10 | m | 5 | 10 | m | 9.9981 | 0.0020326 | -0.0022348 | 0.00056911 | m | -0.0014199 | 0.00017872 | -0.00038394 | 0.0017566 | m | OK | 0.01 | 5 | Hz | 0.01 | 0.001 | 0.0099995 | 2.0405e-06 | -1.8508e-06 | 1.8465e-06 | 0.0099995 | 2.0405e-06 | -0.00018509 | 0.00018466 | 5Hz10m | 63 | 0.01 | 1e-07 | -9.9104e-07 | 5e-07 | 0.01 | 1e-07 | -9.9102e-05 | 4.9999e-05 | 5Hz10m10m | 7.4124e-07 | 0.0040653 | 8.5979e-07 | 0.0011382 | 5Hz10m | 5Hz10m | 81 | 1.0001 | 4.081e-06 | 8.5991e-05 | 0.00037269 | OK
10 | m | 10 | 10 | m | 9.9998 | 0.0019827 | -0.00018872 | 0.0018857 | m | -0.00049685 | 0.00085959 | -0.00029132 | 0.00073404 | m | OK | 0.01 | 10 | Hz | 0.01 | 0.001 | 0.01 | 2.161e-06 | 1.0259e-07 | 2.0235e-06 | 0.01 | 2.161e-06 | 1.0259e-05 | 0.00020235 | 10Hz10m | 64 | 0.010001 | 1e-07 | 8.9632e-08 | 5e-07 | 0.010001 | 1e-07 | 8.9624e-06 | 4.9995e-05 | 10Hz10m10m | 6.9689e-07 | 0.0039654 | -1.2962e-08 | 0.0037714 | 10Hz10m | 10Hz10m | 82 | 1.0001 | 4.322e-06 | -1.2968e-06 | 0.00040777 | OK
10 | m | 20 | 10 | m | 9.9984 | 0.00046783 | 0.0032939 | 0.0014533 | m | -0.0018929 | 0.0016618 | 0.00072524 | 0.0010563 | m | OK | 0.01 | 20 | Hz | 0.01 | 0.001 | 0.01 | 1.7264e-06 | 2.5687e-06 | 1.7967e-06 | 0.01 | 1.7264e-06 | 0.00025686 | 0.00017966 | 20Hz10m | 65 | 0.0099998 | 1e-07 | 1.6297e-06 | 5e-07 | 0.0099998 | 1e-07 | 0.00016297 | 5.0001e-05 | 20Hz10m10m | -5.6007e-07 | 0.00093566 | -9.39e-07 | 0.0029067 | 20Hz10m | 20Hz10m | 83 | 0.99994 | 3.4529e-06 | -9.3888e-05 | 0.00036278 | OK
10 | m | 50 | 10 | m | 9.999 | 0.0018726 | 0.0039251 | 0.0016944 | m | -0.00012033 | 0.0014548 | -0.00055172 | 0.0019197 | m | OK | 0.01 | 50 | Hz | 0.01 | 0.001 | 0.0099991 | 2.3713e-06 | 4.4768e-06 | 2.5605e-06 | 0.0099991 | 2.3713e-06 | 0.00044772 | 0.00025608 | 50Hz10m | 66 | 0.0099993 | 1e-07 | 4.4984e-06 | 5e-07 | 0.0099994 | 1e-07 | 0.00044987 | 5.0003e-05 | 50Hz10m10m | 2.5877e-07 | 0.0037452 | 2.1659e-08 | 0.0033887 | 50Hz10m | 50Hz10m | 84 | 1 | 4.7426e-06 | 2.1544e-06 | 0.00051459 | OK
10 | m | 100 | 10 | m | 9.9978 | 0.0018388 | 0.0091269 | 0.00089827 | m | -0.001965 | 0.0018266 | 0.00031911 | 0.00095926 | m | OK | 0.01 | 100 | Hz | 0.01 | 0.001 | 0.0099998 | 2.5919e-06 | 8.8078e-06 | 1.3142e-06 | 0.0099998 | 2.5919e-06 | 0.0008808 | 0.00013142 | 100Hz10m | 67 | 0.010001 | 1e-07 | 9.5713e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00095706 | 4.9996e-05 | 100Hz10m10m | 9.5218e-07 | 0.0036777 | 7.6347e-07 | 0.0017965 | 100Hz10m | 100Hz10m | 85 | 1.0001 | 5.1837e-06 | 7.6258e-05 | 0.00026755 | OK
10 | m | 200 | 10 | m | 9.9991 | 0.0020429 | 0.015977 | 0.0013582 | m | -0.00055057 | 0.0013993 | -0.0010169 | 0.00032034 | m | OK | 0.01 | 200 | Hz | 0.01 | 0.001 | 0.0099996 | 2.4762e-06 | 1.6994e-05 | 1.3954e-06 | 0.0099997 | 2.4762e-06 | 0.0016995 | 0.00013955 | 200Hz10m | 68 | 0.01 | 1e-07 | 1.7729e-05 | 5e-07 | 0.01 | 1e-07 | 0.0017728 | 4.9997e-05 | 200Hz10m10m | 8.3222e-07 | 0.0040858 | 7.3506e-07 | 0.0027163 | 200Hz10m | 200Hz10m | 86 | 1.0001 | 4.9523e-06 | 7.3361e-05 | 0.00028354 | OK
10 | m | 500 | 10 | m | 10.004 | 6.6917e-05 | 0.044164 | 0.00047784 | m | -0.0012973 | 0.0014314 | -0.00066973 | 0.00030254 | m | OK | 0.01 | 500 | Hz | 0.01 | 0.001 | 0.010005 | 1.4329e-06 | 4.4833e-05 | 5.6556e-07 | 0.010005 | 1.4329e-06 | 0.004481 | 5.6529e-05 | 500Hz10m | 69 | 0.010006 | 1.5811e-07 | 4.5356e-05 | 5e-07 | 0.010006 | 1.5813e-07 | 0.0045329 | 4.9969e-05 | 500Hz10m10m | 6.6842e-07 | 0.00013383 | 5.2239e-07 | 0.00095568 | 500Hz10m | 500Hz10m | 87 | 1.0001 | 2.8658e-06 | 5.1907e-05 | 0.00012361 | OK
10 | m | 1000 | 10 | m | 10.0177 | 0.00055185 | 0.091546 | 0.00039564 | m | -0.00027255 | 0.00068504 | -0.0013441 | 0.00092029 | m | OK | 0.01 | 1 | kHz | 0.01 | 0.001 | 0.010018 | 8.7967e-07 | 9.289e-05 | 1.0017e-06 | 0.010018 | 8.7968e-07 | 0.0092721 | 9.9988e-05 | 1kHz10m | 70 | 0.010018 | 4.4721e-07 | 9.3275e-05 | 1e-06 | 0.010019 | 4.4729e-07 | 0.0093101 | 9.9808e-05 | 1kHz10m10m | 4.3184e-07 | 0.0011037 | 3.8473e-07 | 0.00079127 | 1kHz10m | 1kHz10m | 88 | 1 | 1.7594e-06 | 3.8e-05 | 0.0002235 | OK
10 | m | 2000 | 10 | m | 10.0494 | 0.0018276 | 0.18801 | 0.0019671 | m | -0.00098371 | 0.0008405 | 0.00026273 | 0.0019752 | m | OK | 0.01 | 2 | kHz | 0.01 | 0.001 | 0.01005 | 2.0116e-06 | 0.00018775 | 2.7876e-06 | 0.010052 | 2.0119e-06 | 0.018679 | 0.00027729 | 2kHz10m | 71 | 0.010051 | 1.2649e-06 | 0.00018716 | 2e-06 | 0.010052 | 1.2652e-06 | 0.01862 | 0.00019894 | 2kHz10m10m | 3.2318e-07 | 0.0036551 | -5.8467e-07 | 0.0039341 | 2kHz10m | 2kHz10m | 89 | 1 | 4.0238e-06 | -5.8752e-05 | 0.00058918 | OK
10 | m | 5000 | 10 | m | 10.1981 | 0.00094658 | 0.46825 | 0.00046664 | m | 6.5625e-05 | 0.00098072 | -0.0011405 | 0.0016417 | m | OK | 0.01 | 5 | kHz | 0.01 | 0.001 | 0.010198 | 1.363e-06 | 0.00046939 | 1.7067e-06 | 0.010209 | 1.3638e-06 | 0.045995 | 0.00016712 | 5kHz10m | 72 | 0.010199 | 5e-06 | 0.00046839 | 5e-06 | 0.01021 | 5e-06 | 0.045892 | 0.00048972 | 5kHz10m10m | 9.8702e-07 | 0.0018932 | -1.0028e-06 | 0.0009333 | 5kHz10m | 5kHz10m | 90 | 1.0001 | 2.7282e-06 | -0.00010256 | 0.00059291 | OK
100 | m | 0.01 | 10 | m | 10.0021 | 0.0016527 | -0.0008145 | 0.0010776 | m | 0.0019321 | 0.00053966 | 0.00024089 | 0.0015101 | m | OK | 0.1 | 10 | mHz | 0.01 | 0.001 | 0.01 | 1.7385e-06 | -1.0554e-06 | 1.8552e-06 | 0.01 | 1.7385e-06 | -0.00010554 | 0.00018551 | 10mHz10m | 55 | 0.01 | 1e-07 | -5.5374e-08 | 5e-07 | 0.01 | 1e-07 | -5.5372e-06 | 4.9998e-05 | 10mHz100m10m | 2.0795e-07 | 0.0033053 | 1e-06 | 0.0021551 | 10mHz100m | 10mHz100m | 109 | 1 | 3.4771e-06 | 9.9999e-05 | 0.00037438 | OK
100 | m | 0.02 | 10 | m | 10.0009 | 0.0015173 | -0.0012551 | 0.0017859 | m | -0.0012832 | 0.00022362 | -0.0015836 | 0.0012155 | m | OK | 0.1 | 20 | mHz | 0.01 | 0.001 | 0.010002 | 1.5337e-06 | 3.2846e-07 | 2.1603e-06 | 0.010002 | 1.5337e-06 | 3.2838e-05 | 0.00021598 | 20mHz10m | 56 | 0.010001 | 1e-07 | 1.3286e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00013284 | 4.9993e-05 | 20mHz100m10m | -7.5251e-07 | 0.0030346 | 1.0001e-06 | 0.0035718 | 20mHz100m | 20mHz100m | 110 | 0.99992 | 3.0673e-06 | 9.9999e-05 | 0.00043485 | OK
100 | m | 0.05 | 10 | m | 10.0001 | 0.00065158 | -0.0032544 | 0.0016579 | m | -0.00012885 | 0.0013969 | -0.0011096 | 0.00089146 | m | OK | 0.1 | 50 | mHz | 0.01 | 0.001 | 0.01 | 1.5414e-06 | -2.1448e-06 | 1.8824e-06 | 0.01 | 1.5414e-06 | -0.00021448 | 0.00018824 | 50mHz10m | 57 | 0.010001 | 1e-07 | -1.1449e-06 | 5e-07 | 0.010001 | 1e-07 | -0.00011448 | 4.9996e-05 | 50mHz100m10m | 4.7128e-07 | 0.0013032 | 9.9997e-07 | 0.0033159 | 50mHz100m | 50mHz100m | 111 | 1 | 3.0828e-06 | 0.0001 | 0.00037978 | OK
100 | m | 0.1 | 10 | m | 9.9994 | 0.0013784 | -0.0025697 | 0.00079133 | m | -0.0017298 | 0.0014965 | -0.0014066 | 0.00146 | m | OK | 0.1 | 100 | mHz | 0.01 | 0.001 | 0.010001 | 2.0346e-06 | -1.1631e-06 | 1.6606e-06 | 0.010001 | 2.0346e-06 | -0.0001163 | 0.00016605 | 100mHz10m | 58 | 0.01 | 1e-07 | -1.6302e-07 | 5e-07 | 0.01 | 1e-07 | -1.6301e-05 | 4.9998e-05 | 100mHz100m10m | -6.9929e-07 | 0.0027568 | 1.0001e-06 | 0.0015827 | 100mHz100m | 100mHz100m | 112 | 0.99993 | 4.0692e-06 | 9.9999e-05 | 0.00033583 | OK
100 | m | 0.2 | 10 | m | 9.999 | 0.0013621 | -0.0012649 | 0.00059684 | m | -0.00070435 | 0.0011228 | -0.0010123 | 0.0015071 | m | OK | 0.1 | 200 | mHz | 0.01 | 0.001 | 0.0099998 | 1.7652e-06 | -2.5261e-07 | 1.621e-06 | 0.0099998 | 1.7652e-06 | -2.5262e-05 | 0.0001621 | 200mHz10m | 59 | 0.0099996 | 1e-07 | 7.4735e-07 | 5e-07 | 0.0099996 | 1e-07 | 7.4739e-05 | 5.0002e-05 | 200mHz100m10m | -2.0189e-07 | 0.0027242 | 9.9997e-07 | 0.0011937 | 200mHz100m | 200mHz100m | 113 | 0.99998 | 3.5304e-06 | 0.0001 | 0.00032803 | OK
100 | m | 0.5 | 10 | m | 10.0003 | 0.0014719 | -0.00079598 | 0.00092605 | m | -0.00062742 | 0.00058519 | 0.00066399 | 0.001535 | m | OK | 0.1 | 500 | mHz | 0.01 | 0.001 | 0.010001 | 1.584e-06 | -1.46e-06 | 1.7927e-06 | 0.010001 | 1.584e-06 | -0.00014598 | 0.00017925 | 500mHz10m | 60 | 0.01 | 1e-07 | -4.5986e-07 | 5e-07 | 0.01 | 1e-07 | -4.5985e-05 | 4.9998e-05 | 500mHz100m10m | -5.3831e-07 | 0.0029439 | 1.0001e-06 | 0.0018521 | 500mHz100m | 500mHz100m | 114 | 0.99995 | 3.168e-06 | 9.9999e-05 | 0.00036198 | OK
100 | m | 1 | 10 | m | 9.9982 | 0.0013777 | -0.0028454 | 0.0011854 | m | -7.0605e-05 | 0.00060244 | -0.0013462 | 0.00092295 | m | OK | 0.1 | 1 | Hz | 0.01 | 0.001 | 0.0099983 | 1.5037e-06 | -1.4993e-06 | 1.5023e-06 | 0.0099983 | 1.5037e-06 | -0.00014995 | 0.00015026 | 1Hz10m | 61 | 0.009999 | 1e-07 | -4.9945e-07 | 5e-07 | 0.009999 | 1e-07 | -4.995e-05 | 5.0005e-05 | 1Hz100m10m | 7.4691e-07 | 0.0027555 | 9.998e-07 | 0.0023707 | 1Hz100m | 1Hz100m | 115 | 1.0001 | 3.0074e-06 | 0.0001 | 0.00030464 | OK
100 | m | 2 | 10 | m | 9.9976 | 0.00095957 | 0.00051123 | 0.0011923 | m | -0.0012763 | 0.0016836 | 0.00056674 | 0.0017998 | m | OK | 0.1 | 2 | Hz | 0.01 | 0.001 | 0.0099989 | 1.9378e-06 | -5.5512e-08 | 2.1589e-06 | 0.0099989 | 1.9378e-06 | -5.5518e-06 | 0.00021591 | 2Hz10m | 62 | 0.009998 | 1e-07 | 9.4431e-07 | 5e-07 | 0.009998 | 1e-07 | 9.445e-05 | 5.001e-05 | 2Hz100m10m | -9.3168e-07 | 0.0019191 | 9.9982e-07 | 0.0023846 | 2Hz100m | 2Hz100m | 116 | 0.99991 | 3.8756e-06 | 0.0001 | 0.00043471 | OK
100 | m | 5 | 10 | m | 9.9982 | 0.0018643 | -0.0036659 | 0.0018183 | m | -0.0014883 | 0.0015881 | -0.001675 | 0.00083769 | m | OK | 0.1 | 5 | Hz | 0.01 | 0.001 | 0.0099996 | 2.449e-06 | -1.9909e-06 | 2.002e-06 | 0.0099996 | 2.449e-06 | -0.0001991 | 0.0002002 | 5Hz10m | 63 | 0.01 | 1e-07 | -9.9104e-07 | 5e-07 | 0.01 | 1e-07 | -9.9102e-05 | 4.9999e-05 | 5Hz100m10m | 5.6629e-07 | 0.0037286 | 9.9991e-07 | 0.0036366 | 5Hz100m | 5Hz100m | 117 | 1.0001 | 4.898e-06 | 0.0001 | 0.00040352 | OK
100 | m | 10 | 10 | m | 10.0031 | 0.0011786 | -0.0017465 | 0.0016673 | m | 0.001426 | 0.00068824 | -0.00083601 | 0.00043344 | m | OK | 0.1 | 10 | Hz | 0.01 | 0.001 | 0.010002 | 1.3648e-06 | -9.1052e-07 | 1.7227e-06 | 0.010002 | 1.3648e-06 | -9.1037e-05 | 0.00017224 | 10Hz10m | 64 | 0.010001 | 1e-07 | 8.9632e-08 | 5e-07 | 0.010001 | 1e-07 | 8.9624e-06 | 4.9995e-05 | 10Hz100m10m | -7.0881e-07 | 0.0023572 | 1.0002e-06 | 0.0033346 | 10Hz100m | 10Hz100m | 118 | 0.99993 | 2.7297e-06 | 9.9999e-05 | 0.0003481 | OK
100 | m | 20 | 10 | m | 10.0011 | 0.0015211 | 0.0022343 | 9.3715e-05 | m | 0.00091446 | 0.0010522 | 0.0016046 | 0.001212 | m | OK | 0.1 | 20 | Hz | 0.01 | 0.001 | 0.01 | 1.8496e-06 | 6.2973e-07 | 1.2157e-06 | 0.01 | 1.8496e-06 | 6.2972e-05 | 0.00012156 | 20Hz10m | 65 | 0.0099998 | 1e-07 | 1.6297e-06 | 5e-07 | 0.0099998 | 1e-07 | 0.00016297 | 5.0001e-05 | 20Hz100m10m | -4.1245e-07 | 0.0030423 | 9.9995e-07 | 0.00018743 | 20Hz100m | 20Hz100m | 119 | 0.99996 | 3.6992e-06 | 0.0001 | 0.00024822 | OK
100 | m | 50 | 10 | m | 10.0011 | 0.00077152 | 0.0033265 | 0.0018453 | m | 0.0010309 | 0.0017462 | -0.00017226 | 0.00081722 | m | OK | 0.1 | 50 | Hz | 0.01 | 0.001 | 0.01 | 1.909e-06 | 3.4987e-06 | 2.0182e-06 | 0.01 | 1.909e-06 | 0.00034987 | 0.00020182 | 50Hz10m | 66 | 0.0099993 | 1e-07 | 4.4984e-06 | 5e-07 | 0.0099994 | 1e-07 | 0.00044987 | 5.0003e-05 | 50Hz100m10m | -6.9444e-07 | 0.001543 | 9.9971e-07 | 0.0036907 | 50Hz100m | 50Hz100m | 120 | 0.99993 | 3.818e-06 | 0.0001 | 0.00040672 | OK
100 | m | 100 | 10 | m | 10.0005 | 0.00082488 | 0.0082409 | 0.0014019 | m | -0.00027149 | 0.00040904 | -0.00033027 | 1.6655e-05 | m | OK | 0.1 | 100 | Hz | 0.01 | 0.001 | 0.010001 | 9.2072e-07 | 8.5712e-06 | 1.402e-06 | 0.010001 | 9.2073e-07 | 0.00085706 | 0.00014019 | 100Hz10m | 67 | 0.010001 | 1e-07 | 9.5713e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00095706 | 4.9996e-05 | 100Hz100m10m | -1.3815e-09 | 0.0016498 | 1.0001e-06 | 0.0028037 | 100Hz100m | 100Hz100m | 121 | 1 | 1.8415e-06 | 9.9999e-05 | 0.0002848 | OK
100 | m | 200 | 10 | m | 9.9992 | 0.0006179 | 0.016846 | 0.0011796 | m | -0.0019305 | 0.0010664 | 0.00011614 | 0.00039479 | m | OK | 0.1 | 200 | Hz | 0.01 | 0.001 | 0.010001 | 1.2325e-06 | 1.673e-05 | 1.2439e-06 | 0.010001 | 1.2325e-06 | 0.0016728 | 0.00012438 | 200Hz10m | 68 | 0.01 | 1e-07 | 1.7729e-05 | 5e-07 | 0.01 | 1e-07 | 0.0017728 | 4.9997e-05 | 200Hz100m10m | -6.8486e-07 | 0.0012358 | 9.989e-07 | 0.0023593 | 200Hz100m | 200Hz100m | 122 | 0.99993 | 2.465e-06 | 9.9999e-05 | 0.00025373 | OK
100 | m | 500 | 10 | m | 10.0058 | 0.0020083 | 0.045832 | 0.0019721 | m | -0.00081519 | 0.0013201 | 0.001474 | 0.00023109 | m | OK | 0.1 | 500 | Hz | 0.01 | 0.001 | 0.010007 | 2.4033e-06 | 4.4358e-05 | 1.9856e-06 | 0.010007 | 2.4033e-06 | 0.0044329 | 0.00019843 | 500Hz10m | 69 | 0.010006 | 1.5811e-07 | 4.5356e-05 | 5e-07 | 0.010006 | 1.5813e-07 | 0.0045329 | 4.9969e-05 | 500Hz100m10m | -6.7937e-07 | 0.0040166 | 9.9753e-07 | 0.0039442 | 500Hz100m | 500Hz100m | 123 | 0.99993 | 4.8066e-06 | 9.9993e-05 | 0.00039999 | OK
100 | m | 1000 | 10 | m | 10.019 | 0.00036304 | 0.091626 | 0.00018911 | m | 0.00011479 | 0.0014916 | -0.00065195 | 0.00077067 | m | OK | 0.1 | 1 | kHz | 0.01 | 0.001 | 0.010019 | 1.5352e-06 | 9.2278e-05 | 7.9353e-07 | 0.010019 | 1.5351e-06 | 0.0092101 | 7.921e-05 | 1kHz10m | 70 | 0.010018 | 4.4721e-07 | 9.3275e-05 | 1e-06 | 0.010019 | 4.4729e-07 | 0.0093101 | 9.9808e-05 | 1kHz100m10m | -4.9574e-07 | 0.00072607 | 9.9715e-07 | 0.00037823 | 1kHz100m | 1kHz100m | 124 | 0.99995 | 3.0702e-06 | 9.9979e-05 | 0.00018724 | OK
100 | m | 2000 | 10 | m | 10.0492 | 0.00028971 | 0.18435 | 5.0915e-05 | m | -0.0015314 | 0.0008732 | -0.0018124 | 0.0019227 | m | OK | 0.1 | 2 | kHz | 0.01 | 0.001 | 0.010051 | 9.2001e-07 | 0.00018616 | 1.9234e-06 | 0.010052 | 9.2054e-07 | 0.01852 | 0.00019131 | 2kHz10m | 71 | 0.010051 | 1.2649e-06 | 0.00018716 | 2e-06 | 0.010052 | 1.2652e-06 | 0.01862 | 0.00019894 | 2kHz100m10m | -2.5323e-08 | 0.00057942 | 1.0047e-06 | 0.00010185 | 2kHz100m | 2kHz100m | 125 | 1 | 1.8411e-06 | 9.9971e-05 | 0.00043125 | OK
100 | m | 5000 | 10 | m | 10.2014 | 0.0018074 | 0.46691 | 6.0981e-05 | m | 0.0018994 | 0.0019157 | -0.00048366 | 0.00077231 | m | OK | 0.1 | 5 | kHz | 0.01 | 0.001 | 0.010199 | 2.6337e-06 | 0.00046739 | 7.7471e-07 | 0.01021 | 2.6312e-06 | 0.045793 | 7.6711e-05 | 5kHz10m | 72 | 0.010199 | 5e-06 | 0.00046839 | 5e-06 | 0.01021 | 5e-06 | 0.045892 | 0.00048972 | 5kHz100m10m | -4.0201e-07 | 0.0036149 | 9.9975e-07 | 0.00012206 | 5kHz100m | 5kHz100m | 126 | 0.99997 | 5.2627e-06 | 9.962e-05 | 0.00051319 | OK
100 | m | 0.01 | 100 | m | 100.0066 | 0.0015274 | 0.010191 | 0.0010187 | m | -0.0017948 | 0.00048443 | 0.00022369 | 0.00025428 | m | OK | 0.1 | 10 | mHz | 0.1 | 0.001 | 0.10001 | 1.6024e-06 | 9.9673e-06 | 1.05e-06 | 0.10001 | 1.6024e-06 | 9.9664e-05 | 1.0499e-05 | 10mHz100m | 73 | 0.099999 | 1e-06 | 7.8026e-07 | 5e-06 | 0.099999 | 1e-06 | 7.8027e-06 | 5e-05 | 10mHz100m100m | -9.3807e-06 | 0.0030549 | -9.187e-06 | 0.0020374 | 10mHz100m | 10mHz100m | 109 | 0.99991 | 3.2064e-06 | -9.1862e-05 | 5.4231e-05 | OK
100 | m | 0.02 | 100 | m | 99.9911 | 0.0020015 | -0.0087356 | 0.0019394 | m | 0.00085004 | 0.00091148 | 0.0018387 | 0.0013519 | m | OK | 0.1 | 20 | mHz | 0.1 | 0.001 | 0.09999 | 2.1993e-06 | -1.0574e-05 | 2.3641e-06 | 0.09999 | 2.1993e-06 | -0.00010575 | 2.3643e-05 | 20mHz100m | 74 | 0.099992 | 1e-06 | -4.5184e-06 | 5e-06 | 0.099992 | 1e-06 | -4.5187e-05 | 5.0004e-05 | 20mHz100m100m | 2.19e-06 | 0.004003 | 6.056e-06 | 0.0038788 | 20mHz100m | 20mHz100m | 110 | 1 | 4.3997e-06 | 6.0567e-05 | 6.8822e-05 | OK
100 | m | 0.05 | 100 | m | 99.9913 | 0.0016193 | -0.030059 | 0.0017864 | m | -0.0011789 | 0.0019102 | -0.0016915 | 0.00029879 | m | OK | 0.1 | 50 | mHz | 0.1 | 0.001 | 0.099992 | 2.5042e-06 | -2.8368e-05 | 1.8113e-06 | 0.099992 | 2.5042e-06 | -0.0002837 | 1.8114e-05 | 50mHz100m | 75 | 0.099995 | 1e-06 | -1.8542e-05 | 5e-06 | 0.099995 | 1e-06 | -0.00018543 | 5.0003e-05 | 50mHz100m100m | 2.0948e-06 | 0.0032386 | 9.8259e-06 | 0.0035729 | 50mHz100m | 50mHz100m | 111 | 1 | 5.0094e-06 | 9.827e-05 | 6.1747e-05 | OK
100 | m | 0.1 | 100 | m | 100.0051 | 0.0027664 | 0.0040218 | 0.0015978 | m | -0.00047417 | 0.00060716 | -0.00040726 | 1.6028e-05 | m | OK | 0.1 | 100 | mHz | 0.1 | 0.001 | 0.10001 | 2.8322e-06 | 4.429e-06 | 1.5979e-06 | 0.10001 | 2.8322e-06 | 4.4288e-05 | 1.5978e-05 | 100mHz100m | 76 | 0.10001 | 1e-06 | -3.6364e-06 | 5e-06 | 0.10001 | 1e-06 | -3.6361e-05 | 4.9997e-05 | 100mHz100m100m | 1.3731e-06 | 0.0055328 | -8.0654e-06 | 0.0031956 | 100mHz100m | 100mHz100m | 112 | 1 | 5.6654e-06 | -8.0649e-05 | 5.9337e-05 | OK
100 | m | 0.2 | 100 | m | 99.9904 | 0.0021295 | -0.010212 | 0.0015946 | m | 0.0013005 | 0.0015164 | 0.00065736 | 0.0011669 | m | OK | 0.1 | 200 | mHz | 0.1 | 0.001 | 0.099989 | 2.6142e-06 | -1.0869e-05 | 1.9759e-06 | 0.099989 | 2.6142e-06 | -0.0001087 | 1.9762e-05 | 200mHz100m | 77 | 0.099983 | 1e-06 | -1.368e-05 | 5e-06 | 0.099983 | 1e-06 | -0.00013682 | 5.0009e-05 | 200mHz100m100m | -6.2211e-06 | 0.004259 | -2.8108e-06 | 0.0031892 | 200mHz100m | 200mHz100m | 113 | 0.99994 | 5.2294e-06 | -2.812e-05 | 6.3741e-05 | OK
100 | m | 0.5 | 100 | m | 99.9922 | 0.0024075 | 0.00036272 | 0.00076669 | m | -0.0017471 | 0.0010182 | -0.00018548 | 0.0016919 | m | OK | 0.1 | 500 | mHz | 0.1 | 0.001 | 0.099994 | 2.614e-06 | 5.482e-07 | 1.8575e-06 | 0.099994 | 2.614e-06 | 5.4823e-06 | 1.8576e-05 | 500mHz100m | 78 | 0.099994 | 1e-06 | -7.8555e-06 | 5e-06 | 0.099994 | 1e-06 | -7.8561e-05 | 5.0003e-05 | 500mHz100m100m | -3.4265e-07 | 0.004815 | -8.4037e-06 | 0.0015334 | 500mHz100m | 500mHz100m | 114 | 1 | 5.2289e-06 | -8.4043e-05 | 6.2294e-05 | OK
100 | m | 1 | 100 | m | 99.9879 | 0.0019921 | -0.0091539 | 0.0012259 | m | 0.0011875 | 0.0010203 | -0.00034368 | 0.00035064 | m | OK | 0.1 | 1 | Hz | 0.1 | 0.001 | 0.099987 | 2.2382e-06 | -8.8103e-06 | 1.275e-06 | 0.099987 | 2.2382e-06 | -8.8114e-05 | 1.2752e-05 | 1Hz100m | 79 | 0.099994 | 1e-06 | -6.1224e-06 | 5e-06 | 0.099994 | 1e-06 | -6.1228e-05 | 5.0003e-05 | 1Hz100m100m | 6.8936e-06 | 0.0039843 | 2.6878e-06 | 0.0024517 | 1Hz100m | 1Hz100m | 115 | 1.0001 | 4.4776e-06 | 2.6886e-05 | 5.6132e-05 | OK
100 | m | 2 | 100 | m | 100.0272 | 0.0023253 | 0.0072499 | 0.0007053 | m | 0.00085553 | 0.0012344 | -8.4082e-05 | 0.0017804 | m | OK | 0.1 | 2 | Hz | 0.1 | 0.001 | 0.10003 | 2.6327e-06 | 7.3339e-06 | 1.915e-06 | 0.10003 | 2.6327e-06 | 7.332e-05 | 1.9145e-05 | 2Hz100m | 80 | 0.10002 | 1e-06 | 9.6162e-06 | 5e-06 | 0.10002 | 1e-06 | 9.6141e-05 | 4.9989e-05 | 2Hz100m100m | -4.3209e-06 | 0.0046507 | 2.2823e-06 | 0.0014106 | 2Hz100m | 2Hz100m | 116 | 0.99996 | 5.2663e-06 | 2.2821e-05 | 6.2969e-05 | OK
100 | m | 5 | 100 | m | 99.9841 | 0.0015252 | -0.00016611 | 0.0011851 | m | -0.00166 | 0.0014969 | -0.0014404 | 0.00048986 | m | OK | 0.1 | 5 | Hz | 0.1 | 0.001 | 0.099986 | 2.1371e-06 | 1.2743e-06 | 1.2823e-06 | 0.099986 | 2.1371e-06 | 1.2745e-05 | 1.2825e-05 | 5Hz100m | 81 | 0.099991 | 1e-06 | 9.8483e-06 | 5e-06 | 0.099991 | 1e-06 | 9.8492e-05 | 5.0005e-05 | 5Hz100m100m | 5.1461e-06 | 0.0030504 | 8.574e-06 | 0.0023701 | 5Hz100m | 5Hz100m | 117 | 1.0001 | 4.2753e-06 | 8.5747e-05 | 5.62e-05 | OK
100 | m | 10 | 100 | m | 99.9797 | 0.0012293 | -0.0055236 | 0.0013177 | m | -0.0011731 | 0.00061129 | 0.00046309 | 8.0096e-05 | m | OK | 0.1 | 10 | Hz | 0.1 | 0.001 | 0.099981 | 1.3729e-06 | -5.9867e-06 | 1.3201e-06 | 0.099981 | 1.3729e-06 | -5.9879e-05 | 1.3204e-05 | 10Hz100m | 82 | 0.099977 | 1e-06 | -1.7626e-06 | 5e-06 | 0.099977 | 1e-06 | -1.763e-05 | 5.0011e-05 | 10Hz100m100m | -3.7017e-06 | 0.0024586 | 4.2241e-06 | 0.0026354 | 10Hz100m | 10Hz100m | 118 | 0.99996 | 2.7476e-06 | 4.2249e-05 | 5.6555e-05 | OK
100 | m | 20 | 100 | m | 100.0197 | 0.0014527 | 0.016756 | 0.0006996 | m | 0.0002229 | 0.00058692 | -0.0015136 | 0.00028047 | m | OK | 0.1 | 20 | Hz | 0.1 | 0.001 | 0.10002 | 1.5668e-06 | 1.827e-05 | 7.5372e-07 | 0.10002 | 1.5668e-06 | 0.00018266 | 7.5358e-06 | 20Hz100m | 83 | 0.10002 | 1e-06 | 1.1488e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00011486 | 4.9992e-05 | 20Hz100m100m | -2.4634e-06 | 0.0029054 | -6.7817e-06 | 0.0013992 | 20Hz100m | 20Hz100m | 119 | 0.99998 | 3.1351e-06 | -6.7801e-05 | 5.2214e-05 | OK
100 | m | 50 | 100 | m | 99.9982 | 0.0012871 | 0.059477 | 0.00094142 | m | 0.00034325 | 0.0014274 | -0.0016878 | 0.0014203 | m | OK | 0.1 | 50 | Hz | 0.1 | 0.001 | 0.099998 | 1.9221e-06 | 6.1165e-05 | 1.7039e-06 | 0.099998 | 1.9221e-06 | 0.00061166 | 1.704e-05 | 50Hz100m | 84 | 0.1 | 1e-06 | 5.2501e-05 | 5e-06 | 0.1 | 1e-06 | 0.000525 | 5e-05 | 50Hz100m100m | 2.842e-06 | 0.0025742 | -8.6638e-06 | 0.0018828 | 50Hz100m | 50Hz100m | 120 | 1 | 3.8454e-06 | -8.6654e-05 | 6.0509e-05 | OK
100 | m | 100 | 100 | m | 100.0089 | 0.00087995 | 0.095338 | 0.0011987 | m | -0.0010457 | 0.0018684 | -1.8913e-05 | 0.0011274 | m | OK | 0.1 | 100 | Hz | 0.1 | 0.001 | 0.10001 | 2.0653e-06 | 9.5357e-05 | 1.6456e-06 | 0.10001 | 2.0653e-06 | 0.00095347 | 1.6454e-05 | 100Hz100m | 85 | 0.10002 | 1e-06 | 8.934e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00089326 | 4.9992e-05 | 100Hz100m100m | 6.3062e-06 | 0.0017599 | -6.0165e-06 | 0.0023973 | 100Hz100m | 100Hz100m | 121 | 1.0001 | 4.1317e-06 | -6.0215e-05 | 5.9851e-05 | OK
100 | m | 200 | 100 | m | 100.0114 | 0.0019878 | 0.1911 | 0.0016387 | m | -0.00022593 | 2.295e-05 | -0.0019002 | 0.00064596 | m | OK | 0.1 | 200 | Hz | 0.1 | 0.001 | 0.10001 | 1.9879e-06 | 0.000193 | 1.7614e-06 | 0.10001 | 1.9879e-06 | 0.0019298 | 1.7612e-05 | 200Hz100m | 86 | 0.10001 | 1e-06 | 0.000189 | 5e-06 | 0.10001 | 1e-06 | 0.0018898 | 4.9994e-05 | 200Hz100m100m | -2.9553e-07 | 0.0039756 | -4.0027e-06 | 0.0032774 | 200Hz100m | 200Hz100m | 122 | 1 | 3.9771e-06 | -4.0017e-05 | 6.1157e-05 | OK
100 | m | 500 | 100 | m | 100.0311 | 0.00093129 | 0.45966 | 0.0014001 | m | 0.0011768 | 0.00089999 | 0.0018408 | 0.0010821 | m | OK | 0.1 | 500 | Hz | 0.1 | 0.001 | 0.10003 | 1.2951e-06 | 0.00045782 | 1.7695e-06 | 0.10003 | 1.2951e-06 | 0.0045768 | 1.769e-05 | 500Hz100m | 87 | 0.10004 | 1.5811e-06 | 0.00045737 | 5e-06 | 0.10004 | 1.5813e-06 | 0.0045719 | 4.998e-05 | 500Hz100m100m | 8.1089e-06 | 0.0018626 | -4.5758e-07 | 0.0028002 | 500Hz100m | 500Hz100m | 123 | 1.0001 | 2.5951e-06 | -4.945e-06 | 6.1235e-05 | OK
100 | m | 1000 | 100 | m | 100.1777 | 0.0012848 | 0.91918 | 0.0011715 | m | -0.0015329 | 0.0011192 | -0.00018531 | 0.00076816 | m | OK | 0.1 | 1 | kHz | 0.1 | 0.001 | 0.10018 | 1.704e-06 | 0.00091937 | 1.4009e-06 | 0.10018 | 1.7039e-06 | 0.009177 | 1.3983e-05 | 1kHz100m | 88 | 0.10017 | 4.4721e-06 | 0.00091905 | 1e-05 | 0.10018 | 4.4729e-06 | 0.0091744 | 9.982e-05 | 1kHz100m100m | -6.5227e-06 | 0.0025697 | -3.2015e-07 | 0.002343 | 1kHz100m | 1kHz100m | 124 | 0.99993 | 3.4372e-06 | -2.5982e-06 | 0.00010366 | OK
100 | m | 2000 | 100 | m | 100.4892 | 0.00099623 | 1.8414 | 0.0013939 | m | -0.00062187 | 0.001777 | -0.00027636 | 0.0018825 | m | OK | 0.1 | 2 | kHz | 0.1 | 0.001 | 0.10049 | 2.0372e-06 | 0.0018417 | 2.3424e-06 | 0.10051 | 2.0373e-06 | 0.018325 | 2.3305e-05 | 2kHz100m | 89 | 0.1005 | 1.2649e-05 | 0.0018469 | 2e-05 | 0.10051 | 1.2652e-05 | 0.018376 | 0.00019896 | 2kHz100m100m | 7.1609e-06 | 0.0019925 | 5.2033e-06 | 0.0027879 | 2kHz100m | 2kHz100m | 125 | 1.0001 | 4.2684e-06 | 5.0453e-05 | 0.00020434 | OK
100 | m | 5000 | 100 | m | 101.9946 | 0.00096269 | 4.6654 | 0.001065 | m | 0.0019403 | 0.00049797 | 0.00087175 | 0.00052482 | m | OK | 0.1 | 5 | kHz | 0.1 | 0.001 | 0.10199 | 1.0839e-06 | 0.0046645 | 1.1873e-06 | 0.1021 | 1.0841e-06 | 0.045702 | 1.1627e-05 | 5kHz100m | 90 | 0.10199 | 5e-05 | 0.0046575 | 5e-05 | 0.1021 | 5e-05 | 0.045635 | 0.00048974 | 5kHz100m100m | -3.7561e-06 | 0.001926 | -6.999e-06 | 0.0021307 | 5kHz100m | 5kHz100m | 126 | 0.99996 | 5.5463e-06 | -6.6801e-05 | 0.00049029 | OK

## Ref Z
This is sheet containing calibration data of reference standards to be used for calibration of RLC meter.
Calibration data of ref. Z standards
f | Nominal Z |  | Rs | U(Rs) | Xs | U(Xs) | Rs-Xs unit | f |  | Nom Z | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | tag
Hz | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | - |  | Ohm | - | Ohm | Ohm | Ohm | Ohm | -
0.01 | 0 | m | 0 | 0 | 0 | 0 | m | 10 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 10mHz0m
0.02 | 0 | m | 0 | 0 | 0 | 0 | m | 20 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 20mHz0m
0.05 | 0 | m | 0 | 0 | 0 | 0 | m | 50 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 50mHz0m
0.1 | 0 | m | 0 | 0 | 0 | 0 | m | 100 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 100mHz0m
0.2 | 0 | m | 0 | 0 | 0 | 0 | m | 200 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 200mHz0m
0.5 | 0 | m | 0 | 0 | 0 | 0 | m | 500 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 500mHz0m
1 | 0 | m | 0 | 0 | 0 | 0 | m | 1 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 1Hz0m
2 | 0 | m | 0 | 0 | 0 | 0 | m | 2 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 2Hz0m
5 | 0 | m | 0 | 0 | 0 | 0 | m | 5 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 5Hz0m
10 | 0 | m | 0 | 0 | 0 | 0 | m | 10 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 10Hz0m
20 | 0 | m | 0 | 0 | 0 | 0 | m | 20 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 20Hz0m
50 | 0 | m | 0 | 0 | 0 | 0 | m | 50 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 50Hz0m
100 | 0 | m | 0 | 0 | 0 | 0 | m | 100 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 100Hz0m
200 | 0 | m | 0 | 0 | 0 | 0 | m | 200 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 200Hz0m
500 | 0 | m | 0 | 0 | 0 | 0 | m | 500 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 500Hz0m
1000 | 0 | m | 0 | 0 | 0 | 0 | m | 1 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 1kHz0m
2000 | 0 | m | 0 | 0 | 0 | 0 | m | 2 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 2kHz0m
5000 | 0 | m | 0 | 0 | 0 | 0 | m | 5 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 5kHz0m
0.01 | 1 | m | 1.0001 | 1e-05 | 5.0774e-05 | 5e-05 | m | 10 | mHz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 5.0774e-08 | 5e-08 | 10mHz1m
0.02 | 1 | m | 1.0001 | 1e-05 | 6.3549e-05 | 5e-05 | m | 20 | mHz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 6.3549e-08 | 5e-08 | 20mHz1m
0.05 | 1 | m | 0.99991 | 1e-05 | -7.6899e-05 | 5e-05 | m | 50 | mHz | 0.001 | 0.001 | 0.00099991 | 1e-08 | -7.6899e-08 | 5e-08 | 50mHz1m
0.1 | 1 | m | 1 | 1e-05 | -3.8534e-05 | 5e-05 | m | 100 | mHz | 0.001 | 0.001 | 0.001 | 1e-08 | -3.8534e-08 | 5e-08 | 100mHz1m
0.2 | 1 | m | 0.99998 | 1e-05 | -2.0679e-05 | 5e-05 | m | 200 | mHz | 0.001 | 0.001 | 0.00099998 | 1e-08 | -2.0679e-08 | 5e-08 | 200mHz1m
0.5 | 1 | m | 1 | 1e-05 | 0.00013157 | 5e-05 | m | 500 | mHz | 0.001 | 0.001 | 0.001 | 1e-08 | 1.3157e-07 | 5e-08 | 500mHz1m
1 | 1 | m | 1.0002 | 1e-05 | -3.7714e-05 | 5e-05 | m | 1 | Hz | 0.001 | 0.001 | 0.0010002 | 1e-08 | -3.7714e-08 | 5e-08 | 1Hz1m
2 | 1 | m | 0.99996 | 1e-05 | 8.2644e-05 | 5e-05 | m | 2 | Hz | 0.001 | 0.001 | 0.00099996 | 1e-08 | 8.2644e-08 | 5e-08 | 2Hz1m
5 | 1 | m | 0.99993 | 1e-05 | -5.3103e-05 | 5e-05 | m | 5 | Hz | 0.001 | 0.001 | 0.00099993 | 1e-08 | -5.3103e-08 | 5e-08 | 5Hz1m
10 | 1 | m | 0.99987 | 1e-05 | -2.27e-05 | 5e-05 | m | 10 | Hz | 0.001 | 0.001 | 0.00099987 | 1e-08 | -2.27e-08 | 5e-08 | 10Hz1m
20 | 1 | m | 1 | 1e-05 | 0.00019142 | 5e-05 | m | 20 | Hz | 0.001 | 0.001 | 0.001 | 1e-08 | 1.9142e-07 | 5e-08 | 20Hz1m
50 | 1 | m | 1 | 1e-05 | 0.0003026 | 5e-05 | m | 50 | Hz | 0.001 | 0.001 | 0.001 | 1e-08 | 3.026e-07 | 5e-08 | 50Hz1m
100 | 1 | m | 1 | 1e-05 | 0.00097347 | 5e-05 | m | 100 | Hz | 0.001 | 0.001 | 0.001 | 1e-08 | 9.7347e-07 | 5e-08 | 100Hz1m
200 | 1 | m | 1 | 1e-05 | 0.0018839 | 5e-05 | m | 200 | Hz | 0.001 | 0.001 | 0.001 | 1e-08 | 1.8839e-06 | 5e-08 | 200Hz1m
500 | 1 | m | 1.0007 | 1.5811e-05 | 0.0045229 | 5e-05 | m | 500 | Hz | 0.001 | 0.001 | 0.0010007 | 1.5811e-08 | 4.5229e-06 | 5e-08 | 500Hz1m
1000 | 1 | m | 1.0017 | 4.4721e-05 | 0.0093571 | 0.0001 | m | 1 | kHz | 0.001 | 0.001 | 0.0010017 | 4.4721e-08 | 9.3571e-06 | 1e-07 | 1kHz1m
2000 | 1 | m | 1.0049 | 0.00012649 | 0.018571 | 0.0002 | m | 2 | kHz | 0.001 | 0.001 | 0.0010049 | 1.2649e-07 | 1.8571e-05 | 2e-07 | 2kHz1m
5000 | 1 | m | 1.0199 | 0.0005 | 0.046831 | 0.0005 | m | 5 | kHz | 0.001 | 0.001 | 0.0010199 | 5e-07 | 4.6831e-05 | 5e-07 | 5kHz1m
0.01 | 3 | m | 2.9999 | 3e-05 | 7.3301e-05 | 0.00015 | m | 10 | mHz | 0.003 | 0.001 | 0.0029999 | 3e-08 | 7.3301e-08 | 1.5e-07 | 10mHz3m
0.02 | 3 | m | 2.9997 | 3e-05 | 4.1886e-05 | 0.00015 | m | 20 | mHz | 0.003 | 0.001 | 0.0029997 | 3e-08 | 4.1886e-08 | 1.5e-07 | 20mHz3m
0.05 | 3 | m | 3.0001 | 3e-05 | 9.8063e-05 | 0.00015 | m | 50 | mHz | 0.003 | 0.001 | 0.0030001 | 3e-08 | 9.8063e-08 | 1.5e-07 | 50mHz3m
0.1 | 3 | m | 3 | 3e-05 | 0.00029191 | 0.00015 | m | 100 | mHz | 0.003 | 0.001 | 0.003 | 3e-08 | 2.9191e-07 | 1.5e-07 | 100mHz3m
0.2 | 3 | m | 2.9995 | 3e-05 | 0.00012602 | 0.00015 | m | 200 | mHz | 0.003 | 0.001 | 0.0029995 | 3e-08 | 1.2602e-07 | 1.5e-07 | 200mHz3m
0.5 | 3 | m | 3.0002 | 3e-05 | 0.00026309 | 0.00015 | m | 500 | mHz | 0.003 | 0.001 | 0.0030002 | 3e-08 | 2.6309e-07 | 1.5e-07 | 500mHz3m
1 | 3 | m | 3.0002 | 3e-05 | -7.9262e-05 | 0.00015 | m | 1 | Hz | 0.003 | 0.001 | 0.0030002 | 3e-08 | -7.9262e-08 | 1.5e-07 | 1Hz3m
2 | 3 | m | 3 | 3e-05 | -8.1116e-05 | 0.00015 | m | 2 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | -8.1116e-08 | 1.5e-07 | 2Hz3m
5 | 3 | m | 2.9998 | 3e-05 | 0.0005804 | 0.00015 | m | 5 | Hz | 0.003 | 0.001 | 0.0029998 | 3e-08 | 5.804e-07 | 1.5e-07 | 5Hz3m
10 | 3 | m | 3.0007 | 3e-05 | 0.00019934 | 0.00015 | m | 10 | Hz | 0.003 | 0.001 | 0.0030007 | 3e-08 | 1.9934e-07 | 1.5e-07 | 10Hz3m
20 | 3 | m | 3.0005 | 3e-05 | 0.00037537 | 0.00015 | m | 20 | Hz | 0.003 | 0.001 | 0.0030005 | 3e-08 | 3.7537e-07 | 1.5e-07 | 20Hz3m
50 | 3 | m | 3 | 3e-05 | 0.0017363 | 0.00015 | m | 50 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | 1.7363e-06 | 1.5e-07 | 50Hz3m
100 | 3 | m | 3.0002 | 3e-05 | 0.002848 | 0.00015 | m | 100 | Hz | 0.003 | 0.001 | 0.0030002 | 3e-08 | 2.848e-06 | 1.5e-07 | 100Hz3m
200 | 3 | m | 3.0011 | 3e-05 | 0.0055955 | 0.00015 | m | 200 | Hz | 0.003 | 0.001 | 0.0030011 | 3e-08 | 5.5955e-06 | 1.5e-07 | 200Hz3m
500 | 3 | m | 3.0017 | 4.7434e-05 | 0.013645 | 0.00015 | m | 500 | Hz | 0.003 | 0.001 | 0.0030017 | 4.7434e-08 | 1.3645e-05 | 1.5e-07 | 500Hz3m
1000 | 3 | m | 3.0054 | 0.00013416 | 0.027866 | 0.0003 | m | 1 | kHz | 0.003 | 0.001 | 0.0030054 | 1.3416e-07 | 2.7866e-05 | 3e-07 | 1kHz3m
2000 | 3 | m | 3.0147 | 0.00037947 | 0.056069 | 0.0006 | m | 2 | kHz | 0.003 | 0.001 | 0.0030147 | 3.7947e-07 | 5.6069e-05 | 6e-07 | 2kHz3m
5000 | 3 | m | 3.0601 | 0.0015 | 0.14009 | 0.0015 | m | 5 | kHz | 0.003 | 0.001 | 0.0030601 | 1.5e-06 | 0.00014009 | 1.5e-06 | 5kHz3m
0.01 | 10 | m | 10.0004 | 0.0001 | -5.5374e-05 | 0.0005 | m | 10 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | -5.5374e-08 | 5e-07 | 10mHz10m
0.02 | 10 | m | 10.0015 | 0.0001 | 0.0013286 | 0.0005 | m | 20 | mHz | 0.01 | 0.001 | 0.010001 | 1e-07 | 1.3286e-06 | 5e-07 | 20mHz10m
0.05 | 10 | m | 10.0007 | 0.0001 | -0.0011449 | 0.0005 | m | 50 | mHz | 0.01 | 0.001 | 0.010001 | 1e-07 | -1.1449e-06 | 5e-07 | 50mHz10m
0.1 | 10 | m | 10.0004 | 0.0001 | -0.00016302 | 0.0005 | m | 100 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | -1.6302e-07 | 5e-07 | 100mHz10m
0.2 | 10 | m | 9.9996 | 0.0001 | 0.00074735 | 0.0005 | m | 200 | mHz | 0.01 | 0.001 | 0.0099996 | 1e-07 | 7.4735e-07 | 5e-07 | 200mHz10m
0.5 | 10 | m | 10.0004 | 0.0001 | -0.00045986 | 0.0005 | m | 500 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | -4.5986e-07 | 5e-07 | 500mHz10m
1 | 10 | m | 9.999 | 0.0001 | -0.00049945 | 0.0005 | m | 1 | Hz | 0.01 | 0.001 | 0.009999 | 1e-07 | -4.9945e-07 | 5e-07 | 1Hz10m
2 | 10 | m | 9.998 | 0.0001 | 0.00094431 | 0.0005 | m | 2 | Hz | 0.01 | 0.001 | 0.009998 | 1e-07 | 9.4431e-07 | 5e-07 | 2Hz10m
5 | 10 | m | 10.0002 | 0.0001 | -0.00099104 | 0.0005 | m | 5 | Hz | 0.01 | 0.001 | 0.01 | 1e-07 | -9.9104e-07 | 5e-07 | 5Hz10m
10 | 10 | m | 10.001 | 0.0001 | 8.9632e-05 | 0.0005 | m | 10 | Hz | 0.01 | 0.001 | 0.010001 | 1e-07 | 8.9632e-08 | 5e-07 | 10Hz10m
20 | 10 | m | 9.9998 | 0.0001 | 0.0016297 | 0.0005 | m | 20 | Hz | 0.01 | 0.001 | 0.0099998 | 1e-07 | 1.6297e-06 | 5e-07 | 20Hz10m
50 | 10 | m | 9.9993 | 0.0001 | 0.0044984 | 0.0005 | m | 50 | Hz | 0.01 | 0.001 | 0.0099993 | 1e-07 | 4.4984e-06 | 5e-07 | 50Hz10m
100 | 10 | m | 10.0007 | 0.0001 | 0.0095713 | 0.0005 | m | 100 | Hz | 0.01 | 0.001 | 0.010001 | 1e-07 | 9.5713e-06 | 5e-07 | 100Hz10m
200 | 10 | m | 10.0005 | 0.0001 | 0.017729 | 0.0005 | m | 200 | Hz | 0.01 | 0.001 | 0.01 | 1e-07 | 1.7729e-05 | 5e-07 | 200Hz10m
500 | 10 | m | 10.0059 | 0.00015811 | 0.045356 | 0.0005 | m | 500 | Hz | 0.01 | 0.001 | 0.010006 | 1.5811e-07 | 4.5356e-05 | 5e-07 | 500Hz10m
1000 | 10 | m | 10.0184 | 0.00044721 | 0.093275 | 0.001 | m | 1 | kHz | 0.01 | 0.001 | 0.010018 | 4.4721e-07 | 9.3275e-05 | 1e-06 | 1kHz10m
2000 | 10 | m | 10.0507 | 0.0012649 | 0.18716 | 0.002 | m | 2 | kHz | 0.01 | 0.001 | 0.010051 | 1.2649e-06 | 0.00018716 | 2e-06 | 2kHz10m
5000 | 10 | m | 10.1991 | 0.005 | 0.46839 | 0.005 | m | 5 | kHz | 0.01 | 0.001 | 0.010199 | 5e-06 | 0.00046839 | 5e-06 | 5kHz10m
0.01 | 100 | m | 99.9991 | 0.001 | 0.00078026 | 0.005 | m | 10 | mHz | 0.1 | 0.001 | 0.099999 | 1e-06 | 7.8026e-07 | 5e-06 | 10mHz100m
0.02 | 100 | m | 99.9924 | 0.001 | -0.0045184 | 0.005 | m | 20 | mHz | 0.1 | 0.001 | 0.099992 | 1e-06 | -4.5184e-06 | 5e-06 | 20mHz100m
0.05 | 100 | m | 99.9945 | 0.001 | -0.018542 | 0.005 | m | 50 | mHz | 0.1 | 0.001 | 0.099995 | 1e-06 | -1.8542e-05 | 5e-06 | 50mHz100m
0.1 | 100 | m | 100.0069 | 0.001 | -0.0036364 | 0.005 | m | 100 | mHz | 0.1 | 0.001 | 0.10001 | 1e-06 | -3.6364e-06 | 5e-06 | 100mHz100m
0.2 | 100 | m | 99.9829 | 0.001 | -0.01368 | 0.005 | m | 200 | mHz | 0.1 | 0.001 | 0.099983 | 1e-06 | -1.368e-05 | 5e-06 | 200mHz100m
0.5 | 100 | m | 99.9936 | 0.001 | -0.0078555 | 0.005 | m | 500 | mHz | 0.1 | 0.001 | 0.099994 | 1e-06 | -7.8555e-06 | 5e-06 | 500mHz100m
1 | 100 | m | 99.9936 | 0.001 | -0.0061224 | 0.005 | m | 1 | Hz | 0.1 | 0.001 | 0.099994 | 1e-06 | -6.1224e-06 | 5e-06 | 1Hz100m
2 | 100 | m | 100.0221 | 0.001 | 0.0096162 | 0.005 | m | 2 | Hz | 0.1 | 0.001 | 0.10002 | 1e-06 | 9.6162e-06 | 5e-06 | 2Hz100m
5 | 100 | m | 99.9909 | 0.001 | 0.0098483 | 0.005 | m | 5 | Hz | 0.1 | 0.001 | 0.099991 | 1e-06 | 9.8483e-06 | 5e-06 | 5Hz100m
10 | 100 | m | 99.9772 | 0.001 | -0.0017626 | 0.005 | m | 10 | Hz | 0.1 | 0.001 | 0.099977 | 1e-06 | -1.7626e-06 | 5e-06 | 10Hz100m
20 | 100 | m | 100.017 | 0.001 | 0.011488 | 0.005 | m | 20 | Hz | 0.1 | 0.001 | 0.10002 | 1e-06 | 1.1488e-05 | 5e-06 | 20Hz100m
50 | 100 | m | 100.0007 | 0.001 | 0.052501 | 0.005 | m | 50 | Hz | 0.1 | 0.001 | 0.1 | 1e-06 | 5.2501e-05 | 5e-06 | 50Hz100m
100 | 100 | m | 100.0163 | 0.001 | 0.08934 | 0.005 | m | 100 | Hz | 0.1 | 0.001 | 0.10002 | 1e-06 | 8.934e-05 | 5e-06 | 100Hz100m
200 | 100 | m | 100.0113 | 0.001 | 0.189 | 0.005 | m | 200 | Hz | 0.1 | 0.001 | 0.10001 | 1e-06 | 0.000189 | 5e-06 | 200Hz100m
500 | 100 | m | 100.0381 | 0.0015811 | 0.45737 | 0.005 | m | 500 | Hz | 0.1 | 0.001 | 0.10004 | 1.5811e-06 | 0.00045737 | 5e-06 | 500Hz100m
1000 | 100 | m | 100.1727 | 0.0044721 | 0.91905 | 0.01 | m | 1 | kHz | 0.1 | 0.001 | 0.10017 | 4.4721e-06 | 0.00091905 | 1e-05 | 1kHz100m
2000 | 100 | m | 100.497 | 0.012649 | 1.8469 | 0.02 | m | 2 | kHz | 0.1 | 0.001 | 0.1005 | 1.2649e-05 | 0.0018469 | 2e-05 | 2kHz100m
5000 | 100 | m | 101.9889 | 0.05 | 4.6575 | 0.05 | m | 5 | kHz | 0.1 | 0.001 | 0.10199 | 5e-05 | 0.0046575 | 5e-05 | 5kHz100m

## Ref Z list
Ref. Z list
¦Z¦ |  | ¦Z¦ | Z tag
Ohm |  | Ohm | -
0 | m | 0 | 0m
1 | m | 0.001 | 1m
3 | m | 0.003 | 3m
10 | m | 0.01 | 10m
100 | m | 0.1 | 100m

## Ranges
Range list
Range |  | Range Z | Min ¦Z¦ |  | Min Z | Max ¦Z¦ |  | Max Z
Ohm |  | Ohm | Ohm |  | Ohm | Ohm |  | Ohm
1 | m | 0.001 | 0 | m | 0 | 1 | m | 0.001
3 | m | 0.003 | 1 | m | 0.001 | 3 | m | 0.003
10 | m | 0.01 | 3 | m | 0.003 | 10 | m | 0.01
100 | m | 0.1 | 10 | m | 0.01 | 100 | m | 0.1