## Measurement
This sheet uses correction data to fix bridge errors.
|  |  | Measured Z |  |  |  |  | Measured residual SHORT Z: |  |  |  |  | Corrected Z (results) | [5 empty] | Validation: ref data |  |  |  |  | Deviation of XLS calculated from Octave calculated: |  |  |  |  | Deviation with no correction: | [6 empty] | Measured Z (basic units) | [27 empty] | Interpolated correction |  |  |  | corrected Z | [11 empty] | Interpolated correction (gain-phase) |  |  |  | Corrected Z (gain-phase method)
Range |  | f | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Fixed Rs | Fixed U(Rs) | Fixed Xs | Fixed U(Xs) |  | Status | Valid Rs |  | Valid Xs |  |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | f |  | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | ua(¦Z¦) | phi | ua(phi) | lower Z ID | lower Z tag | higher Z tag | lower Z | higher Z | lower tag | higher tag | lower ID | higher ID | low check | low cor(Rs) | low U(cor(Rs)) | low cor(Xs) | low U(cor(Xs)) | high check | high cor(Rs) | high U(cor(Rs)) | high cor(Xs) | high U(cor(Xs)) | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | Rs | U(Rs) | Xs | U(Xs) | low cor(gain) |  | low cor(phi) |  | hi cor(gain) |  | hi cor(phi) |  | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi)) | ¦Z¦ | U(¦Z¦) | phi | U(phi) | Rs | U(Rs) | Xs | U(Xs)
Ohm |  | Hz | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm |  | - | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | - |  | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | - | Ohm | Ohm | - | - | - | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - |  | rad |  | - |  | rad |  | - | - | rad | rad | Ohm | Ohm | rad | rad | Ohm | Ohm | Ohm | Ohm
100 | m | 200 | -22.8988 | -6.4492e-05 | -2.1366 | 0.0011449 | m | -0.00036822 | 0.00041428 | 0.001378 | 0.00058777 | m | -22.9 | 0.0010229 | -2.14 | 0.0058068 | m | OK | -22.9 |  | -2.14 |  | m | 1.0849e-07 | 0.0010229 | 2.2799e-07 | 0.0058068 | m | 0.0016056 |  | 0.0019954 |  | m | 200 | Hz | 0.001 | -0.022898 | 4.1927e-07 | -0.002138 | 1.287e-06 | 0.022998 | 4.3426e-07 | -3.0485 | 5.5744e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | 9.2281e-07 | 0.0017247 | 1.0005e-06 | 0.0025224 | 200Hz100m100m | -9.6164e-07 | 0.002522 | -3.9289e-06 | 0.00056323 | 6.5065e-07 | 0.002522 | 2.8863e-07 | 0.0022394 | -0.022898 | 0.002522 | -0.0021377 | 0.0022394 | 1.0001 | 1.8385e-06 | 9.9995e-05 | 0.00025756 | 0.99999 | 2.8219e-06 | -3.9274e-05 | 5.098e-05 | 1.0001 | 2.8219e-06 | 7.9882e-05 | 0.00022772 | 0.023 | 8.7094e-07 | -3.0484 | 0.00025355 | -0.0229 | 1.0229e-06 | -0.00214 | 5.8068e-06
100 | m | 5 | -94.401 | 0.00093356 | -1.8589 | 0.0011956 | m | -0.0018528 | 0.0013446 | 0.0018253 | 0.001696 | m | -94.4 | 0.003328 | -1.87 | 0.00839 | m | OK | -94.4 |  | -1.87 |  | m | 6.1745e-07 | 0.003328 | -7.7477e-09 | 0.00839 | m | 0.00086581 |  | 0.0092716 |  | m | 5 | Hz | 0.001 | -0.094399 | 1.6369e-06 | -0.0018607 | 2.075e-06 | 0.094417 | 1.6371e-06 | -3.1219 | 2.1975e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 5Hz100m10m | 5Hz100m100m | 117 | 135 | 5Hz100m10m | -5.3364e-07 | 0.00063677 | 9.9995e-07 | 0.0032479 | 5Hz100m100m | 1.5396e-06 | 0.0053574 | 9.785e-06 | 0.0031523 | 1.411e-06 | 0.0053574 | 9.2401e-06 | 0.0031582 | -0.094398 | 0.0053574 | -0.0018515 | 0.0031582 | 0.99995 | 2.8566e-06 | 0.0001 | 0.00034697 | 1 | 6.1029e-06 | 9.7865e-05 | 5.9409e-05 | 1 | 6.1029e-06 | 9.7998e-05 | 7.7246e-05 | 0.094419 | 3.3245e-06 | -3.1218 | 8.8874e-05 | -0.0944 | 3.328e-06 | -0.00187 | 8.39e-06
10 | m | 0.1 | -1.4116 | 0.000996 | 9.3393 | 0.0015415 | m | -0.0018869 | 0.00044104 | -0.0013516 | 0.0018031 | m | -1.41 | 0.0043535 | 9.34 | 0.004695 | m | OK | -1.41 |  | 9.34 |  | m | 3.0021e-08 | 0.0043535 | -1.766e-08 | 0.004695 | m | 0.00031421 |  | 0.00064232 |  | m | 100 | mHz | 0.001 | -0.0014097 | 1.0893e-06 | 0.0093406 | 2.3722e-06 | 0.0094464 | 2.3513e-06 | 1.7206 | 0.00012002 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | -3.5068e-08 | 0.0018993 | 3e-07 | 0.0016416 | 100mHz10m10m | -6.6623e-07 | 0.0028136 | 3.8148e-07 | 0.0036452 | -6.1632e-07 | 0.0028136 | 3.7504e-07 | 0.0034867 | -0.0014103 | 0.0028136 | 0.009341 | 0.0034867 | 0.99999 | 2.916e-06 | 0.0001 | 0.00066159 | 0.99993 | 2.9318e-06 | 3.8145e-05 | 0.0003693 | 0.99994 | 2.9318e-06 | 4.3037e-05 | 0.00039242 | 0.0094458 | 4.7027e-06 | 1.7206 | 0.00046001 | -0.00141 | 4.3535e-06 | 0.00934 | 4.695e-06
100 | m | 500 | 0.11632 | 0.0013757 | -16.6987 | 0.0010483 | m | 0.00096358 | 0.0012041 | 0.00060665 | 0.00046052 | m | 0.117 | 0.0048293 | -16.7 | 0.0022917 | m | OK | 0.117 |  | -16.7 |  | m | -1.1671e-07 | 0.0048293 | 4.8564e-07 | 0.0022917 | m | -0.0016398 |  | 0.00074206 |  | m | 500 | Hz | 0.001 | 0.00011536 | 1.8282e-06 | -0.016699 | 1.145e-06 | 0.0167 | 1.1451e-06 | -1.5639 | 0.00010948 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | 5.4485e-07 | 0.0013148 | 1.002e-06 | 6.2609e-05 | 500Hz100m100m | -7.2572e-06 | 0.0027372 | 7.1503e-06 | 0.0012973 | -3.5943e-08 | 0.0027372 | 1.4597e-06 | 0.00015452 | 0.00011532 | 0.0027372 | -0.016698 | 0.00015454 | 1.0001 | 1.9945e-06 | 9.9996e-05 | 0.00019995 | 0.99993 | 4.7336e-06 | 7.1514e-05 | 5.1597e-05 | 1 | 4.7336e-06 | 9.7875e-05 | 0.0001889 | 0.0167 | 2.2915e-06 | -1.5638 | 0.00028918 | 0.000117 | 4.8293e-06 | -0.0167 | 2.2917e-06
3 | m | 2000 | -1.1019 | 0.00083987 | -0.56538 | 8.176e-05 | m | -0.001738 | 0.00029151 | -0.0014446 | 4.7227e-05 | m | -1.1 | 0.0014982 | -0.564 | 0.0012239 | m | OK | -1.1 |  | -0.564 |  | m | -4.9057e-07 | 0.0014982 | 3.6721e-07 | 0.0012239 | m | -0.00012737 |  | 6.1224e-05 |  | m | 2 | kHz | 0.001 | -0.0011001 | 8.8902e-07 | -0.00056394 | 9.4419e-08 | 0.0012362 | 7.923e-07 | -2.6679 | 0.00033501 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | -7.4486e-08 | 0.00079912 | 1.0022e-07 | 0.00047925 | 2kHz3m3m | -1.0953e-07 | 0.0028076 | 7.0263e-08 | 0.00084968 | -7.8625e-08 | 0.0028076 | 9.668e-08 | 0.00052301 | -0.0011002 | 0.0028076 | -0.00056384 | 0.00052301 | 0.99993 | 9.0744e-07 | 9.9639e-05 | 0.00052319 | 0.99996 | 4.3725e-06 | 2.3388e-05 | 0.0011519 | 0.99993 | 4.3725e-06 | 9.0632e-05 | 0.00059746 | 0.0012362 | 1.5846e-06 | -2.6678 | 0.00089771 | -0.0011 | 1.4982e-06 | -0.000564 | 1.2239e-06
1 | m | 0.5 | -0.0070351 | 0.0001645 | -0.45033 | 0.0014186 | m | 0.001324 | 0.00063473 | 0.0016638 | 0.00046348 | m | -0.008383 | 0.0014932 | -0.452 | 0.0029839 | m | OK | -0.00837 |  | -0.452 |  | m | -1.3021e-05 | 0.0014932 | 2.4198e-07 | 0.0029839 | m | 1.0944e-05 |  | 1.0579e-05 |  | m | 500 | mHz | 0.001 | -8.3591e-06 | 6.557e-07 | -0.00045199 | 1.4924e-06 | 0.00045207 | 1.4922e-06 | -1.5893 | 0.0014515 | 1 | 0m | 1m | 0 | 0.001 | 500mHz1m0m | 500mHz1m1m | 6 | 24 | 500mHz1m0m | 0 | 0.00084516 | 0 | 0.0036068 | 500mHz1m1m | 5.2745e-08 | 0.0020222 | -5.2563e-08 | 0.0013887 | 2.3844e-08 | 0.0020222 | -5.2563e-08 | 0.002604 | -8.3352e-06 | 0.0020222 | -0.00045204 | 0.002604 | 1 | 3.7045e-06 | -5.2578e-05 | 0.0015719 | 1.0001 | 3.317e-06 | -5.2578e-05 | 0.0015719 | 1 | 3.317e-06 | -5.2578e-05 | 0.0015719 | 0.00045208 | 2.9843e-06 | -1.5893 | 0.0033012 | -8.383e-06 | 1.4932e-06 | -0.000452 | 2.9839e-06
3 | m | 5 | -0.64286 | 0.001226 | 0.83819 | 0.0015125 | m | -0.00091284 | 0.00054589 | -0.0019092 | 0.00069897 | m | -0.642 | 0.0039511 | 0.84 | 0.0036248 | m | OK | -0.642 |  | 0.84 |  | m | -2.7613e-10 | 0.0039511 | 2.4486e-09 | 0.0036248 | m | 5.205e-05 |  | 9.9941e-05 |  | m | 5 | Hz | 0.001 | -0.00064195 | 1.342e-06 | 0.0008401 | 1.6662e-06 | 0.0010573 | 1.5546e-06 | 2.2233 | 0.0013902 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | -4.9337e-08 | 0.00079014 | 1.0001e-07 | 0.0022765 | 5Hz3m3m | 2.8449e-07 | 0.0026935 | -6.5054e-08 | 0.0018253 | -3.9775e-08 | 0.0026935 | 9.5279e-08 | 0.0022635 | -0.00064199 | 0.0026935 | 0.0008402 | 0.0022635 | 0.99995 | 2.7916e-06 | 9.9998e-05 | 0.0031263 | 1.0001 | 3.03e-06 | -2.1698e-05 | 0.00068935 | 0.99995 | 3.03e-06 | 9.6512e-05 | 0.0030565 | 0.0010572 | 3.1092e-06 | 2.2234 | 0.004132 | -0.000642 | 3.9511e-06 | 0.00084 | 3.6248e-06
10 | m | 10 | 3.4199 | 0.0014112 | -3.3187 | 0.00034263 | m | 2.2137e-05 | 2.4425e-05 | 0.001635 | 5.2517e-05 | m | 3.42 | 0.003688 | -3.32 | 0.0037641 | m | OK | 3.42 |  | -3.32 |  | m | -3.6126e-09 | 0.003688 | -8.0463e-09 | 0.0037641 | m | -0.00011579 |  | -0.00037748 |  | m | 10 | Hz | 0.001 | 0.0034199 | 1.4114e-06 | -0.0033204 | 3.4663e-07 | 0.0047666 | 1.0411e-06 | -0.77064 | 0.00021276 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -1.3483e-07 | 0.0019237 | 3.0003e-07 | 0.0032939 | 10Hz10m10m | -1.6362e-07 | 0.0015051 | -4.0397e-08 | 0.00018818 | -1.421e-07 | 0.0015051 | 2.1411e-07 | 0.0025101 | 0.0034197 | 0.0015051 | -0.0033202 | 0.0025101 | 0.99996 | 3.1742e-06 | 9.9997e-05 | 0.0012036 | 0.99998 | 3.4234e-06 | -4.0363e-06 | 8.8483e-05 | 0.99996 | 3.4234e-06 | 7.3742e-05 | 0.00092218 | 0.0047664 | 2.0822e-06 | -0.77056 | 0.0010156 | 0.00342 | 3.688e-06 | -0.00332 | 3.7641e-06
3 | m | 50 | 1.5981 | 0.00023616 | 1.741 | 0.00051156 | m | -0.0018716 | 0.00036532 | 0.0011054 | 0.0017796 | m | 1.6 | 0.0050543 | 1.74 | 0.0047751 | m | OK | 1.6 |  | 1.74 |  | m | 2.7628e-08 | 0.0050543 | 6.2169e-09 | 0.0047751 | m | -5.3223e-05 |  | -6.3379e-05 |  | m | 50 | Hz | 0.001 | 0.0015999 | 4.3501e-07 | 0.0017399 | 1.8517e-06 | 0.0023637 | 1.3944e-06 | 0.82729 | 0.00054727 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | -5.3917e-08 | 0.0013231 | 9.9995e-08 | 0.0037478 | 50Hz3m3m | 2.2939e-07 | 0.0020975 | -1.3301e-07 | 0.0033185 | 1.3926e-07 | 0.0020975 | -5.8883e-08 | 0.0034551 | 0.0016001 | 0.0020975 | 0.0017399 | 0.0034551 | 0.99995 | 3.1478e-06 | 9.9994e-05 | 0.0044451 | 1.0001 | 2.5574e-06 | -4.4355e-05 | 0.0015371 | 1 | 2.5574e-06 | 1.5677e-06 | 0.0024622 | 0.0023638 | 2.7889e-06 | 0.82729 | 0.0026946 | 0.0016 | 5.0543e-06 | 0.00174 | 4.7751e-06
10 | m | 2000 | 9.14 | 0.00033063 | 1.8696 | 0.0010811 | m | -0.00089024 | 0.0016796 | 0.00017383 | 0.0014445 | m | 9.14 | 0.0034803 | 1.87 | 0.0044571 | m | OK | 9.14 |  | 1.87 |  | m | -5.4231e-06 | 0.0034803 | -3.329e-06 | 0.0044571 | m | 0.00092803 |  | -0.00060618 |  | m | 2 | kHz | 0.001 | 0.0091409 | 1.7118e-06 | 0.0018694 | 1.8043e-06 | 0.0093301 | 1.7156e-06 | 0.20173 | 0.00019299 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.2113e-07 | 0.0031596 | 3.0201e-07 | 0.00059968 | 2kHz10m10m | -9.1201e-07 | 0.0021293 | 8.1914e-07 | 0.00049468 | -8.3633e-07 | 0.0021293 | 7.6965e-07 | 0.00050473 | 0.0091401 | 0.0021293 | 0.0018702 | 0.00050474 | 0.99996 | 3.1912e-06 | 9.9926e-05 | 0.00040135 | 0.99991 | 2.7666e-06 | 8.1602e-05 | 0.00027639 | 0.99991 | 2.7666e-06 | 8.3356e-05 | 0.00028835 | 0.0093293 | 3.4313e-06 | 0.20181 | 0.0004818 | 0.00914 | 3.4803e-06 | 0.00187 | 4.4571e-06
3 | m | 1000 | -2.6296 | 2.8198e-05 | -0.95362 | 0.0016599 | m | 0.00031811 | 0.0017462 | 0.00033102 | 0.0018542 | m | -2.63 | 0.0040436 | -0.954 | 0.0058394 | m | OK | -2.63 |  | -0.954 |  | m | -6.3761e-07 | 0.0040436 | 1.3125e-06 | 0.0058394 | m | 0.00010062 |  | 5.2335e-05 |  | m | 1 | kHz | 0.001 | -0.0026299 | 1.7464e-06 | -0.00095395 | 2.4887e-06 | 0.0027976 | 1.8481e-06 | -2.7936 | 0.00086293 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 3.0166e-08 | 0.00050233 | 1.0023e-07 | 0.0010421 | 1kHz3m3m | 1.2457e-07 | 0.0038805 | -1.7428e-08 | 0.003317 | 1.1501e-07 | 0.0038805 | -5.519e-09 | 0.0030867 | -0.0026298 | 0.0038805 | -0.00095395 | 0.0030867 | 1 | 1.9109e-06 | 9.9824e-05 | 0.0013661 | 1 | 5.4207e-06 | -5.8979e-06 | 0.001306 | 1 | 5.4207e-06 | 4.8028e-06 | 0.0013121 | 0.0027977 | 3.6962e-06 | -2.7936 | 0.002168 | -0.00263 | 4.0436e-06 | -0.000954 | 5.8394e-06
3 | m | 50 | -1.1697 | 0.00053572 | -2.1107 | 0.00093724 | m | 0.00027085 | 0.00079584 | -0.00073307 | 0.00068642 | m | -1.17 | 0.0054549 | -2.11 | 0.0035508 | m | OK | -1.17 |  | -2.11 |  | m | -2.5783e-08 | 0.0054549 | -1.2467e-08 | 0.0035508 | m | 4.8766e-05 |  | 7.8236e-05 |  | m | 50 | Hz | 0.001 | -0.00117 | 9.5935e-07 | -0.0021099 | 1.1617e-06 | 0.0024126 | 1.1174e-06 | -2.0771 | 0.00041888 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | -5.3917e-08 | 0.0013231 | 9.9995e-08 | 0.0037478 | 50Hz3m3m | 2.2939e-07 | 0.0020975 | -1.3301e-07 | 0.0033185 | 1.4618e-07 | 0.0020975 | -6.4575e-08 | 0.0034446 | -0.0011698 | 0.0020975 | -0.00211 | 0.0034446 | 0.99995 | 3.1478e-06 | 9.9994e-05 | 0.0044451 | 1.0001 | 2.5574e-06 | -4.4355e-05 | 0.0015371 | 1 | 2.5574e-06 | -1.9583e-06 | 0.0023912 | 0.0024127 | 2.2349e-06 | -2.0771 | 0.0025337 | -0.00117 | 5.4549e-06 | -0.00211 | 3.5508e-06
3 | m | 1000 | 2.3396 | 0.00021334 | 1.3089 | 0.0019999 | m | -0.00027905 | 0.0014306 | -0.001048 | 0.0018047 | m | 2.34 | 0.0043412 | 1.31 | 0.0055677 | m | OK | 2.34 |  | 1.31 |  | m | 7.7221e-07 | 0.0043412 | -1.0941e-06 | 0.0055677 | m | -7.7576e-05 |  | -7.8551e-05 |  | m | 1 | kHz | 0.001 | 0.0023399 | 1.4465e-06 | 0.0013099 | 2.6938e-06 | 0.0026816 | 1.8233e-06 | 0.51035 | 0.00091529 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 3.0166e-08 | 0.00050233 | 1.0023e-07 | 0.0010421 | 1kHz3m3m | 1.2457e-07 | 0.0038805 | -1.7428e-08 | 0.003317 | 1.0954e-07 | 0.0038805 | 1.3015e-09 | 0.0029548 | 0.00234 | 0.0038805 | 0.0013099 | 0.0029548 | 1 | 1.9109e-06 | 9.9824e-05 | 0.0013661 | 1 | 5.4207e-06 | -5.8979e-06 | 0.001306 | 1 | 5.4207e-06 | 1.0931e-05 | 0.0013156 | 0.0026817 | 3.6467e-06 | 0.51036 | 0.0022543 | 0.00234 | 4.3412e-06 | 0.00131 | 5.5677e-06
3 | m | 5 | -0.057979 | 0.00011999 | -1.1699 | 0.00015151 | m | 0.00032805 | 0.0001764 | 9.1281e-05 | 0.00028223 | m | -0.0582 | 0.00344 | -1.17 | 0.00066193 | m | OK | -0.0582 |  | -1.17 |  | m | -2.2485e-10 | 0.00344 | -2.046e-09 | 0.00066193 | m | -0.00010695 |  | -3.805e-05 |  | m | 5 | Hz | 0.001 | -5.8307e-05 | 2.1335e-07 | -0.00117 | 3.2033e-07 | 0.0011715 | 3.2011e-07 | -1.6206 | 0.0001824 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | -4.9337e-08 | 0.00079014 | 1.0001e-07 | 0.0022765 | 5Hz3m3m | 2.8449e-07 | 0.0026935 | -6.5054e-08 | 0.0018253 | -2.0714e-08 | 0.0026935 | 8.5854e-08 | 0.0022378 | -5.8328e-05 | 0.0026935 | -0.00117 | 0.0022378 | 0.99995 | 2.7916e-06 | 9.9998e-05 | 0.0031263 | 1.0001 | 3.03e-06 | -2.1698e-05 | 0.00068935 | 0.99996 | 3.03e-06 | 8.9563e-05 | 0.0029173 | 0.0011714 | 6.4023e-07 | -1.6205 | 0.0029401 | -5.82e-05 | 3.44e-06 | -0.00117 | 6.6193e-07
3 | m | 1000 | -1.4303 | 0.0011486 | -0.92339 | 0.0017898 | m | -0.0002963 | 0.00025867 | -0.0015081 | 0.0018794 | m | -1.43 | 0.0039948 | -0.922 | 0.0046655 | m | OK | -1.43 |  | -0.922 |  | m | -3.3208e-07 | 0.0039948 | 4.1319e-07 | 0.0046655 | m | -9.2989e-06 |  | 0.00012166 |  | m | 1 | kHz | 0.001 | -0.00143 | 1.1774e-06 | -0.00092188 | 2.5953e-06 | 0.0017014 | 1.7195e-06 | -2.569 | 0.0013358 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 3.0166e-08 | 0.00050233 | 1.0023e-07 | 0.0010421 | 1kHz3m3m | 1.2457e-07 | 0.0038805 | -1.7428e-08 | 0.003317 | 6.3273e-08 | 0.0038805 | 5.8966e-08 | 0.0018399 | -0.0014299 | 0.0038805 | -0.00092182 | 0.0018399 | 1 | 1.9109e-06 | 9.9824e-05 | 0.0013661 | 1 | 5.4207e-06 | -5.8979e-06 | 0.001306 | 1 | 5.4207e-06 | 6.2747e-05 | 0.001345 | 0.0017015 | 3.439e-06 | -2.5689 | 0.002991 | -0.00143 | 3.9948e-06 | -0.000922 | 4.6655e-06
10 | m | 5 | -3.3099 | 0.00031152 | -9.1113 | -1.4902e-05 | m | 4.9548e-05 | 0.0013847 | -0.0011938 | 0.00069093 | m | -3.31 | 0.0044732 | -9.11 | 0.0022188 | m | OK | -3.31 |  | -9.11 |  | m | 1.2748e-07 | 0.0044732 | -4.7442e-08 | 0.0022188 | m | 1.6142e-05 |  | -0.00010495 |  | m | 5 | Hz | 0.001 | -0.00331 | 1.4193e-06 | -0.0091101 | 6.9109e-07 | 0.0096928 | 8.1044e-07 | -1.9193 | 0.00013976 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -2.6382e-08 | 0.00058565 | 2.9998e-07 | 0.00064819 | 5Hz10m10m | -9.645e-08 | 0.003922 | -1.008e-07 | 0.0028868 | -9.3375e-08 | 0.003922 | -8.3207e-08 | 0.0027886 | -0.0033101 | 0.003922 | -0.0091102 | 0.0027886 | 0.99999 | 1.3944e-06 | 0.0001 | 0.00029063 | 0.99999 | 3.9336e-06 | -1.008e-05 | 0.00040408 | 0.99999 | 3.9336e-06 | -5.249e-06 | 0.0003991 | 0.0096927 | 1.6213e-06 | -1.9193 | 0.00048725 | -0.00331 | 4.4732e-06 | -0.00911 | 2.2188e-06
10 | m | 500 | 3.9883 | 0.0013074 | 0.88378 | 0.0019045 | m | -0.0018237 | 0.0017516 | 0.0010831 | 0.0013018 | m | 3.99 | 0.004492 | 0.883 | 0.0062408 | m | OK | 3.99 |  | 0.883 |  | m | 5.8804e-08 | 0.004492 | -4.541e-07 | 0.0062408 | m | 7.8719e-05 |  | -0.00030619 |  | m | 500 | Hz | 0.001 | 0.0039901 | 2.1857e-06 | 0.00088269 | 2.3069e-06 | 0.0040865 | 2.1915e-06 | 0.21772 | 0.00056316 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | -2.1582e-09 | 0.0010202 | 3.0014e-07 | 0.0021028 | 500Hz10m10m | -1.3647e-07 | 0.0029467 | -4.6502e-07 | 0.001364 | -2.3006e-08 | 0.0029467 | 1.8137e-07 | 0.0019881 | 0.0039901 | 0.0029467 | 0.00088288 | 0.0019882 | 1 | 2.0742e-06 | 9.9922e-05 | 0.0011846 | 0.99999 | 4.3495e-06 | -4.6419e-05 | 0.00037546 | 1 | 4.3495e-06 | 7.7206e-05 | 0.001059 | 0.0040865 | 4.3831e-06 | 0.21779 | 0.001546 | 0.00399 | 4.492e-06 | 0.000883 | 6.2408e-06
1 | m | 2 | -0.32484 | 0.0011537 | -0.11835 | 0.0018214 | m | -0.00084396 | 0.00055078 | -0.0013449 | 0.00087447 | m | -0.324 | 0.0029547 | -0.117 | 0.0039776 | m | OK | -0.324 |  | -0.117 |  | m | -1.5246e-06 | 0.0029547 | 4.22e-06 | 0.0039776 | m | 6.1152e-06 |  | -2.9825e-07 |  | m | 2 | Hz | 0.001 | -0.00032399 | 1.2784e-06 | -0.000117 | 2.0205e-06 | 0.00034447 | 1.3845e-06 | -2.795 | 0.005659 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.00082151 | 0 | 0.00085633 | 2Hz1m1m | 4.7626e-08 | 0.0021182 | -1.9869e-08 | 0.0034282 | 1.6406e-08 | 0.0021182 | -1.9869e-08 | 0.0017422 | -0.00032398 | 0.0021182 | -0.00011702 | 0.0017423 | 1 | 1.1867e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 0.00034448 | 2.769e-06 | -2.7951 | 0.011928 | -0.000324 | 2.9547e-06 | -0.000117 | 3.9776e-06
100 | m | 0.02 | -62.4969 | 0.00024271 | 15.9989 | 0.0012879 | m | 0.00087687 | 0.00089579 | -0.0013904 | 0.0010434 | m | -62.5 | 0.0032511 | 16 | 0.010264 | m | OK | -62.5 |  | 16 |  | m | -4.5126e-07 | 0.0032511 | -2.723e-07 | 0.010264 | m | 0.0022152 |  | 0.00030939 |  | m | 20 | mHz | 0.001 | -0.062498 | 9.2808e-07 | 0.016 | 1.6575e-06 | 0.064513 | 9.8861e-07 | 2.891 | 2.5144e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 20mHz100m10m | 20mHz100m100m | 110 | 128 | 20mHz100m10m | 8.7605e-07 | 0.0023787 | 9.9977e-07 | 0.00033772 | 20mHz100m100m | -4.0604e-07 | 0.0017536 | -4.3356e-06 | 0.0028984 | 9.948e-08 | 0.0017536 | -2.2319e-06 | 0.0018887 | -0.062498 | 0.0017536 | 0.015998 | 0.0018887 | 1.0001 | 2.8399e-06 | 0.0001 | 0.00030124 | 1 | 1.7835e-06 | -4.3359e-05 | 6.1683e-05 | 1 | 1.7835e-06 | 1.3167e-05 | 0.00015614 | 0.064516 | 1.9806e-06 | 2.891 | 0.00016404 | -0.0625 | 3.2511e-06 | 0.016 | 1.0264e-05
3 | m | 10 | 1.6815 | 0.0013644 | 1.4389 | 0.0010992 | m | 0.0015254 | 0.0016807 | -0.0010931 | 0.0017309 | m | 1.68 | 0.0053013 | 1.44 | 0.0056372 | m | OK | 1.68 |  | 1.44 |  | m | 7.08e-09 | 0.0053013 | -3.1363e-09 | 0.0056372 | m | -2.6934e-05 |  | -2.9242e-05 |  | m | 10 | Hz | 0.001 | 0.00168 | 2.1648e-06 | 0.00144 | 2.0505e-06 | 0.0022127 | 2.1171e-06 | 0.70862 | 0.00094893 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 9.6858e-08 | 0.000518 | 1e-07 | 0.0010271 | 10Hz3m3m | -1.0045e-07 | 0.0025811 | -1.8439e-07 | 0.0031799 | -2.2777e-08 | 0.0025811 | -7.2432e-08 | 0.0023324 | 0.00168 | 0.0025811 | 0.0014399 | 0.0023324 | 1.0001 | 1.3075e-06 | 0.00010001 | 0.0035108 | 0.99997 | 2.6599e-06 | -6.1458e-05 | 0.0013962 | 1 | 2.6599e-06 | 2.1091e-06 | 0.0022287 | 0.0022127 | 4.2343e-06 | 0.70863 | 0.0029273 | 0.00168 | 5.3013e-06 | 0.00144 | 5.6372e-06
3 | m | 100 | 1.0989 | 0.00084611 | 2.7602 | 0.0011859 | m | -0.0012699 | 0.00062299 | -2.2972e-05 | 0.0009193 | m | 1.1 | 0.0036173 | 2.76 | 0.0030207 | m | OK | 1.1 |  | 2.76 |  | m | 8.8158e-09 | 0.0036173 | -5.4112e-08 | 0.0030207 | m | 0.00014692 |  | 0.00023175 |  | m | 100 | Hz | 0.001 | 0.0011001 | 1.0507e-06 | 0.0027602 | 1.5005e-06 | 0.0029714 | 1.4471e-06 | 1.1915 | 0.00037797 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | 4.9863e-08 | 0.0037035 | 1.0002e-07 | 0.0028793 | 100Hz3m3m | -2.785e-07 | 0.0034446 | 4.7449e-08 | 0.0013805 | -2.7381e-07 | 0.0034446 | 4.82e-08 | 0.001402 | 0.0010999 | 0.0034446 | 0.0027603 | 0.001402 | 1 | 3.7508e-06 | 9.9996e-05 | 0.0047442 | 0.99991 | 4.4492e-06 | 1.5841e-05 | 0.00094344 | 0.99991 | 4.4492e-06 | 1.7045e-05 | 0.00099779 | 0.0029711 | 2.8943e-06 | 1.1915 | 0.0012518 | 0.0011 | 3.6173e-06 | 0.00276 | 3.0207e-06
3 | m | 0.1 | -2.1006 | 0.00080984 | -0.97119 | 0.0014248 | m | -0.00080752 | 0.00051574 | 0.00078795 | 0.0013606 | m | -2.1 | 0.0032208 | -0.972 | 0.0052189 | m | OK | -2.1 |  | -0.972 |  | m | 4.7408e-09 | 0.0032208 | -7.8405e-09 | 0.0052189 | m | 0.00016822 |  | 2.649e-05 |  | m | 100 | mHz | 0.001 | -0.0020998 | 9.6012e-07 | -0.00097197 | 1.9701e-06 | 0.0023139 | 1.2017e-06 | -2.7081 | 0.00079207 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 5.162e-08 | 0.00079602 | 9.999e-08 | 0.003609 | 100mHz3m3m | 2.4238e-07 | 0.0035406 | -2.4863e-07 | 0.00047042 | 1.7694e-07 | 0.0035406 | -1.2903e-07 | 0.0015472 | -0.0020997 | 0.0035406 | -0.0009721 | 0.0015472 | 1.0001 | 1.6186e-06 | 0.00010001 | 0.0040947 | 1.0001 | 3.8242e-06 | -8.2887e-05 | 0.00068313 | 1.0001 | 3.8242e-06 | -2.0144e-05 | 0.0018535 | 0.002314 | 2.4034e-06 | -2.7081 | 0.0024382 | -0.0021 | 3.2208e-06 | -0.000972 | 5.2189e-06
10 | m | 500 | -3.3293 | 0.00025166 | -1.3489 | 0.00096068 | m | 0.00083503 | 0.0015064 | 0.00085674 | 0.0010991 | m | -3.33 | 0.0033772 | -1.35 | 0.0047467 | m | OK | -3.33 |  | -1.35 |  | m | -1.0161e-07 | 0.0033772 | 3.3311e-07 | 0.0047467 | m | -0.00012393 |  | 0.00028947 |  | m | 500 | Hz | 0.001 | -0.0033301 | 1.5273e-06 | -0.0013497 | 1.4598e-06 | 0.0035932 | 1.518e-06 | -2.7565 | 0.00040895 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | -2.1582e-09 | 0.0010202 | 3.0014e-07 | 0.0021028 | 500Hz10m10m | -1.3647e-07 | 0.0029467 | -4.6502e-07 | 0.001364 | -1.3541e-08 | 0.0029467 | 2.3529e-07 | 0.0020402 | -0.0033301 | 0.0029467 | -0.0013495 | 0.0020402 | 1 | 2.0742e-06 | 9.9922e-05 | 0.0011846 | 0.99999 | 4.3495e-06 | -4.6419e-05 | 0.00037546 | 1 | 4.3495e-06 | 8.7519e-05 | 0.001116 | 0.0035932 | 3.036e-06 | -2.7564 | 0.0013837 | -0.00333 | 3.3772e-06 | -0.00135 | 4.7467e-06
100 | m | 0.02 | 64.1984 | 0.0017342 | 34.5987 | 0.00039612 | m | -8.4972e-05 | 0.00039153 | -0.00047506 | 0.00021275 | m | 64.2 | 0.0054719 | 34.6 | 0.0088692 | m | OK | 64.2 |  | 34.6 |  | m | 1.5619e-07 | 0.0054719 | 6.339e-07 | 0.0088692 | m | -0.0015174 |  | -0.0007976 |  | m | 20 | mHz | 0.001 | 0.064198 | 1.7779e-06 | 0.034599 | 4.4964e-07 | 0.072928 | 1.5795e-06 | 0.49431 | 1.2776e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 20mHz100m10m | 20mHz100m100m | 110 | 128 | 20mHz100m10m | 8.7605e-07 | 0.0023787 | 9.9977e-07 | 0.00033772 | 20mHz100m100m | -4.0604e-07 | 0.0017536 | -4.3356e-06 | 0.0028984 | -2.0395e-08 | 0.0017536 | -2.7307e-06 | 0.0021281 | 0.064198 | 0.0017536 | 0.034596 | 0.0021281 | 1.0001 | 2.8399e-06 | 0.0001 | 0.00030124 | 1 | 1.7835e-06 | -4.3359e-05 | 6.1683e-05 | 1 | 1.7835e-06 | -2.3674e-07 | 0.00013374 | 0.07293 | 3.1618e-06 | 0.49431 | 0.00013616 | 0.0642 | 5.4719e-06 | 0.0346 | 8.8692e-06
100 | m | 0.05 | -64.5115 | 0.0013732 | -57.8963 | 0.0005678 | m | -0.001626 | 4.2867e-05 | 0.0012882 | 0.0017948 | m | -64.5 | 0.0060403 | -57.9 | 0.0065359 | m | OK | -64.5 |  | -57.9 |  | m | 3.1826e-09 | 0.0060403 | -6.7462e-10 | 0.0065359 | m | -0.0098583 |  | 0.0023991 |  | m | 50 | mHz | 0.001 | -0.06451 | 1.3738e-06 | -0.057898 | 1.8824e-06 | 0.086681 | 1.6206e-06 | -2.4102 | 1.9321e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -6.8791e-07 | 0.001896 | 1e-06 | 0.00092261 | 50mHz100m100m | -6.5693e-06 | 0.0036121 | 9.5977e-06 | 0.001982 | -5.699e-06 | 0.0036121 | 8.3253e-06 | 0.0018253 | -0.064516 | 0.0036121 | -0.057889 | 0.0018253 | 0.99993 | 1.9425e-06 | 0.0001 | 0.00024812 | 0.99993 | 3.8821e-06 | 9.5973e-05 | 5.9466e-05 | 0.99993 | 3.8821e-06 | 9.6569e-05 | 8.7384e-05 | 0.086676 | 3.2586e-06 | -2.4101 | 9.5546e-05 | -0.0645 | 6.0403e-06 | -0.0579 | 6.5359e-06
10 | m | 2000 | -5.0702 | 1.5998e-05 | 1.3408 | 0.00044604 | m | -7.8308e-05 | 0.0019507 | 0.00029231 | 0.00067422 | m | -5.07 | 0.0037306 | 1.34 | 0.0027401 | m | OK | -5.07 |  | 1.34 |  | m | 1.9893e-06 | 0.0037306 | 1.8096e-07 | 0.0027401 | m | -0.00015409 |  | 0.00055161 |  | m | 2 | kHz | 0.001 | -0.0050702 | 1.9508e-06 | 0.0013406 | 8.0841e-07 | 0.0052444 | 1.8973e-06 | 2.8831 | 0.00017678 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.2113e-07 | 0.0031596 | 3.0201e-07 | 0.00059968 | 2kHz10m10m | -9.1201e-07 | 0.0021293 | 8.1914e-07 | 0.00049468 | -3.7471e-07 | 0.0021293 | 4.6781e-07 | 0.00056601 | -0.0050705 | 0.0021293 | 0.001341 | 0.00056601 | 0.99996 | 3.1912e-06 | 9.9926e-05 | 0.00040135 | 0.99991 | 2.7666e-06 | 8.1602e-05 | 0.00027639 | 0.99994 | 2.7666e-06 | 9.4051e-05 | 0.00036128 | 0.0052441 | 3.7946e-06 | 2.8832 | 0.00050549 | -0.00507 | 3.7306e-06 | 0.00134 | 2.7401e-06
1 | m | 1 | -0.25292 | 0.0011823 | -0.78397 | 0.0019278 | m | -0.00091155 | 0.0014233 | -0.000976 | 0.00099823 | m | -0.252 | 0.0050223 | -0.783 | 0.0043678 | m | OK | -0.252 |  | -0.783 |  | m | 2.2224e-06 | 0.0050223 | -7.1527e-07 | 0.0043678 | m | -9.9009e-06 |  | 4.5585e-06 |  | m | 1 | Hz | 0.001 | -0.00025201 | 1.8503e-06 | -0.000783 | 2.1709e-06 | 0.00082255 | 2.1429e-06 | -1.8822 | 0.0022889 | 1 | 0m | 1m | 0 | 0.001 | 1Hz1m0m | 1Hz1m1m | 7 | 25 | 1Hz1m0m | 0 | 0.0026759 | 0 | 0.0018268 | 1Hz1m1m | 1.9306e-09 | 0.0034827 | 1.5994e-08 | 0.0014964 | 1.5881e-09 | 0.0034827 | 1.5994e-08 | 0.001555 | -0.00025201 | 0.0034827 | -0.00078298 | 0.001555 | 1 | 3.2401e-06 | 1.5994e-05 | 0.0041681 | 1 | 4.6948e-06 | 1.5994e-05 | 0.0041681 | 1 | 4.6948e-06 | 1.5994e-05 | 0.0041681 | 0.00082255 | 4.2857e-06 | -1.8822 | 0.006191 | -0.000252 | 5.0223e-06 | -0.000783 | 4.3678e-06
3 | m | 200 | -0.82921 | 0.00076967 | -1.249 | 0.00031373 | m | -0.0011205 | 0.0012931 | 0.00092436 | 0.00053526 | m | -0.828 | 0.0025114 | -1.25 | 0.0022178 | m | OK | -0.828 |  | -1.25 |  | m | 1.6928e-08 | 0.0025114 | 4.7213e-08 | 0.0022178 | m | -8.5559e-05 |  | 0.0001201 |  | m | 200 | Hz | 0.001 | -0.00082809 | 1.5048e-06 | -0.0012499 | 6.2043e-07 | 0.0014993 | 9.7891e-07 | -2.1559 | 0.00086734 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | 7.2959e-08 | 0.0035056 | 1.0007e-07 | 0.00029641 | 200Hz3m3m | -2.3484e-07 | 0.0040287 | 2.0149e-07 | 0.0022555 | -3.8857e-09 | 0.0040287 | 1.2539e-07 | 0.00078551 | -0.00082809 | 0.0040287 | -0.0012498 | 0.00078551 | 1.0001 | 4.323e-06 | 9.9993e-05 | 0.00039337 | 0.99992 | 4.0604e-06 | 6.7189e-05 | 0.00093484 | 1 | 4.0604e-06 | 9.1803e-05 | 0.00052855 | 0.0014994 | 1.9578e-06 | -2.1558 | 0.0018134 | -0.000828 | 2.5114e-06 | -0.00125 | 2.2178e-06
100 | m | 0.2 | -47.2005 | -0.00036215 | -56.5022 | -0.00031779 | m | -0.00011226 | 0.001458 | -7.8647e-05 | 0.0012481 | m | -47.2 | 0.0093173 | -56.5 | 0.0079329 | m | OK | -47.2 |  | -56.5 |  | m | -3.518e-07 | 0.0093173 | 7.8261e-07 | 0.0079329 | m | -0.00043441 |  | -0.0021316 |  | m | 200 | mHz | 0.001 | -0.0472 | 1.5023e-06 | -0.056502 | 1.2879e-06 | 0.073623 | 1.3801e-06 | -2.2667 | 1.9262e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | 8.5582e-07 | 0.0026992 | 1e-06 | 0.0016788 | 200mHz100m100m | -7.227e-06 | 0.0047385 | -6.1333e-06 | 0.0037133 | -4.8581e-06 | 0.0047385 | -4.0427e-06 | 0.003117 | -0.047205 | 0.0047385 | -0.056506 | 0.003117 | 1.0001 | 2.8474e-06 | 9.9996e-05 | 0.00038612 | 0.99993 | 5.2098e-06 | -6.1324e-05 | 6.2289e-05 | 0.99997 | 5.2098e-06 | -1.4045e-05 | 0.0001572 | 0.073621 | 2.7867e-06 | -2.2668 | 0.00016185 | -0.0472 | 9.3173e-06 | -0.0565 | 7.9329e-06
3 | m | 200 | 0.50703 | 0.0011544 | -2.9315 | 0.0017752 | m | 0.0011946 | 0.00098783 | -0.0012045 | 0.0010215 | m | 0.506 | 0.0041285 | -2.93 | 0.0040712 | m | OK | 0.506 |  | -2.93 |  | m | -5.5414e-08 | 0.0041285 | 1.3638e-07 | 0.0040712 | m | -0.00015963 |  | -0.00025744 |  | m | 200 | Hz | 0.001 | 0.00050584 | 1.5193e-06 | -0.0029303 | 2.0481e-06 | 0.0029736 | 2.0347e-06 | -1.3999 | 0.00051695 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | 7.2959e-08 | 0.0035056 | 1.0007e-07 | 0.00029641 | 200Hz3m3m | -2.3484e-07 | 0.0040287 | 2.0149e-07 | 0.0022555 | -2.3078e-07 | 0.0040287 | 2.0015e-07 | 0.0022297 | 0.00050561 | 0.0040287 | -0.0029301 | 0.0022297 | 1.0001 | 4.323e-06 | 9.9993e-05 | 0.00039337 | 0.99992 | 4.0604e-06 | 6.7189e-05 | 0.00093484 | 0.99992 | 4.0604e-06 | 6.7622e-05 | 0.00092769 | 0.0029734 | 4.0695e-06 | -1.3998 | 0.0013891 | 0.000506 | 4.1285e-06 | -0.00293 | 4.0712e-06
10 | m | 0.05 | -2.7108 | 0.00061116 | -2.7314 | 0.0019194 | m | -0.0005428 | 0.001057 | -0.0015622 | 0.00061347 | m | -2.71 | 0.0041321 | -2.73 | 0.0041216 | m | OK | -2.71 |  | -2.73 |  | m | -2.4081e-08 | 0.0041321 | 2.1176e-08 | 0.0041216 | m | -0.00028173 |  | 0.00020243 |  | m | 50 | mHz | 0.001 | -0.0027103 | 1.221e-06 | -0.0027298 | 2.0151e-06 | 0.0038467 | 1.6688e-06 | -2.3526 | 0.00043238 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -4.4673e-08 | 0.00062352 | 3.0006e-07 | 0.0029111 | 50mHz10m10m | -9.4711e-08 | 0.0037153 | 9.476e-08 | 0.00093829 | -5.0726e-08 | 0.0037153 | 2.7523e-07 | 0.0026724 | -0.0027103 | 0.0037153 | -0.0027295 | 0.0026724 | 0.99999 | 3.0408e-06 | 9.999e-05 | 0.0009815 | 0.99999 | 4.0101e-06 | 9.4764e-06 | 0.00026758 | 0.99999 | 4.0101e-06 | 8.9041e-05 | 0.00089515 | 0.0038467 | 3.3376e-06 | -2.3525 | 0.0012446 | -0.00271 | 4.1321e-06 | -0.00273 | 4.1216e-06
1 | m | 5 | 0.063729 | 0.0015588 | -0.8878 | 0.001682 | m | -0.0012989 | 0.0012989 | 0.0012173 | 0.00016319 | m | 0.064997 | 0.0055878 | -0.889 | 0.0033994 | m | OK | 0.065 |  | -0.889 |  | m | -3.1853e-06 | 0.0055878 | -2.3243e-07 | 0.0033994 | m | 2.7388e-05 |  | -1.5216e-05 |  | m | 5 | Hz | 0.001 | 6.5027e-05 | 2.029e-06 | -0.00088902 | 1.6899e-06 | 0.00089139 | 1.6919e-06 | -1.4978 | 0.0022743 | 1 | 0m | 1m | 0 | 0.001 | 5Hz1m0m | 5Hz1m1m | 9 | 27 | 5Hz1m0m | 0 | 0.0013745 | 0 | 0.0011936 | 5Hz1m1m | -2.1614e-08 | 0.00057989 | -3.2983e-08 | 0.001785 | -1.9267e-08 | 0.00057989 | -3.2983e-08 | 0.0017208 | 6.5008e-05 | 0.0005799 | -0.00088905 | 0.0017208 | 1 | 1.8204e-06 | -3.2981e-05 | 0.004329 | 0.99998 | 7.5282e-07 | -3.2981e-05 | 0.004329 | 0.99998 | 7.5282e-07 | -3.2981e-05 | 0.004329 | 0.00089137 | 3.3838e-06 | -1.4978 | 0.0062794 | 6.4997e-05 | 5.5878e-06 | -0.000889 | 3.3994e-06
1 | m | 200 | 0.40838 | 0.0019267 | -0.14237 | 0.0013688 | m | -0.0016275 | 0.0015348 | -0.00035095 | 0.00087921 | m | 0.41001 | 0.0046697 | -0.14198 | 0.0038043 | m | OK | 0.41 |  | -0.142 |  | m | 6.043e-06 | 0.0046697 | 1.747e-05 | 0.0038043 | m | 1.1911e-05 |  | -1.9109e-05 |  | m | 200 | Hz | 0.001 | 0.00041001 | 2.4633e-06 | -0.00014202 | 1.6268e-06 | 0.00043391 | 2.3877e-06 | -0.33344 | 0.0040004 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 2.9562e-05 | 0 | 0.0014921 | 200Hz1m1m | -9.3119e-08 | 0.00051164 | 7.5225e-08 | 0.0013653 | -4.0405e-08 | 0.00051164 | 7.5225e-08 | 0.0014371 | 0.00040997 | 0.00051166 | -0.00014194 | 0.0014371 | 1 | 1.4924e-06 | 7.5234e-05 | 0.0027488 | 0.99991 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.99996 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.00043389 | 4.7754e-06 | -0.33337 | 0.0084598 | 0.00041001 | 4.6697e-06 | -0.00014198 | 3.8043e-06
3 | m | 2000 | -0.37866 | 0.0013793 | 1.4002 | 0.0012026 | m | -0.00074861 | 0.001207 | 0.00010182 | 0.001556 | m | -0.378 | 0.0038156 | 1.4 | 0.0039083 | m | OK | -0.378 |  | 1.4 |  | m | 6.3419e-07 | 0.0038156 | 5.5431e-07 | 0.0039083 | m | 9.1576e-05 |  | 0.00012258 |  | m | 2 | kHz | 0.001 | -0.00037791 | 1.8328e-06 | 0.0014001 | 1.9665e-06 | 0.0014502 | 1.9577e-06 | 1.8344 | 0.0012703 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | -7.4486e-08 | 0.00079912 | 1.0022e-07 | 0.00047925 | 2kHz3m3m | -1.0953e-07 | 0.0028076 | 7.0263e-08 | 0.00084968 | -8.2374e-08 | 0.0028076 | 9.3476e-08 | 0.00056264 | -0.00037799 | 0.0028076 | 0.0014002 | 0.00056266 | 0.99993 | 9.0744e-07 | 9.9639e-05 | 0.00052319 | 0.99996 | 4.3725e-06 | 2.3388e-05 | 0.0011519 | 0.99994 | 4.3725e-06 | 8.2474e-05 | 0.00066472 | 0.0014501 | 3.9155e-06 | 1.8345 | 0.0026261 | -0.000378 | 3.8156e-06 | 0.0014 | 3.9083e-06
10 | m | 0.05 | 4.9912 | 0.0013605 | 4.8212 | 0.0004619 | m | 0.00090569 | 0.00088383 | 0.001338 | 0.00033627 | m | 4.99 | 0.0037061 | 4.82 | 0.0037795 | m | OK | 4.99 |  | 4.82 |  | m | 1.5309e-08 | 0.0037061 | -1.3986e-08 | 0.0037795 | m | 0.00029567 |  | -0.00018777 |  | m | 50 | mHz | 0.001 | 0.0049903 | 1.6224e-06 | 0.0048198 | 5.7133e-07 | 0.0069378 | 1.2326e-06 | 0.76802 | 0.00017292 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -4.4673e-08 | 0.00062352 | 3.0006e-07 | 0.0029111 | 50mHz10m10m | -9.4711e-08 | 0.0037153 | 9.476e-08 | 0.00093829 | -7.2822e-08 | 0.0037153 | 1.8457e-07 | 0.0018013 | 0.0049902 | 0.0037153 | 0.00482 | 0.0018013 | 0.99999 | 3.0408e-06 | 9.999e-05 | 0.0009815 | 0.99999 | 4.0101e-06 | 9.4764e-06 | 0.00026758 | 0.99999 | 4.0101e-06 | 4.9072e-05 | 0.00057989 | 0.0069378 | 2.4654e-06 | 0.76807 | 0.00067518 | 0.00499 | 3.7061e-06 | 0.00482 | 3.7795e-06
10 | m | 2 | -3.5291 | 0.0010029 | -1.2493 | 0.0013687 | m | 0.00089993 | 0.0006533 | 0.0003297 | 0.0015056 | m | -3.53 | 0.003236 | -1.25 | 0.0059269 | m | OK | -3.53 |  | -1.25 |  | m | -1.0454e-09 | 0.003236 | 1.028e-08 | 0.0059269 | m | -1.6831e-06 |  | 0.00037953 |  | m | 2 | Hz | 0.001 | -0.00353 | 1.1969e-06 | -0.0012496 | 2.0348e-06 | 0.0037447 | 1.3169e-06 | -2.8014 | 0.00052322 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.0342e-07 | 0.0024993 | 3.0003e-07 | 0.0035924 | 2Hz10m10m | 2.4445e-07 | 0.0013271 | 5.9488e-07 | 0.0020341 | 1.1843e-07 | 0.0013271 | 3.314e-07 | 0.0034267 | -0.0035299 | 0.0013271 | -0.0012493 | 0.0034267 | 1 | 2.9655e-06 | 9.9997e-05 | 0.00141 | 1 | 1.3475e-06 | 5.9483e-05 | 0.00027487 | 1 | 1.3475e-06 | 9.5687e-05 | 0.0012892 | 0.0037448 | 2.6337e-06 | -2.8013 | 0.0016604 | -0.00353 | 3.236e-06 | -0.00125 | 5.9269e-06
10 | m | 0.02 | -4.2985 | 0.0019209 | 1.0796 | 0.00079411 | m | 0.0012638 | 0.0006758 | -0.0007075 | 0.0016597 | m | -4.3 | 0.0040937 | 1.08 | 0.0046909 | m | OK | -4.3 |  | 1.08 |  | m | -1.3029e-09 | 0.0040937 | -1.8663e-08 | 0.0046909 | m | 0.00026122 |  | 0.00026469 |  | m | 20 | mHz | 0.001 | -0.0042997 | 2.0363e-06 | 0.0010803 | 1.8399e-06 | 0.0044334 | 2.0252e-06 | 2.8954 | 0.00041778 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | 1.1033e-07 | 0.00050493 | 2.9998e-07 | 0.00069903 | 20mHz10m10m | 6.5208e-07 | 0.0021735 | -3.5481e-07 | 0.0023208 | 2.2126e-07 | 0.0021735 | 1.659e-07 | 0.0010311 | -0.0042995 | 0.0021735 | 0.0010804 | 0.0010311 | 1 | 2.0522e-06 | 0.0001 | 0.00076835 | 1.0001 | 3.2705e-06 | -3.5476e-05 | 0.00025274 | 1 | 3.2705e-06 | 7.2262e-05 | 0.00066277 | 0.0044336 | 4.0503e-06 | 2.8955 | 0.0010665 | -0.0043 | 4.0937e-06 | 0.00108 | 4.6909e-06
100 | m | 0.2 | -23.3006 | 0.0012788 | 56.5015 | 0.00054171 | m | -0.00092967 | 0.00051552 | 0.0010598 | 0.00042461 | m | -23.3 | 0.011692 | 56.5 | 0.0050592 | m | OK | -23.3 |  | 56.5 |  | m | 5.8357e-07 | 0.011692 | 5.2581e-08 | 0.0050592 | m | 0.00037872 |  | 0.00042545 |  | m | 200 | mHz | 0.001 | -0.0233 | 1.3788e-06 | 0.0565 | 6.8829e-07 | 0.061116 | 8.2534e-07 | 1.9619 | 2.1294e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | 8.5582e-07 | 0.0026992 | 1e-06 | 0.0016788 | 200mHz100m100m | -7.227e-06 | 0.0047385 | -6.1333e-06 | 0.0037133 | -3.7349e-06 | 0.0047385 | -3.0514e-06 | 0.0028343 | -0.023303 | 0.0047385 | 0.056497 | 0.0028343 | 1.0001 | 2.8474e-06 | 9.9996e-05 | 0.00038612 | 0.99993 | 5.2098e-06 | -6.1324e-05 | 6.2289e-05 | 1 | 5.2098e-06 | 8.3736e-06 | 0.0002022 | 0.061116 | 1.6811e-06 | 1.9619 | 0.00020663 | -0.0233 | 1.1692e-05 | 0.0565 | 5.0592e-06
100 | m | 0.5 | -24.7998 | 0.00076238 | 27.3018 | 0.00070079 | m | -0.0016253 | 0.00081104 | 0.0016235 | 0.0009408 | m | -24.8 | 0.0065224 | 27.3 | 0.0060027 | m | OK | -24.8 |  | 27.3 |  | m | -2.6031e-07 | 0.0065224 | 6.2446e-08 | 0.0060027 | m | 0.0018291 |  | 0.00021584 |  | m | 500 | mHz | 0.001 | -0.024798 | 1.1131e-06 | 0.0273 | 1.1731e-06 | 0.036882 | 1.1464e-06 | 2.3082 | 3.0927e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 500mHz100m10m | 500mHz100m100m | 114 | 132 | 500mHz100m10m | 6.6981e-07 | 0.0019708 | 1e-06 | 0.0024639 | 500mHz100m100m | -6.0135e-06 | 0.0038832 | -9.871e-06 | 0.00038142 | -1.3264e-06 | 0.0038832 | -2.247e-06 | 0.0018419 | -0.024799 | 0.0038832 | 0.027298 | 0.0018419 | 1.0001 | 1.9805e-06 | 0.0001 | 0.00029193 | 0.99994 | 5.4565e-06 | -9.8714e-05 | 6.3495e-05 | 1 | 5.4565e-06 | 4.0647e-05 | 0.0002237 | 0.036883 | 2.3016e-06 | 2.3082 | 0.00023209 | -0.0248 | 6.5224e-06 | 0.0273 | 6.0027e-06
3 | m | 100 | -1.5283 | 0.0017986 | -0.57133 | 0.0003123 | m | 0.0017732 | 0.00024925 | -0.0014479 | 0.00091592 | m | -1.53 | 0.0039043 | -0.57 | 0.0059272 | m | OK | -1.53 |  | -0.57 |  | m | 4.689e-09 | 0.0039043 | 1.1582e-08 | 0.0059272 | m | -3.4558e-05 |  | 0.00011498 |  | m | 100 | Hz | 0.001 | -0.00153 | 1.8158e-06 | -0.00056989 | 9.677e-07 | 0.0016327 | 1.7348e-06 | -2.785 | 0.00067762 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | 4.9863e-08 | 0.0037035 | 1.0002e-07 | 0.0028793 | 100Hz3m3m | -2.785e-07 | 0.0034446 | 4.7449e-08 | 0.0013805 | -5.4019e-08 | 0.0034446 | 8.3386e-08 | 0.0024051 | -0.0015301 | 0.0034446 | -0.0005698 | 0.0024051 | 1 | 3.7508e-06 | 9.9996e-05 | 0.0047442 | 0.99991 | 4.4492e-06 | 1.5841e-05 | 0.00094344 | 1 | 4.4492e-06 | 7.3373e-05 | 0.0035418 | 0.0016327 | 3.4696e-06 | -2.785 | 0.0037922 | -0.00153 | 3.9043e-06 | -0.00057 | 5.9272e-06
1 | m | 0.1 | -0.035456 | 0.00063268 | 0.10568 | 0.00031354 | m | 4.374e-05 | 0.0013773 | 0.00067762 | 0.00088682 | m | -0.035501 | 0.0028631 | 0.105 | 0.0021412 | m | OK | -0.0355 |  | 0.105 |  | m | -6.9218e-07 | 0.0028631 | -2.3393e-07 | 0.0021412 | m | -1.7309e-07 |  | 7.9634e-07 |  | m | 100 | mHz | 0.001 | -3.55e-05 | 1.5157e-06 | 0.000105 | 9.4061e-07 | 0.00011084 | 1.0147e-06 | 1.8968 | 0.013236 | 1 | 0m | 1m | 0 | 0.001 | 100mHz1m0m | 100mHz1m1m | 4 | 22 | 100mHz1m0m | 0 | 0.00080668 | 0 | 0.00065996 | 100mHz1m1m | -6.591e-08 | 0.0030622 | 7.4188e-09 | 0.0010126 | -7.3054e-09 | 0.0030622 | 7.4188e-09 | 0.00069905 | -3.5507e-05 | 0.0030622 | 0.00010501 | 0.00069905 | 1 | 1.0422e-06 | 7.4137e-06 | 0.0020973 | 0.99993 | 4.0096e-06 | 7.4137e-06 | 0.0020973 | 0.99999 | 4.0096e-06 | 7.4137e-06 | 0.0020973 | 0.00011084 | 2.0294e-06 | 1.8968 | 0.026556 | -3.5501e-05 | 2.8631e-06 | 0.000105 | 2.1412e-06
10 | m | 1 | 5.4808 | 0.0020241 | 7.7687 | 0.001824 | m | 0.0005159 | 0.00062851 | -0.0012318 | 0.0010142 | m | 5.48 | 0.0051254 | 7.77 | 0.0046813 | m | OK | 5.48 |  | 7.77 |  | m | 1.2593e-08 | 0.0051254 | -3.6995e-08 | 0.0046813 | m | 0.00032258 |  | -4.5206e-05 |  | m | 1 | Hz | 0.001 | 0.0054803 | 2.1194e-06 | 0.00777 | 2.087e-06 | 0.0095082 | 2.0978e-06 | 0.95651 | 0.00022178 | 3 | 3m | 10m | 0.003 | 0.01 | 1Hz10m3m | 1Hz10m10m | 79 | 97 | 1Hz10m3m | 1.1918e-07 | 0.0012472 | 2.9995e-07 | 0.0010183 | 1Hz10m10m | -1.9857e-07 | 0.0035657 | 2.5208e-07 | 0.00099493 | -1.7625e-07 | 0.0035657 | 2.5544e-07 | 0.00099657 | 0.0054801 | 0.0035657 | 0.0077702 | 0.00099658 | 1 | 3.4605e-06 | 0.0001 | 0.00080072 | 0.99998 | 3.5806e-06 | 2.5206e-05 | 0.00034411 | 0.99998 | 3.5806e-06 | 3.0461e-05 | 0.00037619 | 0.0095081 | 4.1958e-06 | 0.95654 | 0.0005816 | 0.00548 | 5.1254e-06 | 0.00777 | 4.6813e-06
3 | m | 0.5 | -1.5308 | 0.0011381 | 0.12443 | 0.0015389 | m | -0.00073979 | 0.0015944 | -0.0006942 | 0.0018729 | m | -1.53 | 0.0039626 | 0.125 | 0.0077478 | m | OK | -1.53 |  | 0.125 |  | m | -3.1272e-09 | 0.0039626 | -6.0885e-09 | 0.0077478 | m | -1.231e-05 |  | 0.00012883 |  | m | 500 | mHz | 0.001 | -0.00153 | 1.959e-06 | 0.00012513 | 2.424e-06 | 0.0015351 | 1.9624e-06 | 3.06 | 0.0015772 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | -8.2739e-09 | 0.003378 | 9.9993e-08 | 0.0032475 | 500mHz3m3m | -9.8373e-08 | 0.00084011 | 1.0931e-07 | 0.0026046 | -3.2381e-08 | 0.00084011 | 1.0249e-07 | 0.0030755 | -0.00153 | 0.00084012 | 0.00012523 | 0.0030755 | 0.99999 | 4.1811e-06 | 0.0001 | 0.0050388 | 0.99997 | 3.2335e-06 | 3.6436e-05 | 0.00099174 | 0.99999 | 3.2335e-06 | 8.2995e-05 | 0.0039559 | 0.0015351 | 3.9248e-06 | 3.0601 | 0.0050596 | -0.00153 | 3.9626e-06 | 0.000125 | 7.7478e-06
10 | m | 0.5 | 7.4589 | 0.0012013 | -3.6704 | 0.00061182 | m | -0.0014991 | 0.001047 | 0.00044178 | 0.0013771 | m | 7.46 | 0.0036723 | -3.67 | 0.0049561 | m | OK | 7.46 |  | -3.67 |  | m | 2.6449e-10 | 0.0036723 | -4.8133e-09 | 0.0049561 | m | 0.00038953 |  | -0.00079932 |  | m | 500 | mHz | 0.001 | 0.0074604 | 1.5935e-06 | -0.0036708 | 1.5069e-06 | 0.0083146 | 1.577e-06 | -0.45726 | 0.00018331 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | -2.9748e-07 | 0.0027488 | 3.0002e-07 | 0.0013977 | 500mHz10m10m | -7.983e-07 | 0.0019224 | 5.4661e-07 | 0.0032719 | -6.7772e-07 | 0.0019224 | 4.8724e-07 | 0.0028206 | 0.0074597 | 0.0019224 | -0.0036703 | 0.0028206 | 0.9999 | 2.7529e-06 | 0.0001 | 0.00077851 | 0.99992 | 4.1819e-06 | 5.4664e-05 | 0.0004402 | 0.99992 | 4.1819e-06 | 6.558e-05 | 0.00052166 | 0.0083139 | 3.1542e-06 | -0.45719 | 0.0006376 | 0.00746 | 3.6723e-06 | -0.00367 | 4.9561e-06
3 | m | 5 | -1.5286 | 0.00055428 | -0.91303 | 0.0011632 | m | 0.0014787 | 0.00093517 | 0.0018801 | 0.0010516 | m | -1.53 | 0.0032651 | -0.915 | 0.0043477 | m | OK | -1.53 |  | -0.915 |  | m | 8.1081e-09 | 0.0032651 | -7.0951e-09 | 0.0043477 | m | -3.7065e-05 |  | 8.6607e-05 |  | m | 5 | Hz | 0.001 | -0.00153 | 1.0871e-06 | -0.00091491 | 1.5681e-06 | 0.0017827 | 1.2321e-06 | -2.6027 | 0.00081722 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | -4.9337e-08 | 0.00079014 | 1.0001e-07 | 0.0022765 | 5Hz3m3m | 2.8449e-07 | 0.0026935 | -6.5054e-08 | 0.0018253 | 8.1308e-08 | 0.0026935 | 3.5409e-08 | 0.0020999 | -0.00153 | 0.0026935 | -0.00091488 | 0.0020999 | 0.99995 | 2.7916e-06 | 9.9998e-05 | 0.0031263 | 1.0001 | 3.03e-06 | -2.1698e-05 | 0.00068935 | 1 | 3.03e-06 | 5.2371e-05 | 0.0021726 | 0.0017827 | 2.4643e-06 | -2.6026 | 0.0027187 | -0.00153 | 3.2651e-06 | -0.000915 | 4.3477e-06
1 | m | 50 | 0.63847 | 8.7537e-05 | -0.23509 | 0.0013194 | m | 0.00043698 | 0.00079728 | -0.0010646 | 0.0015163 | m | 0.638 | 0.0025303 | -0.23399 | 0.004547 | m | OK | 0.638 |  | -0.234 |  | m | 1.9689e-06 | 0.0025303 | 5.375e-06 | 0.004547 | m | 3.0003e-05 |  | -2.3934e-05 |  | m | 50 | Hz | 0.001 | 0.00063803 | 8.0207e-07 | -0.00023402 | 2.0099e-06 | 0.0006796 | 1.0228e-06 | -0.35155 | 0.0028063 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | 0 | 0.0015478 | 0 | 0.00027825 | 50Hz1m1m | -7.8859e-08 | 0.0027617 | 2.6271e-08 | 0.0027705 | -5.3592e-08 | 0.0027617 | 2.6271e-08 | 0.001972 | 0.00063798 | 0.0027617 | -0.000234 | 0.001972 | 1 | 1.5727e-06 | 2.6285e-05 | 0.0042513 | 0.99992 | 2.989e-06 | 2.6285e-05 | 0.0042513 | 0.99995 | 2.989e-06 | 2.6285e-05 | 0.0042513 | 0.00067956 | 2.0456e-06 | -0.35153 | 0.0070409 | 0.000638 | 2.5303e-06 | -0.00023399 | 4.547e-06
3 | m | 1000 | 0.58099 | 0.0018023 | 1.7994 | 0.0018584 | m | -8.6953e-05 | 0.00099136 | -0.00047852 | 0.00050922 | m | 0.581 | 0.004731 | 1.8 | 0.0039763 | m | OK | 0.581 |  | 1.8 |  | m | 6.4396e-07 | 0.004731 | -1.3586e-07 | 0.0039763 | m | 7.5038e-05 |  | -9.4178e-05 |  | m | 1 | kHz | 0.001 | 0.00058108 | 2.057e-06 | 0.0017999 | 1.9269e-06 | 0.0018914 | 1.9395e-06 | 1.2585 | 0.0010812 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 3.0166e-08 | 0.00050233 | 1.0023e-07 | 0.0010421 | 1kHz3m3m | 1.2457e-07 | 0.0038805 | -1.7428e-08 | 0.003317 | 7.224e-08 | 0.0038805 | 4.7791e-08 | 0.002056 | 0.00058115 | 0.0038805 | 0.0018 | 0.002056 | 1 | 1.9109e-06 | 9.9824e-05 | 0.0013661 | 1 | 5.4207e-06 | -5.8979e-06 | 0.001306 | 1 | 5.4207e-06 | 5.2705e-05 | 0.0013393 | 0.0018914 | 3.879e-06 | 1.2586 | 0.0025436 | 0.000581 | 4.731e-06 | 0.0018 | 3.9763e-06
1 | m | 0.05 | 0.11698 | 7.1889e-05 | -0.80048 | 0.0013292 | m | 0.0019561 | 0.0010257 | 0.00053515 | 5.5686e-05 | m | 0.115 | 0.0042495 | -0.801 | 0.0026924 | m | OK | 0.115 |  | -0.801 |  | m | -4.4428e-06 | 0.0042495 | -6.4031e-07 | 0.0026924 | m | 2.139e-05 |  | -1.4906e-05 |  | m | 50 | mHz | 0.001 | 0.00011502 | 1.0282e-06 | -0.00080101 | 1.3303e-06 | 0.00080923 | 1.3249e-06 | -1.4282 | 0.0012792 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.00036442 | 0 | 0.0010831 | 50mHz1m1m | -2.7165e-08 | 0.0017934 | -2.9092e-08 | 0.003713 | -2.1983e-08 | 0.0017934 | -2.9092e-08 | 0.0032113 | 0.000115 | 0.0017934 | -0.00080104 | 0.0032113 | 1 | 1.1427e-06 | -2.9093e-05 | 0.0046238 | 0.99997 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.99998 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.00080921 | 2.6498e-06 | -1.4282 | 0.0052844 | 0.000115 | 4.2495e-06 | -0.000801 | 2.6924e-06
3 | m | 10 | 0.47411 | 0.00078153 | -2.0091 | 0.00077776 | m | 0.00014969 | 0.0010368 | 0.00082002 | 3.3309e-05 | m | 0.474 | 0.0054118 | -2.01 | 0.0020329 | m | OK | 0.474 |  | -2.01 |  | m | -5.8607e-10 | 0.0054118 | -7.2605e-09 | 0.0020329 | m | -4.1205e-05 |  | 4.8548e-05 |  | m | 10 | Hz | 0.001 | 0.00047396 | 1.2983e-06 | -0.00201 | 7.7847e-07 | 0.0020651 | 8.1418e-07 | -1.3392 | 0.00061802 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 9.6858e-08 | 0.000518 | 1e-07 | 0.0010271 | 10Hz3m3m | -1.0045e-07 | 0.0025811 | -1.8439e-07 | 0.0031799 | -8.2179e-09 | 0.0025811 | -5.1447e-08 | 0.0021735 | 0.00047395 | 0.0025811 | -0.00201 | 0.0021735 | 1.0001 | 1.3075e-06 | 0.00010001 | 0.0035108 | 0.99997 | 2.6599e-06 | -6.1458e-05 | 0.0013962 | 1 | 2.6599e-06 | 1.4024e-05 | 0.0023847 | 0.0020651 | 1.6284e-06 | -1.3392 | 0.002686 | 0.000474 | 5.4118e-06 | -0.00201 | 2.0329e-06
100 | m | 0.1 | -17.7012 | 0.0014769 | -19.2985 | 0.0010813 | m | 0.00032001 | 0.00013588 | 0.00039142 | 0.0010164 | m | -17.7 | 0.0036688 | -19.3 | 0.0035669 | m | OK | -17.7 |  | -19.3 |  | m | -8.3641e-08 | 0.0036688 | -4.2372e-08 | 0.0035669 | m | -0.00156 |  | 0.001089 |  | m | 100 | mHz | 0.001 | -0.017702 | 1.4831e-06 | -0.019299 | 1.484e-06 | 0.026188 | 1.4836e-06 | -2.3131 | 5.665e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 100mHz100m10m | 100mHz100m100m | 112 | 130 | 100mHz100m10m | -3.2234e-07 | 0.0035671 | 1.0001e-06 | 0.00099585 | 100mHz100m100m | 9.3522e-06 | 0.0047814 | -5.5623e-06 | 0.0011289 | 1.4177e-06 | 0.0047814 | -1.8027e-07 | 0.0010198 | -0.0177 | 0.0047814 | -0.019299 | 0.0010198 | 0.99997 | 3.8623e-06 | 9.9999e-05 | 0.00012488 | 1.0001 | 4.8348e-06 | -5.5628e-05 | 5.17e-05 | 0.99999 | 4.8348e-06 | 7.2008e-05 | 0.00011172 | 0.026187 | 2.9699e-06 | -2.313 | 0.00015911 | -0.0177 | 3.6688e-06 | -0.0193 | 3.5669e-06
10 | m | 100 | -3.7392 | 0.00055549 | 4.2805 | 0.00052174 | m | 0.00052916 | 0.0011475 | 0.00047124 | 8.0977e-05 | m | -3.74 | 0.0042529 | 4.28 | 0.0038244 | m | OK | -3.74 |  | 4.28 |  | m | 8.9724e-08 | 0.0042529 | 3.5833e-08 | 0.0038244 | m | 0.0002635 |  | -4.785e-06 |  | m | 100 | Hz | 0.001 | -0.0037397 | 1.2748e-06 | 0.00428 | 5.2798e-07 | 0.0056837 | 9.2828e-07 | 2.2889 | 0.00017962 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 1.7855e-07 | 0.000589 | 3.0004e-07 | 0.003002 | 100Hz10m10m | -1.4533e-07 | 0.003003 | -7.1273e-07 | 0.00061156 | 5.4376e-08 | 0.003003 | -8.8233e-08 | 0.0020855 | -0.0037397 | 0.003003 | 0.0042799 | 0.0020855 | 1.0001 | 2.2555e-06 | 9.9994e-05 | 0.0012567 | 0.99999 | 3.2181e-06 | -7.1258e-05 | 0.00027777 | 1 | 3.2181e-06 | 3.434e-05 | 0.00088138 | 0.0056838 | 1.8566e-06 | 2.289 | 0.00095179 | -0.00374 | 4.2529e-06 | 0.00428 | 3.8244e-06
3 | m | 500 | -1.6018 | 0.0019013 | -0.21638 | 6.0937e-05 | m | -0.0017006 | 0.0013595 | 0.0004963 | 0.0002689 | m | -1.6 | 0.004626 | -0.217 | 0.0042397 | m | OK | -1.6 |  | -0.217 |  | m | -1.5316e-07 | 0.004626 | 5.9456e-08 | 0.0042397 | m | -0.00012097 |  | 0.00012115 |  | m | 500 | Hz | 0.001 | -0.0016001 | 2.3373e-06 | -0.00021688 | 2.7572e-07 | 0.0016148 | 2.3165e-06 | -3.0069 | 0.00025773 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | -9.1984e-08 | 0.0023561 | 9.9972e-08 | 0.0019901 | 500Hz3m3m | -5.275e-09 | 0.00090694 | 1.4777e-07 | 0.0036975 | -6.5331e-08 | 0.00090694 | 1.1466e-07 | 0.0025149 | -0.0016002 | 0.00090695 | -0.00021676 | 0.0025149 | 0.99991 | 4.4539e-06 | 9.9971e-05 | 0.0029075 | 1 | 4.0081e-06 | 4.9196e-05 | 0.0018092 | 0.99994 | 4.0081e-06 | 8.4364e-05 | 0.0025699 | 0.0016146 | 4.6329e-06 | -3.0068 | 0.0026211 | -0.0016 | 4.626e-06 | -0.000217 | 4.2397e-06
3 | m | 20 | 1.8584 | 0.00014523 | 0.13724 | 0.00089922 | m | -0.0015244 | 0.00067533 | -0.0016778 | 0.00096398 | m | 1.86 | 0.0014639 | 0.139 | 0.0062364 | m | OK | 1.86 |  | 0.139 |  | m | 1.6688e-08 | 0.0014639 | -3.36e-08 | 0.0062364 | m | -4.5108e-05 |  | -7.8994e-05 |  | m | 20 | Hz | 0.001 | 0.00186 | 6.9077e-07 | 0.00013892 | 1.3183e-06 | 0.0018651 | 6.9581e-07 | 0.074552 | 0.00070537 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | 5.1085e-09 | 0.0037612 | 9.9997e-08 | 0.0032193 | 20Hz3m3m | 1.6919e-07 | 0.0020234 | -1.1325e-07 | 0.0010612 | 7.6084e-08 | 0.0020234 | 7.7524e-09 | 0.0022858 | 0.00186 | 0.0020234 | 0.00013893 | 0.0022858 | 1 | 4.4939e-06 | 0.00010001 | 0.005064 | 1.0001 | 3.9936e-06 | -3.7756e-05 | 0.00038761 | 1 | 3.9936e-06 | 4.0414e-05 | 0.0030411 | 0.0018652 | 1.3916e-06 | 0.074593 | 0.0033524 | 0.00186 | 1.4639e-06 | 0.000139 | 6.2364e-06
10 | m | 20 | -3.32 | 0.00038738 | 0.026621 | 9.9575e-06 | m | -2.882e-05 | 0.00058573 | -0.0014004 | 0.0019563 | m | -3.32 | 0.0014053 | 0.0277 | 0.0044299 | m | OK | -3.32 |  | 0.0277 |  | m | -3.3628e-08 | 0.0014053 | 2.5024e-08 | 0.0044299 | m | -5.9621e-06 |  | 0.00032167 |  | m | 20 | Hz | 0.001 | -0.00332 | 7.0224e-07 | 2.8022e-05 | 1.9564e-06 | 0.0033201 | 7.0241e-07 | 3.1332 | 0.00058923 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | -2.0254e-08 | 0.0038673 | 3.0005e-07 | 0.0014224 | 20Hz10m10m | 8.368e-07 | 0.00284 | 3.1494e-07 | 0.00021184 | 1.8941e-08 | 0.00284 | 3.0073e-07 | 0.0013671 | -0.00332 | 0.00284 | 2.8322e-05 | 0.0013671 | 0.99999 | 5.4376e-06 | 9.9993e-05 | 0.00064666 | 1.0001 | 2.9017e-06 | 3.1488e-05 | 0.00019065 | 1 | 2.9017e-06 | 9.686e-05 | 0.00062581 | 0.0033201 | 1.4049e-06 | 3.1332 | 0.0013343 | -0.00332 | 1.4053e-06 | 2.77e-05 | 4.4299e-06
1 | m | 200 | 0.26315 | 6.1356e-06 | 0.011967 | 0.00053283 | m | -0.0018578 | 0.00077851 | -0.00022847 | 0.00097686 | m | 0.265 | 0.0015608 | 0.012215 | 0.0023393 | m | OK | 0.265 |  | 0.0122 |  | m | -6.7725e-07 | 0.0015608 | 1.465e-05 | 0.0023393 | m | 6.7815e-06 |  | -4.9859e-06 |  | m | 200 | Hz | 0.001 | 0.00026501 | 7.7853e-07 | 1.2195e-05 | 1.1127e-06 | 0.00026529 | 7.7939e-07 | 0.045985 | 0.0041922 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 2.9562e-05 | 0 | 0.0014921 | 200Hz1m1m | -9.3119e-08 | 0.00051164 | 7.5225e-08 | 0.0013653 | -2.4703e-08 | 0.00051164 | 7.5225e-08 | 0.0014585 | 0.00026498 | 0.00051164 | 1.227e-05 | 0.0014585 | 1 | 1.4924e-06 | 7.5234e-05 | 0.0027488 | 0.99991 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.99998 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.00026528 | 1.5588e-06 | 0.046061 | 0.0088235 | 0.000265 | 1.5608e-06 | 1.2215e-05 | 2.3393e-06
1 | m | 0.05 | 0.20746 | 6.7533e-05 | 0.48206 | 0.00021991 | m | -0.0015323 | 5.5443e-06 | 0.0010465 | 0.0012175 | m | 0.20901 | 0.0025682 | 0.481 | 0.0023293 | m | OK | 0.209 |  | 0.481 |  | m | 6.6542e-06 | 0.0025682 | -2.8902e-06 | 0.0023293 | m | -4.3616e-06 |  | 1.0044e-05 |  | m | 50 | mHz | 0.001 | 0.000209 | 6.776e-08 | 0.00048101 | 1.2372e-06 | 0.00052445 | 1.135e-06 | 1.1609 | 0.00094753 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.00036442 | 0 | 0.0010831 | 50mHz1m1m | -2.7165e-08 | 0.0017934 | -2.9092e-08 | 0.003713 | -1.4247e-08 | 0.0017934 | -2.9092e-08 | 0.0024624 | 0.00020898 | 0.0017934 | 0.00048098 | 0.0024624 | 1 | 1.1427e-06 | -2.9093e-05 | 0.0046238 | 0.99997 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.99999 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.00052444 | 2.2701e-06 | 1.1609 | 0.0049971 | 0.00020901 | 2.5682e-06 | 0.000481 | 2.3293e-06
10 | m | 2 | 0.62041 | 0.0010636 | -3.8981 | 0.0010542 | m | -0.001202 | 0.0016419 | 0.0018523 | 0.0017592 | m | 0.622 | 0.0062752 | -3.9 | 0.0041668 | m | OK | 0.622 |  | -3.9 |  | m | -1.0814e-08 | 0.0062752 | 8.6733e-10 | 0.0041668 | m | -0.00038916 |  | 7.0363e-05 |  | m | 2 | Hz | 0.001 | 0.00062161 | 1.9563e-06 | -0.0038999 | 2.0509e-06 | 0.0039492 | 2.0486e-06 | -1.4127 | 0.00049597 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.0342e-07 | 0.0024993 | 3.0003e-07 | 0.0035924 | 2Hz10m10m | 2.4445e-07 | 0.0013271 | 5.9488e-07 | 0.0020341 | 1.2255e-07 | 0.0013271 | 3.4001e-07 | 0.0033811 | 0.00062173 | 0.0013271 | -0.0038996 | 0.0033811 | 1 | 2.9655e-06 | 9.9997e-05 | 0.00141 | 1 | 1.3475e-06 | 5.9483e-05 | 0.00027487 | 1 | 1.3475e-06 | 9.4504e-05 | 0.001256 | 0.0039493 | 4.0973e-06 | -1.4126 | 0.0016005 | 0.000622 | 6.2752e-06 | -0.0039 | 4.1668e-06
3 | m | 20 | -2.2113 | 0.0010116 | -2.0108 | 0.00053493 | m | -0.0014815 | 0.0008089 | -0.00079253 | 0.0008133 | m | -2.21 | 0.002443 | -2.01 | 0.0024677 | m | OK | -2.21 |  | -2.01 |  | m | -1.3991e-07 | 0.002443 | 7.2153e-08 | 0.0024677 | m | 0.00019783 |  | 3.124e-05 |  | m | 20 | Hz | 0.001 | -0.0022098 | 1.2952e-06 | -0.00201 | 9.7345e-07 | 0.0029872 | 1.1606e-06 | -2.4035 | 0.00037846 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | 5.1085e-09 | 0.0037612 | 9.9997e-08 | 0.0032193 | 20Hz3m3m | 1.6919e-07 | 0.0020234 | -1.1325e-07 | 0.0010612 | 1.6814e-07 | 0.0020234 | -1.1188e-07 | 0.001075 | -0.0022096 | 0.0020234 | -0.0020101 | 0.001075 | 1 | 4.4939e-06 | 0.00010001 | 0.005064 | 1.0001 | 3.9936e-06 | -3.7756e-05 | 0.00038761 | 1.0001 | 3.9936e-06 | -3.6872e-05 | 0.00041761 | 0.0029873 | 2.3213e-06 | -2.4036 | 0.00086448 | -0.00221 | 2.443e-06 | -0.00201 | 2.4677e-06
3 | m | 1 | 0.79968 | 0.00024653 | -0.63636 | 0.00057906 | m | -0.0012919 | 0.0017369 | 0.0017554 | 0.0011898 | m | 0.801 | 0.0041783 | -0.638 | 0.0046488 | m | OK | 0.801 |  | -0.638 |  | m | -6.9045e-11 | 0.0041783 | -7.5982e-11 | 0.0046488 | m | -2.3782e-05 |  | -0.00011164 |  | m | 1 | Hz | 0.001 | 0.00080098 | 1.7543e-06 | -0.00063811 | 1.3232e-06 | 0.0010241 | 1.6008e-06 | -0.67271 | 0.0014699 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | -4.974e-08 | 0.0035546 | 1.0001e-07 | 0.0023717 | 1Hz3m3m | -1.534e-07 | 0.001355 | 2.3557e-07 | 0.0026795 | -5.0988e-08 | 0.001355 | 1.0164e-07 | 0.0023754 | 0.00080093 | 0.001355 | -0.00063801 | 0.0023754 | 0.99995 | 4.9649e-06 | 0.0001 | 0.0043791 | 0.99995 | 2.2464e-06 | 7.8523e-05 | 0.0011693 | 0.99995 | 2.2464e-06 | 9.9741e-05 | 0.0043405 | 0.001024 | 3.2016e-06 | -0.67261 | 0.0052424 | 0.000801 | 4.1783e-06 | -0.000638 | 4.6488e-06
10 | m | 200 | -7.9415 | 0.0001061 | 2.6003 | 0.0010075 | m | -0.0015319 | 0.0012729 | -0.00048982 | 0.00096603 | m | -7.94 | 0.0031984 | 2.6 | 0.006328 | m | OK | -7.94 |  | 2.6 |  | m | -1.6975e-07 | 0.0031984 | 9.0254e-08 | 0.006328 | m | 4.9834e-05 |  | 0.00079065 |  | m | 200 | Hz | 0.001 | -0.00794 | 1.2773e-06 | 0.0026008 | 1.3958e-06 | 0.0083551 | 1.2893e-06 | 2.825 | 0.00016574 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -2.0914e-07 | 0.0035527 | 2.9996e-07 | 0.0039867 | 200Hz10m10m | -9.7041e-08 | 0.0015213 | 8.9286e-07 | 0.0029613 | -1.2338e-07 | 0.0015213 | 7.5353e-07 | 0.0032023 | -0.0079401 | 0.0015213 | 0.0026015 | 0.0032023 | 0.99993 | 3.8094e-06 | 9.9999e-05 | 0.0018739 | 0.99999 | 2.1803e-06 | 8.9261e-05 | 0.00036252 | 0.99998 | 2.1803e-06 | 9.1784e-05 | 0.00071769 | 0.0083549 | 2.5786e-06 | 2.8251 | 0.00079054 | -0.00794 | 3.1984e-06 | 0.0026 | 6.328e-06
3 | m | 2 | -1.6001 | 0.00024578 | 2.3891 | 0.00045318 | m | -0.00037259 | 0.0012915 | -0.00091034 | 0.0014182 | m | -1.6 | 0.004153 | 2.39 | 0.0035056 | m | OK | -1.6 |  | 2.39 |  | m | 9.9128e-09 | 0.004153 | 1.5826e-08 | 0.0035056 | m | 0.00026428 |  | -6.4474e-06 |  | m | 2 | Hz | 0.001 | -0.0015997 | 1.3147e-06 | 0.00239 | 1.4889e-06 | 0.002876 | 1.4372e-06 | 2.1607 | 0.00047668 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | 4.1868e-08 | 1.9674e-05 | 1e-07 | 0.0023153 | 2Hz3m3m | 1.6117e-07 | 0.0013963 | 2.2041e-07 | 0.0014454 | 1.5378e-07 | 0.0013963 | 2.1295e-07 | 0.0014994 | -0.0015996 | 0.0013963 | 0.0023902 | 0.0014994 | 1 | 3.9354e-07 | 9.9999e-05 | 0.0031515 | 1.0001 | 2.0838e-06 | 7.3462e-05 | 0.0011665 | 1.0001 | 2.0838e-06 | 7.5108e-05 | 0.0012896 | 0.0028761 | 2.8745e-06 | 2.1607 | 0.0016037 | -0.0016 | 4.153e-06 | 0.00239 | 3.5056e-06
3 | m | 0.5 | -0.37124 | 0.00053782 | -1.1401 | 0.00040329 | m | 0.0018731 | 0.00074768 | -0.00011788 | 0.00015649 | m | -0.373 | 0.0055544 | -1.14 | 0.0020493 | m | OK | -0.373 |  | -1.14 |  | m | 3.327e-09 | 0.0055544 | -2.7933e-09 | 0.0020493 | m | -0.00011077 |  | 2.2728e-05 |  | m | 500 | mHz | 0.001 | -0.00037311 | 9.2101e-07 | -0.00114 | 4.3259e-07 | 0.0011995 | 5.011e-07 | -1.8871 | 0.00073832 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | -8.2739e-09 | 0.003378 | 9.9993e-08 | 0.0032475 | 500mHz3m3m | -9.8373e-08 | 0.00084011 | 1.0931e-07 | 0.0026046 | -1.7261e-08 | 0.00084011 | 1.0092e-07 | 0.0031834 | -0.00037313 | 0.00084011 | -0.0011399 | 0.0031834 | 0.99999 | 4.1811e-06 | 0.0001 | 0.0050388 | 0.99997 | 3.2335e-06 | 3.6436e-05 | 0.00099174 | 0.99999 | 3.2335e-06 | 9.3663e-05 | 0.0046351 | 0.0011995 | 1.0022e-06 | -1.887 | 0.0048646 | -0.000373 | 5.5544e-06 | -0.00114 | 2.0493e-06
3 | m | 10 | -0.26108 | 0.0014798 | 1.8515 | 0.001552 | m | 0.00085176 | 0.00027278 | 0.0015566 | 0.0010159 | m | -0.262 | 0.0056785 | 1.85 | 0.0037473 | m | OK | -0.262 |  | 1.85 |  | m | -4.5635e-09 | 0.0056785 | 9.798e-09 | 0.0037473 | m | 6.587e-05 |  | -6.6684e-05 |  | m | 10 | Hz | 0.001 | -0.00026193 | 1.5048e-06 | 0.0018499 | 1.8549e-06 | 0.0018684 | 1.8486e-06 | 1.7115 | 0.00080949 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 9.6858e-08 | 0.000518 | 1e-07 | 0.0010271 | 10Hz3m3m | -1.0045e-07 | 0.0025811 | -1.8439e-07 | 0.0031799 | 1.1187e-08 | 0.0025811 | -2.3479e-08 | 0.0019618 | -0.00026192 | 0.0025811 | 0.0018499 | 0.0019618 | 1.0001 | 1.3075e-06 | 0.00010001 | 0.0035108 | 0.99997 | 2.6599e-06 | -6.1458e-05 | 0.0013962 | 1 | 2.6599e-06 | 2.9904e-05 | 0.0025927 | 0.0018685 | 3.6973e-06 | 1.7115 | 0.0030566 | -0.000262 | 5.6785e-06 | 0.00185 | 3.7473e-06
1 | m | 0.01 | -0.62862 | 0.0010536 | 0.75838 | 0.0011551 | m | 0.00043547 | 0.0001877 | -0.0016027 | 0.0017223 | m | -0.629 | 0.0049031 | 0.76 | 0.0045022 | m | OK | -0.629 |  | 0.76 |  | m | 5.688e-07 | 0.0049031 | 4.7307e-07 | 0.0045022 | m | -5.2803e-05 |  | -2.1966e-05 |  | m | 10 | mHz | 0.001 | -0.00062905 | 1.0702e-06 | 0.00075998 | 2.0738e-06 | 0.00098655 | 1.7372e-06 | 2.2622 | 0.0015795 | 1 | 0m | 1m | 0 | 0.001 | 10mHz1m0m | 10mHz1m1m | 1 | 19 | 10mHz1m0m | 0 | 0.0037411 | 0 | 0.00031167 | 10mHz1m1m | -1.7202e-08 | 0.0037075 | -5.6174e-08 | 0.0033974 | -1.6971e-08 | 0.0037075 | -5.6174e-08 | 0.0033559 | -0.00062907 | 0.0037075 | 0.00075992 | 0.0033559 | 1 | 3.7541e-06 | -5.618e-05 | 0.004811 | 0.99998 | 4.8914e-06 | -5.618e-05 | 0.004811 | 0.99998 | 4.8914e-06 | -5.618e-05 | 0.004811 | 0.00098653 | 3.4744e-06 | 2.2622 | 0.0057554 | -0.000629 | 4.9031e-06 | 0.00076 | 4.5022e-06
1 | m | 200 | -0.44986 | 0.0019089 | -0.85727 | 0.0015301 | m | -0.00075619 | 0.00030476 | -0.0012289 | 0.00045706 | m | -0.449 | 0.0043446 | -0.856 | 0.003651 | m | OK | -0.449 |  | -0.856 |  | m | 2.189e-06 | 0.0043446 | -1.106e-06 | 0.003651 | m | -0.0001026 |  | -4.4313e-05 |  | m | 200 | Hz | 0.001 | -0.0004491 | 1.9331e-06 | -0.00085604 | 1.5969e-06 | 0.0009667 | 1.6752e-06 | -2.0539 | 0.0019299 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 2.9562e-05 | 0 | 0.0014921 | 200Hz1m1m | -9.3119e-08 | 0.00051164 | 7.5225e-08 | 0.0013653 | -9.0018e-08 | 0.00051164 | 7.5225e-08 | 0.0013695 | -0.00044919 | 0.00051165 | -0.00085597 | 0.0013695 | 1 | 1.4924e-06 | 7.5234e-05 | 0.0027488 | 0.99991 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.99991 | 2.4461e-06 | 7.5234e-05 | 0.0027488 | 0.00096661 | 3.3504e-06 | -2.0539 | 0.0047386 | -0.000449 | 4.3446e-06 | -0.000856 | 3.651e-06
10 | m | 0.05 | 5.799 | 0.00024166 | 1.9178 | 0.00013334 | m | -0.0011779 | 0.00083548 | -0.0018375 | 0.00023272 | m | 5.8 | 0.0020411 | 1.92 | 0.0039526 | m | OK | 5.8 |  | 1.92 |  | m | 1.0054e-08 | 0.0020411 | -2.5074e-08 | 0.0039526 | m | 0.00018725 |  | -0.00032277 |  | m | 50 | mHz | 0.001 | 0.0058002 | 8.6973e-07 | 0.0019197 | 2.6821e-07 | 0.0061096 | 8.2997e-07 | 0.31962 | 6.1136e-05 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -4.4673e-08 | 0.00062352 | 3.0006e-07 | 0.0029111 | 50mHz10m10m | -9.4711e-08 | 0.0037153 | 9.476e-08 | 0.00093829 | -6.6901e-08 | 0.0037153 | 2.0886e-07 | 0.0020347 | 0.0058001 | 0.0037153 | 0.0019199 | 0.0020347 | 0.99999 | 3.0408e-06 | 9.999e-05 | 0.0009815 | 0.99999 | 4.0101e-06 | 9.4764e-06 | 0.00026758 | 0.99999 | 4.0101e-06 | 5.9781e-05 | 0.00066436 | 0.0061095 | 1.6601e-06 | 0.31968 | 0.00067552 | 0.0058 | 2.0411e-06 | 0.00192 | 3.9526e-06
3 | m | 500 | -1.0416 | 0.0010084 | -2.0315 | 0.001481 | m | -0.0014761 | 0.0016718 | -0.001494 | 0.00037848 | m | -1.04 | 0.0057724 | -2.03 | 0.0040675 | m | OK | -1.04 |  | -2.03 |  | m | -2.7844e-07 | 0.0057724 | -1.865e-07 | 0.0040675 | m | -0.00017264 |  | 8.0581e-07 |  | m | 500 | Hz | 0.001 | -0.0010402 | 1.9524e-06 | -0.00203 | 1.5286e-06 | 0.002281 | 1.6259e-06 | -2.0443 | 0.00082078 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | -9.1984e-08 | 0.0023561 | 9.9972e-08 | 0.0019901 | 500Hz3m3m | -5.275e-09 | 0.00090694 | 1.4777e-07 | 0.0036975 | -3.6448e-08 | 0.00090694 | 1.3058e-07 | 0.0030837 | -0.0010402 | 0.00090694 | -0.0020299 | 0.0030837 | 0.99991 | 4.4539e-06 | 9.9971e-05 | 0.0029075 | 1 | 4.0081e-06 | 4.9196e-05 | 0.0018092 | 0.99997 | 4.0081e-06 | 6.745e-05 | 0.002204 | 0.0022809 | 3.2518e-06 | -2.0442 | 0.0027482 | -0.00104 | 5.7724e-06 | -0.00203 | 4.0675e-06
10 | m | 0.02 | -3.1088 | 0.0015932 | 5.941 | 0.0013692 | m | 0.00089087 | 0.0016439 | 0.0012482 | 0.00055993 | m | -3.11 | 0.0050501 | 5.94 | 0.0039053 | m | OK | -3.11 |  | 5.94 |  | m | -2.8403e-08 | 0.0050501 | -2.1693e-08 | 0.0039053 | m | 0.00032927 |  | -0.00021984 |  | m | 20 | mHz | 0.001 | -0.0031097 | 2.2893e-06 | 0.0059398 | 1.4792e-06 | 0.0067046 | 1.6867e-06 | 2.0531 | 0.00031934 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | 1.1033e-07 | 0.00050493 | 2.9998e-07 | 0.00069903 | 20mHz10m10m | 6.5208e-07 | 0.0021735 | -3.5481e-07 | 0.0023208 | 3.9704e-07 | 0.0021735 | -4.6548e-08 | 0.0015573 | -0.0031093 | 0.0021735 | 0.0059397 | 0.0015573 | 1 | 2.0522e-06 | 0.0001 | 0.00076835 | 1.0001 | 3.2705e-06 | -3.5476e-05 | 0.00025274 | 1.0001 | 3.2705e-06 | 2.8305e-05 | 0.00049548 | 0.0067049 | 3.3734e-06 | 2.0531 | 0.00080835 | -0.00311 | 5.0501e-06 | 0.00594 | 3.9053e-06
1 | m | 0.01 | 0.088126 | 0.00055775 | -0.033755 | 0.0013026 | m | -0.00087404 | 0.0014662 | -0.0010557 | 0.00086454 | m | 0.088998 | 0.0031391 | -0.032705 | 0.0031582 | m | OK | 0.089 |  | -0.0327 |  | m | -1.663e-06 | 0.0031391 | -4.5258e-06 | 0.0031582 | m | 3.194e-07 |  | 4.2079e-07 |  | m | 10 | mHz | 0.001 | 8.9e-05 | 1.5687e-06 | -3.27e-05 | 1.5634e-06 | 9.4817e-05 | 1.568e-06 | -0.3521 | 0.016495 | 1 | 0m | 1m | 0 | 0.001 | 10mHz1m0m | 10mHz1m1m | 1 | 19 | 10mHz1m0m | 0 | 0.0037411 | 0 | 0.00031167 | 10mHz1m1m | -1.7202e-08 | 0.0037075 | -5.6174e-08 | 0.0033974 | -1.6311e-09 | 0.0037075 | -5.6174e-08 | 0.00060425 | 8.8999e-05 | 0.0037075 | -3.2756e-05 | 0.00060426 | 1 | 3.7541e-06 | -5.618e-05 | 0.004811 | 0.99998 | 4.8914e-06 | -5.618e-05 | 0.004811 | 1 | 4.8914e-06 | -5.618e-05 | 0.004811 | 9.4817e-05 | 3.1361e-06 | -0.35216 | 0.03334 | 8.8998e-05 | 3.1391e-06 | -3.2705e-05 | 3.1582e-06
1 | m | 2 | -0.69836 | 0.0016314 | -0.42309 | 0.0010502 | m | -0.00039485 | 0.00037907 | -0.0010963 | 0.0011844 | m | -0.698 | 0.003644 | -0.422 | 0.0041721 | m | OK | -0.698 |  | -0.422 |  | m | -1.5492e-06 | 0.003644 | 2.5565e-06 | 0.0041721 | m | 3.3946e-05 |  | 5.0791e-06 |  | m | 2 | Hz | 0.001 | -0.00069797 | 1.6749e-06 | -0.00042199 | 1.583e-06 | 0.00081562 | 1.6508e-06 | -2.5978 | 0.0019716 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.00082151 | 0 | 0.00085633 | 2Hz1m1m | 4.7626e-08 | 0.0021182 | -1.9869e-08 | 0.0034282 | 3.8844e-08 | 0.0021182 | -1.9869e-08 | 0.002954 | -0.00069793 | 0.0021182 | -0.00042201 | 0.002954 | 1 | 1.1867e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 0.00081565 | 3.3016e-06 | -2.5978 | 0.0054533 | -0.000698 | 3.644e-06 | -0.000422 | 4.1721e-06
10 | m | 0.05 | -5.5886 | 0.001317 | -2.8002 | 0.0017505 | m | 0.0016511 | 0.0013583 | -0.00053086 | 0.0016982 | m | -5.59 | 0.0045461 | -2.8 | 0.0058266 | m | OK | -5.59 |  | -2.8 |  | m | -1.3109e-08 | 0.0045461 | 2.2572e-08 | 0.0058266 | m | -0.00023146 |  | 0.00028923 |  | m | 50 | mHz | 0.001 | -0.0055902 | 1.8919e-06 | -0.0027997 | 2.4389e-06 | 0.0062521 | 2.0136e-06 | -2.6773 | 0.00037419 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -4.4673e-08 | 0.00062352 | 3.0006e-07 | 0.0029111 | 50mHz10m10m | -9.4711e-08 | 0.0037153 | 9.476e-08 | 0.00093829 | -6.792e-08 | 0.0037153 | 2.0468e-07 | 0.0019945 | -0.0055903 | 0.0037153 | -0.0027995 | 0.0019945 | 0.99999 | 3.0408e-06 | 9.999e-05 | 0.0009815 | 0.99999 | 4.0101e-06 | 9.4764e-06 | 0.00026758 | 0.99999 | 4.0101e-06 | 5.7938e-05 | 0.00064982 | 0.006252 | 4.0272e-06 | -2.6772 | 0.00099113 | -0.00559 | 4.5461e-06 | -0.0028 | 5.8266e-06
100 | m | 500 | -42.296 | 0.00022233 | 69.3071 | 0.0014173 | m | 0.00057801 | 0.0010012 | 0.00068351 | 0.00081248 | m | -42.3 | 0.0062904 | 69.3 | 0.0045202 | m | OK | -42.3 |  | 69.3 |  | m | 7.4922e-06 | 0.0062904 | -7.6776e-06 | 0.0045202 | m | 0.0034359 |  | 0.0064478 |  | m | 500 | Hz | 0.001 | -0.042297 | 1.0256e-06 | 0.069306 | 1.6337e-06 | 0.081193 | 1.4933e-06 | 2.1187 | 1.5037e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | 5.4485e-07 | 0.0013148 | 1.002e-06 | 6.2609e-05 | 500Hz100m100m | -7.2572e-06 | 0.0027372 | 7.1503e-06 | 0.0012973 | -5.6269e-06 | 0.0027372 | 5.8656e-06 | 0.0010393 | -0.042302 | 0.0027372 | 0.069312 | 0.0010393 | 1.0001 | 1.9945e-06 | 9.9996e-05 | 0.00019995 | 0.99993 | 4.7336e-06 | 7.1514e-05 | 5.1597e-05 | 0.99995 | 4.7336e-06 | 7.7465e-05 | 8.2597e-05 | 0.08119 | 3.0113e-06 | 2.1188 | 8.7902e-05 | -0.0423 | 6.2904e-06 | 0.0693 | 4.5202e-06
3 | m | 20 | 0.0028831 | 0.001627 | -1.17 | 0.0013456 | m | 0.0011365 | 0.00011274 | -5.2164e-05 | 0.0015467 | m | 0.00185 | 0.0063599 | -1.17 | 0.0041002 | m | OK | 0.00185 |  | -1.17 |  | m | 4.9118e-10 | 0.0063599 | -5.5669e-11 | 0.0041002 | m | -0.00010332 |  | 1.0908e-05 |  | m | 20 | Hz | 0.001 | 1.7467e-06 | 1.6309e-06 | -0.00117 | 2.0501e-06 | 0.00117 | 2.0501e-06 | -1.5693 | 0.0013939 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | 5.1085e-09 | 0.0037612 | 9.9997e-08 | 0.0032193 | 20Hz3m3m | 1.6919e-07 | 0.0020234 | -1.1325e-07 | 0.0010612 | 1.9054e-08 | 0.0020234 | 8.1872e-08 | 0.0030359 | 1.7657e-06 | 0.0020234 | -0.0011699 | 0.0030359 | 1 | 4.4939e-06 | 0.00010001 | 0.005064 | 1.0001 | 3.9936e-06 | -3.7756e-05 | 0.00038761 | 1 | 3.9936e-06 | 8.8296e-05 | 0.0046665 | 0.00117 | 4.1002e-06 | -1.5692 | 0.0054358 | 1.85e-06 | 6.3599e-06 | -0.00117 | 4.1002e-06
10 | m | 0.02 | -5.4714 | 0.0018145 | -6.9913 | 9.9073e-05 | m | -0.0018751 | 0.00098548 | -0.001633 | 0.00018126 | m | -5.47 | 0.0038215 | -6.99 | 0.0033905 | m | OK | -5.47 |  | -6.99 |  | m | 5.0711e-08 | 0.0038215 | -2.5997e-08 | 0.0033905 | m | 0.0004276 |  | 0.0003489 |  | m | 20 | mHz | 0.001 | -0.0054696 | 2.0649e-06 | -0.0069897 | 2.0657e-07 | 0.0088753 | 1.2829e-06 | -2.2348 | 0.00018378 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | 1.1033e-07 | 0.00050493 | 2.9998e-07 | 0.00069903 | 20mHz10m10m | 6.5208e-07 | 0.0021735 | -3.5481e-07 | 0.0023208 | 5.6504e-07 | 0.0021735 | -2.4961e-07 | 0.0020602 | -0.005469 | 0.0021735 | -0.0069899 | 0.0020602 | 1 | 2.0522e-06 | 0.0001 | 0.00076835 | 1.0001 | 3.2705e-06 | -3.5476e-05 | 0.00025274 | 1.0001 | 3.2705e-06 | -1.3709e-05 | 0.00033558 | 0.0088759 | 2.5659e-06 | -2.2348 | 0.00049771 | -0.00547 | 3.8215e-06 | -0.00699 | 3.3905e-06
10 | m | 20 | 0.82749 | 0.001581 | 5.4688 | 0.00035324 | m | -0.0018986 | 0.00017209 | -0.00097645 | 0.0014098 | m | 0.829 | 0.0041215 | 5.47 | 0.0029466 | m | OK | 0.829 |  | 5.47 |  | m | 6.793e-08 | 0.0041215 | 7.4923e-08 | 0.0029466 | m | 0.00038993 |  | -0.00020436 |  | m | 20 | Hz | 0.001 | 0.00082939 | 1.5903e-06 | 0.0054698 | 1.4534e-06 | 0.0055323 | 1.4566e-06 | 1.4203 | 0.00028693 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | -2.0254e-08 | 0.0038673 | 3.0005e-07 | 0.0014224 | 20Hz10m10m | 8.368e-07 | 0.00284 | 3.1494e-07 | 0.00021184 | 2.8979e-07 | 0.00284 | 3.0543e-07 | 0.00098448 | 0.00082968 | 0.00284 | 0.0054701 | 0.00098449 | 0.99999 | 5.4376e-06 | 9.9993e-05 | 0.00064666 | 1.0001 | 2.9017e-06 | 3.1488e-05 | 0.00019065 | 1 | 2.9017e-06 | 7.5211e-05 | 0.00048169 | 0.0055325 | 2.9133e-06 | 1.4204 | 0.00074923 | 0.000829 | 4.1215e-06 | 0.00547 | 2.9466e-06
10 | m | 2 | -2.3083 | 0.0011856 | -2.89 | 0.00030375 | m | 0.001881 | 0.00052812 | -0.00034517 | 0.0008331 | m | -2.31 | 0.0043667 | -2.89 | 0.0037181 | m | OK | -2.31 |  | -2.89 |  | m | -6.2434e-09 | 0.0043667 | 8.0895e-09 | 0.0037181 | m | -0.00019997 |  | 0.00031837 |  | m | 2 | Hz | 0.001 | -0.0023102 | 1.2979e-06 | -0.0028897 | 8.8674e-07 | 0.0036996 | 1.0661e-06 | -2.2452 | 0.00031223 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.0342e-07 | 0.0024993 | 3.0003e-07 | 0.0035924 | 2Hz10m10m | 2.4445e-07 | 0.0013271 | 5.9488e-07 | 0.0020341 | 1.1752e-07 | 0.0013271 | 3.295e-07 | 0.0034367 | -0.0023101 | 0.0013271 | -0.0028894 | 0.0034367 | 1 | 2.9655e-06 | 9.9997e-05 | 0.00141 | 1 | 1.3475e-06 | 5.9483e-05 | 0.00027487 | 1 | 1.3475e-06 | 9.5948e-05 | 0.0012965 | 0.0036998 | 2.1322e-06 | -2.2451 | 0.001439 | -0.00231 | 4.3667e-06 | -0.00289 | 3.7181e-06
3 | m | 0.2 | 2.1794 | 0.0019542 | 0.68048 | 0.0012966 | m | -0.00077673 | 0.00044814 | -0.00046258 | 0.0014039 | m | 2.18 | 0.0040503 | 0.681 | 0.0045423 | m | OK | 2.18 |  | 0.681 |  | m | -5.407e-09 | 0.0040503 | 1.9695e-08 | 0.0045423 | m | 0.00021478 |  | -5.9544e-05 |  | m | 200 | mHz | 0.001 | 0.0021802 | 2.0049e-06 | 0.00068094 | 1.9111e-06 | 0.0022841 | 1.9967e-06 | 0.30273 | 0.00084043 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | -8.0344e-08 | 0.0021164 | 1.0001e-07 | 0.0014907 | 200mHz3m3m | -2.4867e-07 | 0.00017664 | 8.0071e-08 | 0.0026185 | -1.8842e-07 | 0.00017664 | 8.7207e-08 | 0.0022148 | 0.00218 | 0.00017668 | 0.00068103 | 0.0022148 | 0.99992 | 2.2511e-06 | 0.0001 | 0.0014992 | 0.99992 | 4.7804e-07 | 2.6686e-05 | 0.000883 | 0.99992 | 4.7804e-07 | 5.2931e-05 | 0.0011036 | 0.0022839 | 3.9934e-06 | 0.30278 | 0.0020108 | 0.00218 | 4.0503e-06 | 0.000681 | 4.5423e-06
10 | m | 5 | 1.6415 | 0.0015585 | 7.4707 | 0.0015157 | m | 0.0012917 | 0.0011649 | 0.00069755 | 0.0010528 | m | 1.64 | 0.0047416 | 7.47 | 0.0037573 | m | OK | 1.64 |  | 7.47 |  | m | -8.2924e-08 | 0.0047416 | 1.8884e-08 | 0.0037573 | m | 0.00021627 |  | 2.5793e-05 |  | m | 5 | Hz | 0.001 | 0.0016402 | 1.9457e-06 | 0.00747 | 1.8454e-06 | 0.007648 | 1.8502e-06 | 1.3547 | 0.00025382 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -2.6382e-08 | 0.00058565 | 2.9998e-07 | 0.00064819 | 5Hz10m10m | -9.645e-08 | 0.003922 | -1.008e-07 | 0.0028868 | -7.2907e-08 | 0.003922 | 3.3867e-08 | 0.0021346 | 0.0016401 | 0.003922 | 0.0074701 | 0.0021346 | 0.99999 | 1.3944e-06 | 0.0001 | 0.00029063 | 0.99999 | 3.9336e-06 | -1.008e-05 | 0.00040408 | 0.99999 | 3.9336e-06 | 2.6908e-05 | 0.00036596 | 0.0076479 | 3.7005e-06 | 1.3547 | 0.0006258 | 0.00164 | 4.7416e-06 | 0.00747 | 3.7573e-06
1 | m | 2 | 0.20757 | 0.0014342 | -0.04997 | 0.0015393 | m | 0.0015765 | 0.0018621 | -0.00047127 | 0.00085217 | m | 0.206 | 0.0045967 | -0.049503 | 0.00374 | m | OK | 0.206 |  | -0.0495 |  | m | -7.7494e-07 | 0.0045967 | -3.226e-06 | 0.00374 | m | -1.8698e-06 |  | 1.3664e-06 |  | m | 2 | Hz | 0.001 | 0.000206 | 2.3504e-06 | -4.9499e-05 | 1.7594e-06 | 0.00021186 | 2.322e-06 | -0.23582 | 0.0084806 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.00082151 | 0 | 0.00085633 | 2Hz1m1m | 4.7626e-08 | 0.0021182 | -1.9869e-08 | 0.0034282 | 1.009e-08 | 0.0021182 | -1.9869e-08 | 0.0014012 | 0.00020601 | 0.0021182 | -4.9519e-05 | 0.0014012 | 1 | 1.1867e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669 | 0.00021186 | 4.6441e-06 | -0.23584 | 0.017374 | 0.000206 | 4.5967e-06 | -4.9503e-05 | 3.74e-06
3 | m | 5 | 2.929 | 0.0012499 | 0.41904 | 0.0013866 | m | -0.00073766 | 0.00066391 | 2.64e-05 | 0.0010952 | m | 2.93 | 0.0028782 | 0.419 | 0.0041229 | m | OK | 2.93 |  | 0.419 |  | m | -6.5022e-08 | 0.0028782 | 4.1599e-08 | 0.0041229 | m | -0.00027745 |  | 1.7884e-05 |  | m | 5 | Hz | 0.001 | 0.0029297 | 1.4153e-06 | 0.00041902 | 1.7669e-06 | 0.0029595 | 1.4232e-06 | 0.14206 | 0.00059488 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | -4.9337e-08 | 0.00079014 | 1.0001e-07 | 0.0022765 | 5Hz3m3m | 2.8449e-07 | 0.0026935 | -6.5054e-08 | 0.0018253 | 2.7773e-07 | 0.0026935 | -6.1715e-08 | 0.0018344 | 0.00293 | 0.0026935 | 0.00041896 | 0.0018344 | 0.99995 | 2.7916e-06 | 9.9998e-05 | 0.0031263 | 1.0001 | 3.03e-06 | -2.1698e-05 | 0.00068935 | 1.0001 | 3.03e-06 | -1.9236e-05 | 0.00073866 | 0.0029598 | 2.8464e-06 | 0.14204 | 0.0014004 | 0.00293 | 2.8782e-06 | 0.000419 | 4.1229e-06
10 | m | 0.2 | 5.8321 | 0.0017578 | -0.98688 | 0.00070699 | m | 0.001971 | 0.00073121 | 0.00034244 | 0.0011275 | m | 5.83 | 0.0038884 | -0.987 | 0.0065702 | m | OK | 5.83 |  | -0.987 |  | m | -9.9498e-09 | 0.0038884 | -9.1631e-08 | 0.0065702 | m | 0.0001643 |  | -0.00022001 |  | m | 200 | mHz | 0.001 | 0.0058302 | 1.9038e-06 | -0.00098722 | 1.3308e-06 | 0.0059132 | 1.8902e-06 | -0.16774 | 0.00022831 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | -1.1265e-07 | 0.0014165 | 3.0001e-07 | 0.0030477 | 200mHz10m10m | -2.8064e-07 | 0.0018077 | -6.3346e-07 | 0.0023635 | -1.8256e-07 | 0.0018077 | -8.8469e-08 | 0.002763 | 0.00583 | 0.0018077 | -0.00098731 | 0.002763 | 0.99996 | 3.0131e-06 | 0.00010001 | 0.0015782 | 0.99997 | 3.7919e-06 | -6.3333e-05 | 0.00024795 | 0.99997 | 3.7919e-06 | 3.2031e-05 | 0.0010246 | 0.005913 | 3.7805e-06 | -0.16771 | 0.0011218 | 0.00583 | 3.8884e-06 | -0.000987 | 6.5702e-06
100 | m | 2000 | 23.0988 | 0.002148 | -2.7129 | 0.0013979 | m | 9.0904e-05 | 0.00020103 | -0.0011811 | 0.0019026 | m | 23.1 | 0.004465 | -2.71 | 0.0105 | m | OK | 23.1 |  | -2.71 |  | m | -9.3084e-06 | 0.004465 | -7.0196e-06 | 0.0105 | m | -0.00125 |  | -0.0017103 |  | m | 2 | kHz | 0.001 | 0.023099 | 2.1574e-06 | -0.0027117 | 2.3609e-06 | 0.023257 | 2.1603e-06 | -0.11686 | 0.0001014 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | 6.6578e-07 | 0.0026793 | 1.0116e-06 | 0.002223 | 2kHz100m100m | -8.3635e-06 | 0.0019217 | -4.2838e-06 | 0.0023977 | -6.6427e-07 | 0.0019217 | 2.3157e-07 | 0.0022488 | 0.023098 | 0.0019217 | -0.0027115 | 0.0022488 | 1.0001 | 2.6951e-06 | 9.9849e-05 | 0.00044188 | 0.99992 | 3.9669e-06 | -4.2013e-05 | 0.00019973 | 1 | 3.9669e-06 | 7.8953e-05 | 0.00040621 | 0.023258 | 4.3216e-06 | -0.11678 | 0.00045402 | 0.0231 | 4.465e-06 | -0.00271 | 1.05e-05
10 | m | 5 | -3.6903 | 0.001438 | 2.991 | 0.00061845 | m | -0.00052283 | 0.0012338 | 0.00069665 | 0.0018852 | m | -3.69 | 0.0039917 | 2.99 | 0.0040582 | m | OK | -3.69 |  | 2.99 |  | m | -2.1105e-08 | 0.0039917 | -2.5635e-08 | 0.0040582 | m | 0.00018352 |  | 0.00029439 |  | m | 5 | Hz | 0.001 | -0.0036898 | 1.8948e-06 | 0.0029903 | 1.9841e-06 | 0.0047494 | 1.9307e-06 | 2.4605 | 0.0004104 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -2.6382e-08 | 0.00058565 | 2.9998e-07 | 0.00064819 | 5Hz10m10m | -9.645e-08 | 0.003922 | -1.008e-07 | 0.0028868 | -4.3893e-08 | 0.003922 | 1.9982e-07 | 0.0012077 | -0.0036899 | 0.003922 | 0.0029905 | 0.0012077 | 0.99999 | 1.3944e-06 | 0.0001 | 0.00029063 | 0.99999 | 3.9336e-06 | -1.008e-05 | 0.00040408 | 0.99999 | 3.9336e-06 | 7.2493e-05 | 0.00031898 | 0.0047493 | 3.8614e-06 | 2.4606 | 0.00088061 | -0.00369 | 3.9917e-06 | 0.00299 | 4.0582e-06
1 | m | 100 | -0.093142 | 0.0010185 | 0.057605 | 0.00067755 | m | -0.00064218 | 0.00052253 | -0.000494 | 0.0012051 | m | -0.092497 | 0.0024987 | 0.058105 | 0.0025958 | m | OK | -0.0925 |  | 0.0581 |  | m | 3.0412e-06 | 0.0024987 | 4.8418e-06 | 0.0025958 | m | 3.1279e-07 |  | -1.0244e-06 |  | m | 100 | Hz | 0.001 | -9.25e-05 | 1.1447e-06 | 5.8099e-05 | 1.3825e-06 | 0.00010923 | 1.2167e-06 | 2.5808 | 0.01208 | 1 | 0m | 1m | 0 | 0.001 | 100Hz1m0m | 100Hz1m1m | 13 | 31 | 100Hz1m0m | 0 | 0.0026016 | 0 | 4.2522e-05 | 100Hz1m1m | 6.7895e-08 | 0.0031842 | -5.8743e-08 | 0.0016459 | 7.4163e-09 | 0.0031842 | -5.8743e-08 | 0.00021767 | -9.2492e-05 | 0.0031842 | 5.804e-05 | 0.00021768 | 1 | 2.602e-06 | -5.8762e-05 | 0.0028252 | 1.0001 | 4.8143e-06 | -5.8762e-05 | 0.0028252 | 1 | 4.8143e-06 | -5.8762e-05 | 0.0028252 | 0.00010923 | 2.4334e-06 | 2.5807 | 0.024325 | -9.2497e-05 | 2.4987e-06 | 5.8105e-05 | 2.5958e-06
3 | m | 500 | 0.82297 | 0.0017848 | -1.7209 | 0.00027488 | m | 0.0010587 | 0.0010557 | -0.00077242 | 0.00086581 | m | 0.822 | 0.0054927 | -1.72 | 0.003381 | m | OK | 0.822 |  | -1.72 |  | m | -1.3796e-08 | 0.0054927 | -2.3164e-07 | 0.003381 | m | -9.0524e-05 |  | -0.0001511 |  | m | 500 | Hz | 0.001 | 0.00082191 | 2.0737e-06 | -0.0017202 | 9.084e-07 | 0.0019064 | 1.2129e-06 | -1.1251 | 0.0010027 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | -9.1984e-08 | 0.0023561 | 9.9972e-08 | 0.0019901 | 500Hz3m3m | -5.275e-09 | 0.00090694 | 1.4777e-07 | 0.0036975 | -5.2686e-08 | 0.00090694 | 1.2163e-07 | 0.0027639 | 0.00082186 | 0.00090695 | -0.00172 | 0.0027639 | 0.99991 | 4.4539e-06 | 9.9971e-05 | 0.0029075 | 1 | 4.0081e-06 | 4.9196e-05 | 0.0018092 | 0.99995 | 4.0081e-06 | 7.6959e-05 | 0.0024097 | 0.0019063 | 2.4258e-06 | -1.125 | 0.003135 | 0.000822 | 5.4927e-06 | -0.00172 | 3.381e-06
100 | m | 0.2 | -7.5508 | 3.1009e-05 | -10.9002 | 0.00034197 | m | -0.00038687 | 0.001315 | -0.001798 | 0.001511 | m | -7.55 | 0.004975 | -10.9 | 0.0040521 | m | OK | -7.55 |  | -10.9 |  | m | -2.0888e-08 | 0.004975 | 6.4801e-08 | 0.0040521 | m | -0.00042342 |  | 0.0015812 |  | m | 200 | mHz | 0.001 | -0.0075504 | 1.3154e-06 | -0.010898 | 1.5492e-06 | 0.013258 | 1.4774e-06 | -2.1767 | 0.00010525 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | 8.5582e-07 | 0.0026992 | 1e-06 | 0.0016788 | 200mHz100m100m | -7.227e-06 | 0.0047385 | -6.1333e-06 | 0.0037133 | 5.6319e-07 | 0.0047385 | 7.4177e-07 | 0.0017525 | -0.0075499 | 0.0047385 | -0.010898 | 0.0017525 | 1.0001 | 2.8474e-06 | 9.9996e-05 | 0.00038612 | 0.99993 | 5.2098e-06 | -6.1324e-05 | 6.2289e-05 | 1.0001 | 5.2098e-06 | 9.4156e-05 | 0.00037439 | 0.013259 | 2.9557e-06 | -2.1766 | 0.00042952 | -0.00755 | 4.975e-06 | -0.0109 | 4.0521e-06
3 | m | 10 | 2.7712 | 0.00035229 | 0.49071 | 0.0010615 | m | 0.0011759 | 0.00047507 | 0.00057123 | 0.0016496 | m | 2.77 | 0.0016831 | 0.49 | 0.0058361 | m | OK | 2.77 |  | 0.49 |  | m | -1.6273e-08 | 0.0016831 | -3.6023e-08 | 0.0058361 | m | 3.6253e-05 |  | 0.00013884 |  | m | 10 | Hz | 0.001 | 0.00277 | 5.9143e-07 | 0.00049014 | 1.9616e-06 | 0.0028131 | 6.7527e-07 | 0.17513 | 0.00068764 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 9.6858e-08 | 0.000518 | 1e-07 | 0.0010271 | 10Hz3m3m | -1.0045e-07 | 0.0025811 | -1.8439e-07 | 0.0031799 | -8.2011e-08 | 0.0025811 | -1.5781e-07 | 0.0029787 | 0.00277 | 0.0025811 | 0.00048998 | 0.0029787 | 1.0001 | 1.3075e-06 | 0.00010001 | 0.0035108 | 0.99997 | 2.6599e-06 | -6.1458e-05 | 0.0013962 | 0.99998 | 2.6599e-06 | -4.6366e-05 | 0.0015938 | 0.002813 | 1.3506e-06 | 0.17508 | 0.0021052 | 0.00277 | 1.6831e-06 | 0.00049 | 5.8361e-06
100 | m | 1000 | -31.2983 | 0.00095256 | -16.399 | -4.4275e-05 | m | 0.0017801 | 0.0010768 | -0.00074914 | 0.00081105 | m | -31.3 | 0.0055764 | -16.4 | 0.0097262 | m | OK | -31.3 |  | -16.4 |  | m | -1.822e-06 | 0.0055764 | 9.7753e-06 | 0.0097262 | m | -7.8428e-05 |  | 0.001727 |  | m | 1 | kHz | 0.001 | -0.0313 | 1.4377e-06 | -0.016398 | 8.1226e-07 | 0.035336 | 1.3281e-06 | -2.659 | 2.7769e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | 3.7851e-07 | 0.0038469 | 1.0036e-06 | 0.00040995 | 1kHz100m100m | -2.3449e-06 | 0.0026084 | -9.8959e-06 | 0.0027654 | -3.8816e-07 | 0.0026084 | -2.0647e-06 | 0.001073 | -0.0313 | 0.0026084 | -0.0164 | 0.001073 | 1 | 3.9272e-06 | 9.9925e-05 | 0.00037989 | 0.99998 | 2.7549e-06 | -9.8552e-05 | 0.00010743 | 1 | 2.7549e-06 | 4.4053e-05 | 0.00030319 | 0.035336 | 2.658e-06 | -2.659 | 0.00030823 | -0.0313 | 5.5764e-06 | -0.0164 | 9.7262e-06
1 | m | 0.05 | 0.056926 | 0.0014642 | -0.74728 | 0.0019912 | m | -0.00029105 | 0.0016972 | 0.00073002 | 0.00044934 | m | 0.057195 | 0.0056586 | -0.748 | 0.004096 | m | OK | 0.0572 |  | -0.748 |  | m | -5.4338e-06 | 0.0056586 | -4.1764e-07 | 0.004096 | m | 1.7494e-05 |  | -1.4e-05 |  | m | 50 | mHz | 0.001 | 5.7217e-05 | 2.2415e-06 | -0.00074801 | 2.0413e-06 | 0.0007502 | 2.0425e-06 | -1.4945 | 0.0029864 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.00036442 | 0 | 0.0010831 | 50mHz1m1m | -2.7165e-08 | 0.0017934 | -2.9092e-08 | 0.003713 | -2.0379e-08 | 0.0017934 | -2.9092e-08 | 0.0030561 | 5.7197e-05 | 0.0017934 | -0.00074804 | 0.0030561 | 1 | 1.1427e-06 | -2.9093e-05 | 0.0046238 | 0.99997 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.99998 | 3.3597e-06 | -2.9093e-05 | 0.0046238 | 0.00075018 | 4.085e-06 | -1.4945 | 0.0075534 | 5.7195e-05 | 5.6586e-06 | -0.000748 | 4.096e-06
10 | m | 20 | -0.41696 | 0.0018086 | -6.4487 | 0.0015693 | m | 0.00045085 | 0.00063073 | 0.0010089 | 0.00072354 | m | -0.417 | 0.0046936 | -6.45 | 0.0034638 | m | OK | -0.417 |  | -6.45 |  | m | -7.9459e-08 | 0.0046936 | -1.0408e-07 | 0.0034638 | m | -0.00041057 |  | 0.00027254 |  | m | 20 | Hz | 0.001 | -0.00041741 | 1.9155e-06 | -0.0064497 | 1.728e-06 | 0.0064632 | 1.7289e-06 | -1.6354 | 0.00029625 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | -2.0254e-08 | 0.0038673 | 3.0005e-07 | 0.0014224 | 20Hz10m10m | 8.368e-07 | 0.00284 | 3.1494e-07 | 0.00021184 | 4.0377e-07 | 0.00284 | 3.0741e-07 | 0.00082349 | -0.00041701 | 0.00284 | -0.0064494 | 0.0008235 | 0.99999 | 5.4376e-06 | 9.9993e-05 | 0.00064666 | 1.0001 | 2.9017e-06 | 3.1488e-05 | 0.00019065 | 1 | 2.9017e-06 | 6.61e-05 | 0.00042105 | 0.0064635 | 3.4578e-06 | -1.6354 | 0.00072687 | -0.000417 | 4.6936e-06 | -0.00645 | 3.4638e-06
100 | m | 0.05 | -23.5025 | 0.0002908 | 7.0733 | 0.0010139 | m | -0.0015735 | 0.0011167 | 0.00044316 | 0.0013281 | m | -23.5 | 0.0029311 | 7.07 | 0.0060362 | m | OK | -23.5 |  | 7.07 |  | m | 1.0397e-10 | 0.0029311 | -3.4425e-10 | 0.0060362 | m | -0.00090233 |  | 0.0028177 |  | m | 50 | mHz | 0.001 | -0.023501 | 1.1539e-06 | 0.0070728 | 1.6709e-06 | 0.024542 | 1.2053e-06 | 2.8493 | 6.6586e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -6.8791e-07 | 0.001896 | 1e-06 | 0.00092261 | 50mHz100m100m | -6.5693e-06 | 0.0036121 | 9.5977e-06 | 0.001982 | -1.6382e-06 | 0.0036121 | 2.3892e-06 | 0.0010938 | -0.023503 | 0.0036121 | 0.0070752 | 0.0010938 | 0.99993 | 1.9425e-06 | 0.0001 | 0.00024812 | 0.99993 | 3.8821e-06 | 9.5973e-05 | 5.9466e-05 | 0.99993 | 3.8821e-06 | 9.9349e-05 | 0.00021764 | 0.02454 | 2.4125e-06 | 2.8494 | 0.00025515 | -0.0235 | 2.9311e-06 | 0.00707 | 6.0362e-06
100 | m | 0.1 | -33.4989 | 0.0016713 | 80.494 | 0.001869 | m | 0.0012373 | 0.0013511 | 0.0012473 | 0.0018522 | m | -33.5 | 0.0067698 | 80.5 | 0.0054642 | m | OK | -33.5 |  | 80.5 |  | m | 1.3277e-07 | 0.0067698 | 5.197e-07 | 0.0054642 | m | -0.00016086 |  | -0.0072077 |  | m | 100 | mHz | 0.001 | -0.0335 | 2.1491e-06 | 0.080493 | 2.6313e-06 | 0.087186 | 2.5658e-06 | 1.9652 | 2.5542e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 100mHz100m10m | 100mHz100m100m | 112 | 130 | 100mHz100m10m | -3.2234e-07 | 0.0035671 | 1.0001e-06 | 0.00099585 | 100mHz100m100m | 9.3522e-06 | 0.0047814 | -5.5623e-06 | 0.0011289 | 7.9747e-06 | 0.0047814 | -4.628e-06 | 0.00111 | -0.033492 | 0.0047814 | 0.080488 | 0.00111 | 0.99997 | 3.8623e-06 | 9.9999e-05 | 0.00012488 | 1.0001 | 4.8348e-06 | -5.5628e-05 | 5.17e-05 | 1.0001 | 4.8348e-06 | -3.347e-05 | 6.2119e-05 | 0.087192 | 5.1489e-06 | 1.9651 | 8.0426e-05 | -0.0335 | 6.7698e-06 | 0.0805 | 5.4642e-06
10 | m | 1 | 0.46561 | 3.5147e-05 | 5.7586 | 0.0001188 | m | 0.0012076 | 0.0019994 | -0.0013185 | 8.4218e-05 | m | 0.464 | 0.0053408 | 5.76 | 0.00060983 | m | OK | 0.464 |  | 5.76 |  | m | -3.7121e-09 | 0.0053408 | 8.7417e-09 | 0.00060983 | m | 0.00039754 |  | -0.0001252 |  | m | 1 | Hz | 0.001 | 0.0004644 | 1.9997e-06 | 0.0057599 | 1.4562e-07 | 0.0057786 | 2.1656e-07 | 1.4903 | 0.00034495 | 3 | 3m | 10m | 0.003 | 0.01 | 1Hz10m3m | 1Hz10m10m | 79 | 97 | 1Hz10m3m | 1.1918e-07 | 0.0012472 | 2.9995e-07 | 0.0010183 | 1Hz10m10m | -1.9857e-07 | 0.0035657 | 2.5208e-07 | 0.00099493 | -6.947e-09 | 0.0035657 | 2.8095e-07 | 0.001009 | 0.00046439 | 0.0035657 | 0.0057602 | 0.001009 | 1 | 3.4605e-06 | 0.0001 | 0.00080072 | 0.99998 | 3.5806e-06 | 2.5206e-05 | 0.00034411 | 1 | 3.5806e-06 | 7.0314e-05 | 0.00061948 | 0.0057787 | 4.3361e-07 | 1.4904 | 0.0009272 | 0.000464 | 5.3408e-06 | 0.00576 | 6.0983e-07
100 | m | 1000 | -52.1058 | 0.0015293 | 74.1977 | 0.0024688 | m | 0.00080757 | 0.00027769 | 0.00044749 | 0.0012707 | m | -52.0999 | 0.010973 | 74.2 | 0.0084549 | m | OK | -52.1 |  | 74.2 |  | m | 6.2598e-05 | 0.010973 | 2.0169e-05 | 0.0084549 | m | -0.0066208 |  | -0.002763 |  | m | 1 | kHz | 0.001 | -0.052107 | 1.5543e-06 | 0.074197 | 2.7767e-06 | 0.090666 | 2.4416e-06 | 2.183 | 2.2508e-05 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | 3.7851e-07 | 0.0038469 | 1.0036e-06 | 0.00040995 | 1kHz100m100m | -2.3449e-06 | 0.0026084 | -9.8959e-06 | 0.0027654 | -2.0625e-06 | 0.0026084 | -8.7655e-06 | 0.0025211 | -0.052109 | 0.0026084 | 0.074188 | 0.0025211 | 1 | 3.9272e-06 | 9.9925e-05 | 0.00037989 | 0.99998 | 2.7549e-06 | -9.8552e-05 | 0.00010743 | 0.99998 | 2.7549e-06 | -7.7968e-05 | 0.00013568 | 0.090664 | 4.8896e-06 | 2.183 | 0.00014296 | -0.0521 | 1.0973e-05 | 0.0742 | 8.4549e-06
10 | m | 20 | -0.16533 | 0.0018411 | 3.0282 | 0.00024502 | m | -0.00062919 | 0.0010165 | -0.001864 | 0.00053398 | m | -0.165 | 0.0046269 | 3.03 | 0.0012199 | m | OK | -0.165 |  | 3.03 |  | m | 1.9986e-08 | 0.0046269 | 2.972e-08 | 0.0012199 | m | 0.00030094 |  | 3.5509e-05 |  | m | 20 | Hz | 0.001 | -0.0001647 | 2.1031e-06 | 0.00303 | 5.8752e-07 | 0.0030345 | 5.9765e-07 | 1.6251 | 0.00069211 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | -2.0254e-08 | 0.0038673 | 3.0005e-07 | 0.0014224 | 20Hz10m10m | 8.368e-07 | 0.00284 | 3.1494e-07 | 0.00021184 | -1.6029e-08 | 0.00284 | 3.0012e-07 | 0.0014165 | -0.00016472 | 0.00284 | 0.0030303 | 0.0014165 | 0.99999 | 5.4376e-06 | 9.9993e-05 | 0.00064666 | 1.0001 | 2.9017e-06 | 3.1488e-05 | 0.00019065 | 0.99999 | 2.9017e-06 | 9.9655e-05 | 0.00064441 | 0.0030345 | 1.1953e-06 | 1.6252 | 0.0015269 | -0.000165 | 4.6269e-06 | 0.00303 | 1.2199e-06

## Cal Data
Calibration data: measured values of reference standards.
[9 empty] | Measured calibration Z |  |  |  | Measured residual SHORT Z | [13 empty] | Reference Z (auto search from Ref Z sheet) | [9 empty] | Absolute correction |  |  |  |  | Gain-phase correction
Range |  | Range Z | f | f |  | Nominal Z |  | Nom Z | Rs | ua(Rs) | Xs | ua(Xs) | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | U(Zr) | phi | U(phir) | raf tag | ref ID | Rs | U(Rs) | Xs | U(Xs) | Zr | U(Zr) | phir | U(phir) | tag | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi))
Ohm |  | Ohm | Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | Ohm | Ohm | Ohm | Ohm | - | - | rad | rad
1 | m | 0.001 | 0.01 | 10 | mHz | 0 | m | 0 |  | 0.0018706 |  | 0.00015584 |  |  |  |  | m | 0.001 | 0 | 1.8706e-06 | 0 | 1.5584e-07 | 0 | 1.877e-06 | #DIV/0! | 0 | 10mHz0m | 1 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10mHz1m0m | 0 | 0.0037411 | 0 | 0.00031167 | 1 | 3.7541e-06 | -5.618e-05 | 0.004811
1 | m | 0.001 | 0.02 | 20 | mHz | 0 | m | 0 |  | 0.0012832 |  | 0.0014941 |  |  |  |  | m | 0.001 | 0 | 1.2832e-06 | 0 | 1.4941e-06 | 0 | 1.9695e-06 | #DIV/0! | 0 | 20mHz0m | 2 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20mHz1m0m | 0 | 0.0025663 | 0 | 0.0029883 | 1 | 3.939e-06 | -5.406e-05 | 0.0029495
1 | m | 0.001 | 0.05 | 50 | mHz | 0 | m | 0 |  | 0.00018221 |  | 0.00054154 |  |  |  |  | m | 0.001 | 0 | 1.8221e-07 | 0 | 5.4154e-07 | 0 | 5.7137e-07 | #DIV/0! | 0 | 50mHz0m | 3 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50mHz1m0m | 0 | 0.00036442 | 0 | 0.0010831 | 1 | 1.1427e-06 | -2.9093e-05 | 0.0046238
1 | m | 0.001 | 0.1 | 100 | mHz | 0 | m | 0 |  | 0.00040334 |  | 0.00032998 |  |  |  |  | m | 0.001 | 0 | 4.0334e-07 | 0 | 3.2998e-07 | 0 | 5.2112e-07 | #DIV/0! | 0 | 100mHz0m | 4 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100mHz1m0m | 0 | 0.00080668 | 0 | 0.00065996 | 1 | 1.0422e-06 | 7.4137e-06 | 0.0020973
1 | m | 0.001 | 0.2 | 200 | mHz | 0 | m | 0 |  | 0.0014608 |  | 0.0017802 |  |  |  |  | m | 0.001 | 0 | 1.4608e-06 | 0 | 1.7802e-06 | 0 | 2.3028e-06 | #DIV/0! | 0 | 200mHz0m | 5 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200mHz1m0m | 0 | 0.0029215 | 0 | 0.0035605 | 1 | 4.6057e-06 | 9.7461e-05 | 0.0039192
1 | m | 0.001 | 0.5 | 500 | mHz | 0 | m | 0 |  | 0.00042258 |  | 0.0018034 |  |  |  |  | m | 0.001 | 0 | 4.2258e-07 | 0 | 1.8034e-06 | 0 | 1.8522e-06 | #DIV/0! | 0 | 500mHz0m | 6 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500mHz1m0m | 0 | 0.00084516 | 0 | 0.0036068 | 1 | 3.7045e-06 | -5.2578e-05 | 0.0015719
1 | m | 0.001 | 1 | 1 | Hz | 0 | m | 0 |  | 0.001338 |  | 0.00091342 |  |  |  |  | m | 0.001 | 0 | 1.338e-06 | 0 | 9.1342e-07 | 0 | 1.62e-06 | #DIV/0! | 0 | 1Hz0m | 7 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1Hz1m0m | 0 | 0.0026759 | 0 | 0.0018268 | 1 | 3.2401e-06 | 1.5994e-05 | 0.0041681
1 | m | 0.001 | 2 | 2 | Hz | 0 | m | 0 |  | 0.00041076 |  | 0.00042816 |  |  |  |  | m | 0.001 | 0 | 4.1076e-07 | 0 | 4.2816e-07 | 0 | 5.9333e-07 | #DIV/0! | 0 | 2Hz0m | 8 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2Hz1m0m | 0 | 0.00082151 | 0 | 0.00085633 | 1 | 1.1867e-06 | -1.9869e-05 | 0.0037669
1 | m | 0.001 | 5 | 5 | Hz | 0 | m | 0 |  | 0.00068727 |  | 0.00059678 |  |  |  |  | m | 0.001 | 0 | 6.8727e-07 | 0 | 5.9678e-07 | 0 | 9.1021e-07 | #DIV/0! | 0 | 5Hz0m | 9 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5Hz1m0m | 0 | 0.0013745 | 0 | 0.0011936 | 1 | 1.8204e-06 | -3.2981e-05 | 0.004329
1 | m | 0.001 | 10 | 10 | Hz | 0 | m | 0 |  | 0.0014874 |  | 8.7581e-05 |  |  |  |  | m | 0.001 | 0 | 1.4874e-06 | 0 | 8.7581e-08 | 0 | 1.49e-06 | #DIV/0! | 0 | 10Hz0m | 10 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10Hz1m0m | 0 | 0.0029748 | 0 | 0.00017516 | 1 | 2.98e-06 | 9.3299e-05 | 0.0026533
1 | m | 0.001 | 20 | 20 | Hz | 0 | m | 0 |  | 0.00044589 |  | 0.001845 |  |  |  |  | m | 0.001 | 0 | 4.4589e-07 | 0 | 1.845e-06 | 0 | 1.8981e-06 | #DIV/0! | 0 | 20Hz0m | 11 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20Hz1m0m | 0 | 0.00089178 | 0 | 0.00369 | 1 | 3.7962e-06 | -5.0503e-05 | 0.0049175
1 | m | 0.001 | 50 | 50 | Hz | 0 | m | 0 |  | 0.00077392 |  | 0.00013913 |  |  |  |  | m | 0.001 | 0 | 7.7392e-07 | 0 | 1.3913e-07 | 0 | 7.8633e-07 | #DIV/0! | 0 | 50Hz0m | 12 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50Hz1m0m | 0 | 0.0015478 | 0 | 0.00027825 | 1 | 1.5727e-06 | 2.6285e-05 | 0.0042513
1 | m | 0.001 | 100 | 100 | Hz | 0 | m | 0 |  | 0.0013008 |  | 2.1261e-05 |  |  |  |  | m | 0.001 | 0 | 1.3008e-06 | 0 | 2.1261e-08 | 0 | 1.301e-06 | #DIV/0! | 0 | 100Hz0m | 13 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100Hz1m0m | 0 | 0.0026016 | 0 | 4.2522e-05 | 1 | 2.602e-06 | -5.8762e-05 | 0.0028252
1 | m | 0.001 | 200 | 200 | Hz | 0 | m | 0 |  | 1.4781e-05 |  | 0.00074607 |  |  |  |  | m | 0.001 | 0 | 1.4781e-08 | 0 | 7.4607e-07 | 0 | 7.4621e-07 | #DIV/0! | 0 | 200Hz0m | 14 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200Hz1m0m | 0 | 2.9562e-05 | 0 | 0.0014921 | 1 | 1.4924e-06 | 7.5234e-05 | 0.0027488
1 | m | 0.001 | 500 | 500 | Hz | 0 | m | 0 |  | 0.00099724 |  | 0.0013278 |  |  |  |  | m | 0.001 | 0 | 9.9724e-07 | 0 | 1.3278e-06 | 0 | 1.6606e-06 | #DIV/0! | 0 | 500Hz0m | 15 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500Hz1m0m | 0 | 0.0019945 | 0 | 0.0026557 | 1 | 3.3212e-06 | -8.5022e-05 | 0.0041189
1 | m | 0.001 | 1000 | 1 | kHz | 0 | m | 0 |  | 0.0019861 |  | 0.00060196 |  |  |  |  | m | 0.001 | 0 | 1.9861e-06 | 0 | 6.0196e-07 | 0 | 2.0753e-06 | #DIV/0! | 0 | 1kHz0m | 16 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1kHz1m0m | 0 | 0.0039721 | 0 | 0.0012039 | 1 | 4.1506e-06 | 5.521e-05 | 0.00089881
1 | m | 0.001 | 2000 | 2 | kHz | 0 | m | 0 |  | 0.00095977 |  | 0.0018004 |  |  |  |  | m | 0.001 | 0 | 9.5977e-07 | 0 | 1.8004e-06 | 0 | 2.0402e-06 | #DIV/0! | 0 | 2kHz0m | 17 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2kHz1m0m | 0 | 0.0019195 | 0 | 0.0036008 | 1 | 4.0805e-06 | -2.6142e-05 | 0.0039561
1 | m | 0.001 | 5000 | 5 | kHz | 0 | m | 0 |  | 0.0011804 |  | 0.00069 |  |  |  |  | m | 0.001 | 0 | 1.1804e-06 | 0 | 6.9e-07 | 0 | 1.3673e-06 | #DIV/0! | 0 | 5kHz0m | 18 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5kHz1m0m | 0 | 0.0023608 | 0 | 0.00138 | 1 | 2.7346e-06 | -5.6076e-05 | 0.0039806
1 | m | 0.001 | 0.01 | 10 | mHz | 1 | m | 0.001 | 1.0005 | 0.0018537 | 0.00033498 | 0.0016987 | 0.00061958 | 0.0015953 | 0.00030512 | 0.0017027 | m | 0.001 | 0.00099991 | 2.4457e-06 | 2.986e-08 | 2.4051e-06 | 0.00099991 | 2.4457e-06 | 2.9862e-05 | 0.0024054 | 10mHz1m | 19 | 0.00099989 | 1e-08 | -2.6315e-08 | 5e-08 | 0.00099989 | 1e-08 | -2.6317e-05 | 5.0005e-05 | 10mHz1m1m | -1.7202e-08 | 0.0037075 | -5.6174e-08 | 0.0033974 | 0.99998 | 4.8914e-06 | -5.618e-05 | 0.004811
1 | m | 0.001 | 0.02 | 20 | mHz | 1 | m | 0.001 | 0.99923 | 0.0007116 | -0.0001398 | 0.001008 | -0.00068412 | 0.0018586 | -0.00012791 | 0.001076 | m | 0.001 | 0.00099991 | 1.9902e-06 | -1.1892e-08 | 1.4744e-06 | 0.00099991 | 1.9902e-06 | -1.1893e-05 | 0.0014745 | 20mHz1m | 20 | 0.0009999 | 1e-08 | -6.5946e-08 | 5e-08 | 0.0009999 | 1e-08 | -6.5953e-05 | 5.0005e-05 | 20mHz1m1m | -1.8437e-08 | 0.0014232 | -5.4054e-08 | 0.002016 | 0.99998 | 3.9803e-06 | -5.406e-05 | 0.0029495
1 | m | 0.001 | 0.05 | 50 | mHz | 1 | m | 0.001 | 1.0017 | 0.0008967 | -0.0017462 | 0.0018565 | 0.0018172 | 0.0014205 | -0.0018307 | 0.0013771 | m | 0.001 | 0.0009999 | 1.6799e-06 | 8.4534e-08 | 2.3115e-06 | 0.0009999 | 1.6799e-06 | 8.4543e-05 | 0.0023118 | 50mHz1m | 21 | 0.00099987 | 1e-08 | 5.5443e-08 | 5e-08 | 0.00099987 | 1e-08 | 5.545e-05 | 5.0007e-05 | 50mHz1m1m | -2.7165e-08 | 0.0017934 | -2.9092e-08 | 0.003713 | 0.99997 | 3.3597e-06 | -2.9093e-05 | 0.0046238
1 | m | 0.001 | 0.1 | 100 | mHz | 1 | m | 0.001 | 0.99892 | 0.0015311 | 0.00063653 | 0.00050631 | -0.0010928 | 0.0012942 | 0.00072003 | 0.00091802 | m | 0.001 | 0.001 | 2.0048e-06 | -8.3496e-08 | 1.0484e-06 | 0.001 | 2.0048e-06 | -8.3496e-05 | 0.0010484 | 100mHz1m | 22 | 0.00099994 | 1e-08 | -7.6078e-08 | 5e-08 | 0.00099994 | 1e-08 | -7.6082e-05 | 5.0003e-05 | 100mHz1m1m | -6.591e-08 | 0.0030622 | 7.4188e-09 | 0.0010126 | 0.99993 | 4.0096e-06 | 7.4137e-06 | 0.0020973
1 | m | 0.001 | 0.2 | 200 | mHz | 1 | m | 0.001 | 0.99819 | 0.0015507 | -6.1159e-06 | 0.0013497 | -0.0016649 | 0.00033771 | 0.00012332 | 0.0014201 | m | 0.001 | 0.00099985 | 1.587e-06 | -1.2944e-07 | 1.9592e-06 | 0.00099985 | 1.587e-06 | -0.00012946 | 0.0019595 | 200mHz1m | 23 | 0.00099993 | 1e-08 | -3.1994e-08 | 5e-08 | 0.00099993 | 1e-08 | -3.1996e-05 | 5.0003e-05 | 200mHz1m1m | 7.9909e-08 | 0.0031013 | 9.7444e-08 | 0.0026993 | 1.0001 | 3.174e-06 | 9.7461e-05 | 0.0039192
1 | m | 0.001 | 0.5 | 500 | mHz | 1 | m | 0.001 | 1.0011 | 0.0010111 | 0.00095498 | 0.00069434 | 0.0012058 | 0.0013147 | 0.00074904 | 0.00036712 | m | 0.001 | 0.00099986 | 1.6585e-06 | 2.0594e-07 | 7.8542e-07 | 0.00099986 | 1.6585e-06 | 0.00020597 | 0.00078553 | 500mHz1m | 24 | 0.00099991 | 1e-08 | 1.5338e-07 | 5e-08 | 0.00099991 | 1e-08 | 0.00015339 | 5.0005e-05 | 500mHz1m1m | 5.2745e-08 | 0.0020222 | -5.2563e-08 | 0.0013887 | 1.0001 | 3.317e-06 | -5.2578e-05 | 0.0015719
1 | m | 0.001 | 1 | 1 | Hz | 1 | m | 0.001 | 1.0011 | 0.0017413 | -0.001636 | 0.0007482 | 0.0011204 | 0.0015741 | -0.0015996 | 0.001945 | m | 0.001 | 0.001 | 2.3474e-06 | -3.634e-08 | 2.0839e-06 | 0.001 | 2.3474e-06 | -3.634e-05 | 0.0020839 | 1Hz1m | 25 | 0.001 | 1e-08 | -2.0346e-08 | 5e-08 | 0.001 | 1e-08 | -2.0345e-05 | 4.9999e-05 | 1Hz1m1m | 1.9306e-09 | 0.0034827 | 1.5994e-08 | 0.0014964 | 1 | 4.6948e-06 | 1.5994e-05 | 0.0041681
1 | m | 0.001 | 2 | 2 | Hz | 1 | m | 0.001 | 0.99889 | 0.0010591 | 0.0018306 | 0.0017141 | -0.0011503 | 0.0013664 | 0.0017956 | 0.00078034 | m | 0.001 | 0.001 | 1.7288e-06 | 3.4967e-08 | 1.8833e-06 | 0.001 | 1.7288e-06 | 3.4966e-05 | 0.0018833 | 2Hz1m | 26 | 0.0010001 | 1e-08 | 1.5098e-08 | 5e-08 | 0.0010001 | 1e-08 | 1.5097e-05 | 4.9996e-05 | 2Hz1m1m | 4.7626e-08 | 0.0021182 | -1.9869e-08 | 0.0034282 | 1 | 3.4575e-06 | -1.9869e-05 | 0.0037669
1 | m | 0.001 | 5 | 5 | Hz | 1 | m | 0.001 | 1.0019 | 0.00028994 | 0.00043714 | 0.00089252 | 0.0018258 | 0.00024004 | 0.0004241 | 0.0019719 | m | 0.001 | 0.0010001 | 3.7641e-07 | 1.304e-08 | 2.1645e-06 | 0.0010001 | 3.7641e-07 | 1.3039e-05 | 0.0021643 | 5Hz1m | 27 | 0.001 | 1e-08 | -1.9943e-08 | 5e-08 | 0.001 | 1e-08 | -1.9942e-05 | 4.9999e-05 | 5Hz1m1m | -2.1614e-08 | 0.00057989 | -3.2983e-08 | 0.001785 | 0.99998 | 7.5282e-07 | -3.2981e-05 | 0.004329
1 | m | 0.001 | 10 | 10 | Hz | 1 | m | 0.001 | 1.0006 | 0.0012812 | 0.00043936 | 0.0013241 | 0.00069041 | 0.00096534 | 0.00036898 | 7.6726e-05 | m | 0.001 | 0.00099988 | 1.6041e-06 | 7.0376e-08 | 1.3263e-06 | 0.00099988 | 1.6041e-06 | 7.0385e-05 | 0.0013264 | 10Hz1m | 28 | 0.00099983 | 1e-08 | 1.6366e-07 | 5e-08 | 0.00099983 | 1e-08 | 0.00016368 | 5.0009e-05 | 10Hz1m1m | -5.5154e-08 | 0.0025623 | 9.3279e-08 | 0.0026481 | 0.99994 | 3.2083e-06 | 9.3299e-05 | 0.0026533
1 | m | 0.001 | 20 | 20 | Hz | 1 | m | 0.001 | 0.99889 | 0.0018242 | -0.00024493 | 0.0017038 | -0.00095486 | 0.001141 | -0.00026013 | 0.0017719 | m | 0.001 | 0.00099984 | 2.1516e-06 | 1.5207e-08 | 2.4582e-06 | 0.00099984 | 2.1516e-06 | 1.5209e-05 | 0.0024586 | 20Hz1m | 29 | 0.00099993 | 1e-08 | -3.5291e-08 | 5e-08 | 0.00099993 | 1e-08 | -3.5293e-05 | 5.0004e-05 | 20Hz1m1m | 8.5112e-08 | 0.0036483 | -5.0498e-08 | 0.0034077 | 1.0001 | 4.3033e-06 | -5.0503e-05 | 0.0049175
1 | m | 0.001 | 50 | 50 | Hz | 1 | m | 0.001 | 1.0002 | 0.0013808 | 0.00072322 | 0.0013853 | 3.4064e-05 | 0.00057171 | 0.00051919 | 0.0016125 | m | 0.001 | 0.0010002 | 1.4945e-06 | 2.0403e-07 | 2.1258e-06 | 0.0010002 | 1.4945e-06 | 0.000204 | 0.0021255 | 50Hz1m | 30 | 0.0010001 | 1e-08 | 2.303e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00023028 | 4.9996e-05 | 50Hz1m1m | -7.8859e-08 | 0.0027617 | 2.6271e-08 | 0.0027705 | 0.99992 | 2.989e-06 | 2.6285e-05 | 0.0042513
1 | m | 0.001 | 100 | 100 | Hz | 1 | m | 0.001 | 1.0003 | 0.0015921 | -0.00013482 | 0.00082297 | 0.000252 | 0.0018055 | -0.00050122 | 0.0011479 | m | 0.001 | 0.001 | 2.4072e-06 | 3.664e-07 | 1.4124e-06 | 0.001 | 2.4072e-06 | 0.00036639 | 0.0014124 | 100Hz1m | 31 | 0.0010001 | 1e-08 | 3.0766e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00030763 | 4.9995e-05 | 100Hz1m1m | 6.7895e-08 | 0.0031842 | -5.8743e-08 | 0.0016459 | 1.0001 | 4.8143e-06 | -5.8762e-05 | 0.0028252
1 | m | 0.001 | 200 | 200 | Hz | 1 | m | 0.001 | 1.0016 | 0.00025582 | 0.0014257 | 0.00068265 | 0.0011191 | 0.001196 | 0.00097945 | 0.0011935 | m | 0.001 | 0.0010005 | 1.223e-06 | 4.4623e-07 | 1.3749e-06 | 0.0010005 | 1.223e-06 | 0.00044599 | 0.0013742 | 200Hz1m | 32 | 0.0010004 | 1e-08 | 5.2145e-07 | 5e-08 | 0.0010004 | 1e-08 | 0.00052123 | 4.9978e-05 | 200Hz1m1m | -9.3119e-08 | 0.00051164 | 7.5225e-08 | 0.0013653 | 0.99991 | 2.4461e-06 | 7.5234e-05 | 0.0027488
1 | m | 0.001 | 500 | 500 | Hz | 1 | m | 0.001 | 1.0022 | 0.0010804 | 0.00057879 | 0.0018694 | 0.0011281 | 0.0019169 | -0.00083035 | 0.00086885 | m | 0.001 | 0.001001 | 2.2004e-06 | 1.4091e-06 | 2.0614e-06 | 0.001001 | 2.2004e-06 | 0.0014077 | 0.0020593 | 500Hz1m | 33 | 0.0010011 | 1.5811e-08 | 1.3242e-06 | 5e-08 | 0.0010011 | 1.5812e-08 | 0.0013227 | 4.9943e-05 | 500Hz1m1m | 9.1095e-08 | 0.0021608 | -8.4991e-08 | 0.0037388 | 1.0001 | 4.4009e-06 | -8.5022e-05 | 0.0041189
1 | m | 0.001 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0045 | 0.00053223 | 0.0028919 | 2.5249e-05 | 0.0011901 | 0.0019181 | 0.00048951 | 0.00044738 | m | 0.001 | 0.0010033 | 1.9905e-06 | 2.4023e-06 | 4.4809e-07 | 0.0010033 | 1.9905e-06 | 0.0023944 | 0.00044664 | 1kHz1m | 34 | 0.0010033 | 4.4721e-08 | 2.4578e-06 | 1e-07 | 0.0010033 | 4.4722e-08 | 0.0024496 | 9.9666e-05 | 1kHz1m1m | 2.8722e-08 | 0.0010645 | 5.5464e-08 | 5.0498e-05 | 1 | 3.9811e-06 | 5.521e-05 | 0.00089881
1 | m | 0.001 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0091 | 0.00089291 | 0.0066929 | 0.001864 | -0.00035886 | 0.00073474 | 0.0015021 | 0.00070913 | m | 0.001 | 0.0010095 | 1.1563e-06 | 5.1909e-06 | 1.9944e-06 | 0.0010095 | 1.1564e-06 | 0.005142 | 0.0019755 | 2kHz1m | 35 | 0.0010095 | 1.2649e-07 | 5.1647e-06 | 2e-07 | 0.0010096 | 1.2649e-07 | 0.0051159 | 0.0001981 | 2kHz1m1m | 4.6281e-08 | 0.0017858 | -2.6154e-08 | 0.0037281 | 1 | 2.3127e-06 | -2.6142e-05 | 0.0039561
1 | m | 0.001 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.0393 | 0.00018984 | 0.014634 | 0.0019544 | 0.0015859 | 2.5811e-05 | 0.0016322 | 0.00062061 | m | 0.001 | 0.0010378 | 1.9159e-07 | 1.3002e-05 | 2.0506e-06 | 0.0010378 | 1.9329e-07 | 0.012528 | 0.0019757 | 5kHz1m | 36 | 0.0010378 | 5e-07 | 1.2944e-05 | 5e-07 | 0.0010379 | 5e-07 | 0.012472 | 0.00048176 | 5kHz1m1m | 2.546e-08 | 0.00037968 | -5.7884e-08 | 0.0039088 | 1 | 3.8657e-07 | -5.6076e-05 | 0.0039806
3 | m | 0.003 | 0.01 | 10 | mHz | 1 | m | 0.001 | 0.99849 | 0.0013416 | -0.00098678 | 0.00080681 | -0.0014374 | 0.00087069 | -0.00086047 | 0.0016192 | m | 0.001 | 0.00099992 | 1.5993e-06 | -1.2632e-07 | 1.8091e-06 | 0.00099992 | 1.5993e-06 | -0.00012633 | 0.0018092 | 10mHz1m | 19 | 0.00099989 | 1e-08 | -2.6315e-08 | 5e-08 | 0.00099989 | 1e-08 | -2.6317e-05 | 5.0005e-05 | 10mHz3m1m | -2.9179e-08 | 0.0026831 | 1e-07 | 0.0016136 | 0.99997 | 3.1987e-06 | 0.00010001 | 0.0036188
3 | m | 0.003 | 0.02 | 20 | mHz | 1 | m | 0.001 | 0.99984 | 0.00046466 | -0.00038784 | 0.00157 | 1.1309e-05 | 0.001445 | -0.00022191 | 5.6092e-05 | m | 0.001 | 0.00099983 | 1.5179e-06 | -1.6593e-07 | 1.571e-06 | 0.00099983 | 1.5179e-06 | -0.00016595 | 0.0015712 | 20mHz1m | 20 | 0.0009999 | 1e-08 | -6.5946e-08 | 5e-08 | 0.0009999 | 1e-08 | -6.5953e-05 | 5.0005e-05 | 20mHz3m1m | 6.9819e-08 | 0.00092933 | 9.9979e-08 | 0.0031399 | 1.0001 | 3.0358e-06 | 0.0001 | 0.0031429
3 | m | 0.003 | 0.05 | 50 | mHz | 1 | m | 0.001 | 1.0001 | 0.0015489 | -9.5897e-05 | 0.0011964 | 0.00021108 | 0.00017409 | -5.1344e-05 | 0.00077612 | m | 0.001 | 0.00099993 | 1.5587e-06 | -4.4553e-08 | 1.4261e-06 | 0.00099993 | 1.5587e-06 | -4.4556e-05 | 0.0014262 | 50mHz1m | 21 | 0.00099987 | 1e-08 | 5.5443e-08 | 5e-08 | 0.00099987 | 1e-08 | 5.545e-05 | 5.0007e-05 | 50mHz3m1m | -6.3511e-08 | 0.0030978 | 9.9996e-08 | 0.0023929 | 0.99994 | 3.1173e-06 | 0.00010001 | 0.0028529
3 | m | 0.003 | 0.1 | 100 | mHz | 1 | m | 0.001 | 1.0013 | 0.00039801 | -0.00014464 | 0.0018045 | 0.0014395 | 0.00070468 | 3.143e-05 | 0.00096639 | m | 0.001 | 0.00099989 | 8.0931e-07 | -1.7607e-07 | 2.047e-06 | 0.00099989 | 8.0931e-07 | -0.00017609 | 0.0020472 | 100mHz1m | 22 | 0.00099994 | 1e-08 | -7.6078e-08 | 5e-08 | 0.00099994 | 1e-08 | -7.6082e-05 | 5.0003e-05 | 100mHz3m1m | 5.162e-08 | 0.00079602 | 9.999e-08 | 0.003609 | 1.0001 | 1.6186e-06 | 0.00010001 | 0.0040947
3 | m | 0.003 | 0.2 | 200 | mHz | 1 | m | 0.001 | 1.0004 | 0.0010582 | 0.00040342 | 0.00074537 | 0.00043727 | 0.0003836 | 0.00053542 | 7.5681e-05 | m | 0.001 | 0.001 | 1.1256e-06 | -1.32e-07 | 7.4921e-07 | 0.001 | 1.1256e-06 | -0.000132 | 0.0007492 | 200mHz1m | 23 | 0.00099993 | 1e-08 | -3.1994e-08 | 5e-08 | 0.00099993 | 1e-08 | -3.1996e-05 | 5.0003e-05 | 200mHz3m1m | -8.0344e-08 | 0.0021164 | 1.0001e-07 | 0.0014907 | 0.99992 | 2.2511e-06 | 0.0001 | 0.0014992
3 | m | 0.003 | 0.5 | 500 | mHz | 1 | m | 0.001 | 1.0014 | 0.001689 | 0.00090579 | 0.0016237 | 0.0015134 | 0.001232 | 0.0008524 | 0.0019259 | m | 0.001 | 0.00099992 | 2.0905e-06 | 5.3387e-08 | 2.519e-06 | 0.00099992 | 2.0905e-06 | 5.3391e-05 | 0.0025193 | 500mHz1m | 24 | 0.00099991 | 1e-08 | 1.5338e-07 | 5e-08 | 0.00099991 | 1e-08 | 0.00015339 | 5.0005e-05 | 500mHz3m1m | -8.2739e-09 | 0.003378 | 9.9993e-08 | 0.0032475 | 0.99999 | 4.1811e-06 | 0.0001 | 0.0050388
3 | m | 0.003 | 1 | 1 | Hz | 1 | m | 0.001 | 0.99955 | 0.0017773 | 0.00095595 | 0.0011859 | -0.00051416 | 0.0017331 | 0.0010763 | 0.0018406 | m | 0.001 | 0.0010001 | 2.4825e-06 | -1.2035e-07 | 2.1896e-06 | 0.0010001 | 2.4825e-06 | -0.00012035 | 0.0021894 | 1Hz1m | 25 | 0.001 | 1e-08 | -2.0346e-08 | 5e-08 | 0.001 | 1e-08 | -2.0345e-05 | 4.9999e-05 | 1Hz3m1m | -4.974e-08 | 0.0035546 | 1.0001e-07 | 0.0023717 | 0.99995 | 4.9649e-06 | 0.0001 | 0.0043791
3 | m | 0.003 | 2 | 2 | Hz | 1 | m | 0.001 | 0.99817 | 9.8369e-06 | 0.00016843 | 0.0011576 | -0.0018774 | 0.00019652 | 0.00025334 | 0.0010689 | m | 0.001 | 0.001 | 1.9677e-07 | -8.4906e-08 | 1.5756e-06 | 0.001 | 1.9677e-07 | -8.4902e-05 | 0.0015755 | 2Hz1m | 26 | 0.0010001 | 1e-08 | 1.5098e-08 | 5e-08 | 0.0010001 | 1e-08 | 1.5097e-05 | 4.9996e-05 | 2Hz3m1m | 4.1868e-08 | 1.9674e-05 | 1e-07 | 0.0023153 | 1 | 3.9354e-07 | 9.9999e-05 | 0.0031515
3 | m | 0.003 | 5 | 5 | Hz | 1 | m | 0.001 | 0.99965 | 0.00039507 | -0.001885 | 0.0011382 | -0.00042345 | 0.0013387 | -0.001765 | 0.0010713 | m | 0.001 | 0.0010001 | 1.3958e-06 | -1.1995e-07 | 1.5631e-06 | 0.0010001 | 1.3958e-06 | -0.00011994 | 0.0015629 | 5Hz1m | 27 | 0.001 | 1e-08 | -1.9943e-08 | 5e-08 | 0.001 | 1e-08 | -1.9942e-05 | 4.9999e-05 | 5Hz3m1m | -4.9337e-08 | 0.00079014 | 1.0001e-07 | 0.0022765 | 0.99995 | 2.7916e-06 | 9.9998e-05 | 0.0031263
3 | m | 0.003 | 10 | 10 | Hz | 1 | m | 0.001 | 1.0003 | 0.000259 | 0.00024924 | 0.00051353 | 0.00053148 | 0.00060023 | 0.00018559 | 0.0016779 | m | 0.001 | 0.00099973 | 6.5373e-07 | 6.3652e-08 | 1.7548e-06 | 0.00099973 | 6.5373e-07 | 6.3669e-05 | 0.0017552 | 10Hz1m | 28 | 0.00099983 | 1e-08 | 1.6366e-07 | 5e-08 | 0.00099983 | 1e-08 | 0.00016368 | 5.0009e-05 | 10Hz3m1m | 9.6858e-08 | 0.000518 | 1e-07 | 0.0010271 | 1.0001 | 1.3075e-06 | 0.00010001 | 0.0035108
3 | m | 0.003 | 20 | 20 | Hz | 1 | m | 0.001 | 0.99866 | 0.0018806 | 0.00089618 | 0.0016096 | -0.0012576 | 0.0012297 | 0.0010315 | 0.0019541 | m | 0.001 | 0.00099992 | 2.247e-06 | -1.3529e-07 | 2.5317e-06 | 0.00099992 | 2.247e-06 | -0.0001353 | 0.0025319 | 20Hz1m | 29 | 0.00099993 | 1e-08 | -3.5291e-08 | 5e-08 | 0.00099993 | 1e-08 | -3.5293e-05 | 5.0004e-05 | 20Hz3m1m | 5.1085e-09 | 0.0037612 | 9.9997e-08 | 0.0032193 | 1 | 4.4939e-06 | 0.00010001 | 0.005064
3 | m | 0.003 | 50 | 50 | Hz | 1 | m | 0.001 | 0.99959 | 0.00066156 | 0.00013915 | 0.0018739 | -0.0005564 | 0.0014281 | 8.8426e-06 | 0.0011955 | m | 0.001 | 0.0010001 | 1.5739e-06 | 1.3031e-07 | 2.2227e-06 | 0.0010001 | 1.5739e-06 | 0.00013029 | 0.0022224 | 50Hz1m | 30 | 0.0010001 | 1e-08 | 2.303e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00023028 | 4.9996e-05 | 50Hz3m1m | -5.3917e-08 | 0.0013231 | 9.9995e-08 | 0.0037478 | 0.99995 | 3.1478e-06 | 9.9994e-05 | 0.0044451
3 | m | 0.003 | 100 | 100 | Hz | 1 | m | 0.001 | 1.0017 | 0.0018518 | 0.0011429 | 0.0014396 | 0.0016148 | 0.00029672 | 0.0009353 | 0.0018852 | m | 0.001 | 0.001 | 1.8754e-06 | 2.0764e-07 | 2.3721e-06 | 0.001 | 1.8754e-06 | 0.00020764 | 0.002372 | 100Hz1m | 31 | 0.0010001 | 1e-08 | 3.0766e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00030763 | 4.9995e-05 | 100Hz3m1m | 4.9863e-08 | 0.0037035 | 1.0002e-07 | 0.0028793 | 1 | 3.7508e-06 | 9.9996e-05 | 0.0047442
3 | m | 0.003 | 200 | 200 | Hz | 1 | m | 0.001 | 0.99957 | 0.0017528 | -0.00091209 | 0.00014821 | -0.00079383 | 0.0012648 | -0.0013335 | 0.00012698 | m | 0.001 | 0.0010004 | 2.1615e-06 | 4.2139e-07 | 1.9516e-07 | 0.0010004 | 2.1615e-06 | 0.00042124 | 0.00019509 | 200Hz1m | 32 | 0.0010004 | 1e-08 | 5.2145e-07 | 5e-08 | 0.0010004 | 1e-08 | 0.00052123 | 4.9978e-05 | 200Hz3m1m | 7.2959e-08 | 0.0035056 | 1.0007e-07 | 0.00029641 | 1.0001 | 4.323e-06 | 9.9993e-05 | 0.00039337
3 | m | 0.003 | 500 | 500 | Hz | 1 | m | 0.001 | 1.0004 | 0.001178 | -8.6052e-05 | 0.00099504 | -0.00082794 | 0.0018899 | -0.0013102 | 0.001062 | m | 0.001 | 0.0010012 | 2.227e-06 | 1.2242e-06 | 1.4553e-06 | 0.0010012 | 2.227e-06 | 0.0012227 | 0.0014535 | 500Hz1m | 33 | 0.0010011 | 1.5811e-08 | 1.3242e-06 | 5e-08 | 0.0010011 | 1.5812e-08 | 0.0013227 | 4.9943e-05 | 500Hz3m1m | -9.1984e-08 | 0.0023561 | 9.9972e-08 | 0.0019901 | 0.99991 | 4.4539e-06 | 9.9971e-05 | 0.0029075
3 | m | 0.003 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0047 | 0.00025116 | 0.0011535 | 0.00052106 | 0.0014244 | 0.00092184 | -0.0012041 | 0.00044231 | m | 0.001 | 0.0010033 | 9.5544e-07 | 2.3576e-06 | 6.8348e-07 | 0.0010033 | 9.5544e-07 | 0.0023498 | 0.00068122 | 1kHz1m | 34 | 0.0010033 | 4.4721e-08 | 2.4578e-06 | 1e-07 | 0.0010033 | 4.4722e-08 | 0.0024496 | 9.9666e-05 | 1kHz3m1m | 3.0166e-08 | 0.00050233 | 1.0023e-07 | 0.0010421 | 1 | 1.9109e-06 | 9.9824e-05 | 0.0013661
3 | m | 0.003 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0086 | 0.00039956 | 0.0042858 | 0.00023963 | -0.0010546 | 0.00021498 | -0.00077866 | 4.8269e-05 | m | 0.001 | 0.0010096 | 4.5372e-07 | 5.0645e-06 | 2.4444e-07 | 0.0010096 | 4.5372e-07 | 0.0050162 | 0.00024212 | 2kHz1m | 35 | 0.0010095 | 1.2649e-07 | 5.1647e-06 | 2e-07 | 0.0010096 | 1.2649e-07 | 0.0051159 | 0.0001981 | 2kHz3m1m | -7.4486e-08 | 0.00079912 | 1.0022e-07 | 0.00047925 | 0.99993 | 9.0744e-07 | 9.9639e-05 | 0.00052319
3 | m | 0.003 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.0371 | 0.0010529 | 0.014487 | 0.00075835 | -0.00064179 | 0.00083053 | 0.0016465 | 0.0017315 | m | 0.001 | 0.0010378 | 1.341e-06 | 1.2841e-05 | 1.8903e-06 | 0.0010378 | 1.3411e-06 | 0.012373 | 0.0018213 | 5kHz1m | 36 | 0.0010378 | 5e-07 | 1.2944e-05 | 5e-07 | 0.0010379 | 5e-07 | 0.012472 | 0.00048176 | 5kHz3m1m | 8.8731e-09 | 0.0021058 | 1.032e-07 | 0.0015167 | 1 | 2.6823e-06 | 9.9318e-05 | 0.0036744
3 | m | 0.003 | 0.01 | 10 | mHz | 3 | m | 0.003 | 3.0008 | 0.00027679 | -0.0011069 | 0.0016039 | 0.00086883 | 0.00032786 | -0.0016265 | 0.00077736 | m | 0.001 | 0.003 | 4.2907e-07 | 5.1954e-07 | 1.7823e-06 | 0.003 | 4.2907e-07 | 0.00017318 | 0.00059412 | 10mHz3m | 37 | 0.003 | 3e-08 | 3.8588e-07 | 1.5e-07 | 0.003 | 3e-08 | 0.00012863 | 5.0001e-05 | 10mHz3m3m | -2.546e-09 | 0.00055357 | -1.3366e-07 | 0.0032078 | 1 | 8.5814e-07 | -4.4554e-05 | 0.0011893
3 | m | 0.003 | 0.02 | 20 | mHz | 3 | m | 0.003 | 3.0019 | 0.0013624 | 0.00030675 | 0.00074542 | 0.0019343 | 0.00029594 | 0.00052745 | 0.0019468 | m | 0.001 | 0.003 | 1.3942e-06 | -2.207e-07 | 2.0846e-06 | 0.003 | 1.3942e-06 | -7.3568e-05 | 0.00069488 | 20mHz3m | 38 | 0.0029998 | 3e-08 | -3.6688e-09 | 1.5e-07 | 0.0029998 | 3e-08 | -1.223e-06 | 5.0003e-05 | 20mHz3m3m | -1.565e-07 | 0.0027249 | 2.1703e-07 | 0.0014908 | 0.99995 | 2.7884e-06 | 7.2345e-05 | 0.0013907
3 | m | 0.003 | 0.05 | 50 | mHz | 3 | m | 0.003 | 3.0012 | 0.00014596 | -0.00056291 | 0.0011832 | 0.00041993 | 0.0013259 | 1.9034e-05 | 0.0018726 | m | 0.001 | 0.0030008 | 1.3339e-06 | -5.8194e-07 | 2.215e-06 | 0.0030008 | 1.3339e-06 | -0.00019393 | 0.00073815 | 50mHz3m | 39 | 0.0030008 | 3e-08 | -5.7727e-07 | 1.5e-07 | 0.0030008 | 3e-08 | -0.00019237 | 4.9987e-05 | 50mHz3m3m | -6.3964e-08 | 0.00029192 | 4.6722e-09 | 0.0023663 | 0.99998 | 2.6679e-06 | 1.5529e-06 | 0.0014771
3 | m | 0.003 | 0.1 | 100 | mHz | 3 | m | 0.003 | 3 | 0.0017703 | 0.00021524 | 0.00023521 | 0.00025285 | 0.00072259 | -0.00014913 | 0.00099441 | m | 0.001 | 0.0029997 | 1.9121e-06 | 3.6438e-07 | 1.0218e-06 | 0.0029997 | 1.9121e-06 | 0.00012147 | 0.00034065 | 100mHz3m | 40 | 0.003 | 3e-08 | 1.1575e-07 | 1.5e-07 | 0.003 | 3e-08 | 3.8583e-05 | 5.0001e-05 | 100mHz3m3m | 2.4238e-07 | 0.0035406 | -2.4863e-07 | 0.00047042 | 1.0001 | 3.8242e-06 | -8.2887e-05 | 0.00068313
3 | m | 0.003 | 0.2 | 200 | mHz | 3 | m | 0.003 | 2.9985 | 8.8319e-05 | 1.0392e-05 | 0.0013093 | -0.0014177 | 0.0002221 | 0.00025665 | 0.00018541 | m | 0.001 | 0.0029999 | 2.3902e-07 | -2.4626e-07 | 1.3223e-06 | 0.0029999 | 2.3902e-07 | -8.209e-05 | 0.00044079 | 200mHz3m | 41 | 0.0029997 | 3e-08 | -1.6619e-07 | 1.5e-07 | 0.0029997 | 3e-08 | -5.5403e-05 | 5.0006e-05 | 200mHz3m3m | -2.4867e-07 | 0.00017664 | 8.0071e-08 | 0.0026185 | 0.99992 | 4.7804e-07 | 2.6686e-05 | 0.000883
3 | m | 0.003 | 0.5 | 500 | mHz | 3 | m | 0.003 | 3.0009 | 0.00042005 | -0.0011995 | 0.0013023 | 0.0010619 | 0.0015612 | -0.0010348 | 0.00071498 | m | 0.001 | 0.0029999 | 1.6168e-06 | -1.6467e-07 | 1.4856e-06 | 0.0029999 | 1.6168e-06 | -5.4893e-05 | 0.00049524 | 500mHz3m | 42 | 0.0029998 | 3e-08 | -5.5365e-08 | 1.5e-07 | 0.0029998 | 3e-08 | -1.8456e-05 | 5.0004e-05 | 500mHz3m3m | -9.8373e-08 | 0.00084011 | 1.0931e-07 | 0.0026046 | 0.99997 | 3.2335e-06 | 3.6436e-05 | 0.00099174
3 | m | 0.003 | 1 | 1 | Hz | 3 | m | 0.003 | 3.001 | 0.00067751 | 0.0012379 | 0.0013397 | 0.0012292 | 0.00089584 | 0.0017716 | 0.0011293 | m | 0.001 | 0.0029998 | 1.1232e-06 | -5.3373e-07 | 1.7522e-06 | 0.0029998 | 1.1232e-06 | -0.00017792 | 0.00058411 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -2.9816e-07 | 1.5e-07 | 0.0029996 | 3e-08 | -9.94e-05 | 5.0006e-05 | 1Hz3m3m | -1.534e-07 | 0.001355 | 2.3557e-07 | 0.0026795 | 0.99995 | 2.2464e-06 | 7.8523e-05 | 0.0011693
3 | m | 0.003 | 2 | 2 | Hz | 3 | m | 0.003 | 2.9985 | 0.00069817 | -0.0016101 | 0.00072271 | -0.0018563 | 0.00077338 | -0.0014603 | 0.0015919 | m | 0.001 | 0.0030003 | 1.0419e-06 | -1.4978e-07 | 1.7483e-06 | 0.0030003 | 1.0419e-06 | -4.9922e-05 | 0.00058269 | 2Hz3m | 44 | 0.0030005 | 3e-08 | 7.0631e-08 | 1.5e-07 | 0.0030005 | 3e-08 | 2.354e-05 | 4.9992e-05 | 2Hz3m3m | 1.6117e-07 | 0.0013963 | 2.2041e-07 | 0.0014454 | 1.0001 | 2.0838e-06 | 7.3462e-05 | 0.0011665
3 | m | 0.003 | 5 | 5 | Hz | 3 | m | 0.003 | 2.9996 | 0.0013468 | 0.00011939 | 0.00091264 | 0.00021164 | 0.00069392 | -0.00024633 | 0.00047988 | m | 0.001 | 0.0029994 | 1.515e-06 | 3.6573e-07 | 1.0311e-06 | 0.0029994 | 1.515e-06 | 0.00012193 | 0.00034377 | 5Hz3m | 45 | 0.0029997 | 3e-08 | 3.0067e-07 | 1.5e-07 | 0.0029997 | 3e-08 | 0.00010023 | 5.0005e-05 | 5Hz3m3m | 2.8449e-07 | 0.0026935 | -6.5054e-08 | 0.0018253 | 1.0001 | 3.03e-06 | -2.1698e-05 | 0.00068935
3 | m | 0.003 | 10 | 10 | Hz | 3 | m | 0.003 | 2.9993 | 0.0012906 | -0.0015204 | 0.00159 | -0.0010252 | 0.00032133 | -0.001649 | 0.0013614 | m | 0.001 | 0.0030003 | 1.33e-06 | 1.2866e-07 | 2.0932e-06 | 0.0030003 | 1.33e-06 | 4.2883e-05 | 0.00069765 | 10Hz3m | 46 | 0.0030002 | 3e-08 | -5.5728e-08 | 1.5e-07 | 0.0030002 | 3e-08 | -1.8575e-05 | 4.9997e-05 | 10Hz3m3m | -1.0045e-07 | 0.0025811 | -1.8439e-07 | 0.0031799 | 0.99997 | 2.6599e-06 | -6.1458e-05 | 0.0013962
3 | m | 0.003 | 20 | 20 | Hz | 3 | m | 0.003 | 3.0019 | 0.0010117 | 0.0025477 | 0.00053059 | 0.001355 | 0.0017215 | 0.0017866 | 0.00022588 | m | 0.001 | 0.0030006 | 1.9968e-06 | 7.6114e-07 | 5.7667e-07 | 0.0030006 | 1.9968e-06 | 0.00025367 | 0.00019219 | 20Hz3m | 47 | 0.0030007 | 3e-08 | 6.4789e-07 | 1.5e-07 | 0.0030007 | 3e-08 | 0.00021591 | 4.9988e-05 | 20Hz3m3m | 1.6919e-07 | 0.0020234 | -1.1325e-07 | 0.0010612 | 1.0001 | 3.9936e-06 | -3.7756e-05 | 0.00038761
3 | m | 0.003 | 50 | 50 | Hz | 3 | m | 0.003 | 3.0005 | 0.0010487 | 0.0027458 | 0.0016593 | 0.00062379 | 0.0007316 | 0.0019978 | 0.0015989 | m | 0.001 | 0.0029999 | 1.2787e-06 | 7.4795e-07 | 2.3043e-06 | 0.0029999 | 1.2787e-06 | 0.00024933 | 0.00076813 | 50Hz3m | 48 | 0.0030001 | 3e-08 | 6.1494e-07 | 1.5e-07 | 0.0030001 | 3e-08 | 0.00020497 | 4.9998e-05 | 50Hz3m3m | 2.2939e-07 | 0.0020975 | -1.3301e-07 | 0.0033185 | 1.0001 | 2.5574e-06 | -4.4355e-05 | 0.0015371
3 | m | 0.003 | 100 | 100 | Hz | 3 | m | 0.003 | 3.0023 | 0.0017223 | 0.0004664 | 0.00069026 | 0.0017978 | 0.001408 | -0.00038781 | 0.0012334 | m | 0.001 | 0.0030005 | 2.2246e-06 | 8.5421e-07 | 1.4134e-06 | 0.0030005 | 2.2246e-06 | 0.00028469 | 0.00047106 | 100Hz3m | 49 | 0.0030002 | 3e-08 | 9.0165e-07 | 1.5e-07 | 0.0030002 | 3e-08 | 0.00030053 | 4.9996e-05 | 100Hz3m3m | -2.785e-07 | 0.0034446 | 4.7449e-08 | 0.0013805 | 0.99991 | 4.4492e-06 | 1.5841e-05 | 0.00094344
3 | m | 0.003 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0027 | 0.0020144 | 0.0028709 | 0.0011278 | 0.0018164 | 0.00025304 | 0.0013433 | 0.00083067 | m | 0.001 | 0.0030009 | 2.0302e-06 | 1.5276e-06 | 1.4007e-06 | 0.0030009 | 2.0302e-06 | 0.00050905 | 0.00046675 | 200Hz3m | 50 | 0.0030006 | 3e-08 | 1.7291e-06 | 1.5e-07 | 0.0030006 | 3e-08 | 0.00057624 | 4.9989e-05 | 200Hz3m3m | -2.3484e-07 | 0.0040287 | 2.0149e-07 | 0.0022555 | 0.99992 | 4.0604e-06 | 6.7189e-05 | 0.00093484
3 | m | 0.003 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0053 | 0.00045347 | 0.0045102 | 0.0018488 | 0.0015817 | 0.0019521 | 0.00092335 | 0.0019898 | m | 0.001 | 0.0030038 | 2.0041e-06 | 3.5869e-06 | 2.7161e-06 | 0.0030038 | 2.0041e-06 | 0.0011941 | 0.00090423 | 500Hz3m | 51 | 0.0030038 | 4.7434e-08 | 3.7346e-06 | 1.5e-07 | 0.0030038 | 4.7434e-08 | 0.0012433 | 4.9937e-05 | 500Hz3m3m | -5.275e-09 | 0.00090694 | 1.4777e-07 | 0.0036975 | 1 | 4.0081e-06 | 4.9196e-05 | 0.0018092
3 | m | 0.003 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0115 | 0.0019403 | 0.0069779 | 0.0016585 | 0.0010154 | 0.0018925 | -0.00096114 | 0.0010447 | m | 0.001 | 0.0030104 | 2.7104e-06 | 7.939e-06 | 1.9601e-06 | 0.0030105 | 2.7104e-06 | 0.0026371 | 0.0006511 | 1kHz3m | 52 | 0.0030106 | 1.3416e-07 | 7.9216e-06 | 3e-07 | 0.0030106 | 1.3417e-07 | 0.0026312 | 9.9648e-05 | 1kHz3m3m | 1.2457e-07 | 0.0038805 | -1.7428e-08 | 0.003317 | 1 | 5.4207e-06 | -5.8979e-06 | 0.001306
3 | m | 0.003 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0282 | 0.0014038 | 0.015699 | 0.00042484 | -0.00034782 | 0.001676 | 7.7562e-06 | 0.001665 | m | 0.001 | 0.0030285 | 2.1863e-06 | 1.5691e-05 | 1.7183e-06 | 0.0030286 | 2.1862e-06 | 0.0051811 | 0.00056738 | 2kHz3m | 53 | 0.0030284 | 3.7947e-07 | 1.5762e-05 | 6e-07 | 0.0030285 | 3.7948e-07 | 0.0052045 | 0.00019812 | 2kHz3m3m | -1.0953e-07 | 0.0028076 | 7.0263e-08 | 0.00084968 | 0.99996 | 4.3725e-06 | 2.3388e-05 | 0.0011519
3 | m | 0.003 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.1127 | 0.00068239 | 0.036938 | 0.00026475 | -0.00067684 | 0.00095165 | -0.0017369 | 0.00053349 | m | 0.001 | 0.0031134 | 1.171e-06 | 3.8675e-05 | 5.9557e-07 | 0.0031137 | 1.171e-06 | 0.012421 | 0.00019132 | 5kHz3m | 54 | 0.0031137 | 1.5e-06 | 3.8872e-05 | 1.5e-06 | 0.003114 | 1.5e-06 | 0.012483 | 0.0004817 | 5kHz3m3m | 3.1244e-07 | 0.0013648 | 1.9679e-07 | 0.00052951 | 1.0001 | 2.3419e-06 | 6.1945e-05 | 0.00061518
10 | m | 0.01 | 0.01 | 10 | mHz | 3 | m | 0.003 | 2.9982 | 0.00045465 | -0.00070653 | 0.00028862 | -0.001716 | 0.00053517 | -0.00079242 | 0.0018105 | m | 0.001 | 0.0029999 | 7.0222e-07 | 8.5885e-08 | 1.8334e-06 | 0.0029999 | 7.0222e-07 | 2.8629e-05 | 0.00061114 | 10mHz3m | 37 | 0.003 | 3e-08 | 3.8588e-07 | 1.5e-07 | 0.003 | 3e-08 | 0.00012863 | 5.0001e-05 | 10mHz10m3m | 1.876e-08 | 0.00090929 | 3e-07 | 0.00057723 | 1 | 1.4044e-06 | 0.0001 | 0.0012233
10 | m | 0.01 | 0.02 | 20 | mHz | 3 | m | 0.003 | 2.9992 | 0.00025246 | 0.00017659 | 0.00034952 | -0.0004971 | 0.00099455 | 0.00048024 | 0.0010956 | m | 0.001 | 0.0029997 | 1.0261e-06 | -3.0365e-07 | 1.15e-06 | 0.0029997 | 1.0261e-06 | -0.00010123 | 0.00038336 | 20mHz3m | 38 | 0.0029998 | 3e-08 | -3.6688e-09 | 1.5e-07 | 0.0029998 | 3e-08 | -1.223e-06 | 5.0003e-05 | 20mHz10m3m | 1.1033e-07 | 0.00050493 | 2.9998e-07 | 0.00069903 | 1 | 2.0522e-06 | 0.0001 | 0.00076835
10 | m | 0.01 | 0.05 | 50 | mHz | 3 | m | 0.003 | 2.9988 | 0.00031176 | 0.00029069 | 0.0014555 | -0.0019727 | 0.0014881 | 0.001168 | 0.00021092 | m | 0.001 | 0.0030008 | 1.5204e-06 | -8.7733e-07 | 1.4707e-06 | 0.0030008 | 1.5204e-06 | -0.00029236 | 0.00049012 | 50mHz3m | 39 | 0.0030008 | 3e-08 | -5.7727e-07 | 1.5e-07 | 0.0030008 | 3e-08 | -0.00019237 | 4.9987e-05 | 50mHz10m3m | -4.4673e-08 | 0.00062352 | 3.0006e-07 | 0.0029111 | 0.99999 | 3.0408e-06 | 9.999e-05 | 0.0009815
10 | m | 0.01 | 0.1 | 100 | mHz | 3 | m | 0.003 | 2.9998 | 0.00094966 | -1.4001e-05 | 0.00082081 | -0.00022869 | 0.0011063 | 0.00017025 | 0.00055269 | m | 0.001 | 0.003 | 1.458e-06 | -1.8425e-07 | 9.8954e-07 | 0.003 | 1.458e-06 | -6.1418e-05 | 0.00032985 | 100mHz3m | 40 | 0.003 | 3e-08 | 1.1575e-07 | 1.5e-07 | 0.003 | 3e-08 | 3.8583e-05 | 5.0001e-05 | 100mHz10m3m | -3.5068e-08 | 0.0018993 | 3e-07 | 0.0016416 | 0.99999 | 2.916e-06 | 0.0001 | 0.00066159
10 | m | 0.01 | 0.2 | 200 | mHz | 3 | m | 0.003 | 3.0008 | 0.00070824 | -0.001407 | 0.0015238 | 0.00098933 | 0.0013297 | -0.00094077 | 0.0018099 | m | 0.001 | 0.0029998 | 1.5066e-06 | -4.662e-07 | 2.366e-06 | 0.0029998 | 1.5066e-06 | -0.00015541 | 0.00078872 | 200mHz3m | 41 | 0.0029997 | 3e-08 | -1.6619e-07 | 1.5e-07 | 0.0029997 | 3e-08 | -5.5403e-05 | 5.0006e-05 | 200mHz10m3m | -1.1265e-07 | 0.0014165 | 3.0001e-07 | 0.0030477 | 0.99996 | 3.0131e-06 | 0.00010001 | 0.0015782
10 | m | 0.01 | 0.5 | 500 | mHz | 3 | m | 0.003 | 3.0012 | 0.0013744 | 0.00079482 | 0.00069885 | 0.0011818 | 7.4441e-05 | 0.0011502 | 0.00093259 | m | 0.001 | 0.0030001 | 1.3764e-06 | -3.5538e-07 | 1.1654e-06 | 0.0030001 | 1.3764e-06 | -0.00011846 | 0.00038845 | 500mHz3m | 42 | 0.0029998 | 3e-08 | -5.5365e-08 | 1.5e-07 | 0.0029998 | 3e-08 | -1.8456e-05 | 5.0004e-05 | 500mHz10m3m | -2.9748e-07 | 0.0027488 | 3.0002e-07 | 0.0013977 | 0.9999 | 2.7529e-06 | 0.0001 | 0.00077851
10 | m | 0.01 | 1 | 1 | Hz | 3 | m | 0.003 | 2.9984 | 0.00062361 | 0.0012093 | 0.00050916 | -0.0011377 | 0.001614 | 0.0018074 | 0.001085 | m | 0.001 | 0.0029995 | 1.7303e-06 | -5.9812e-07 | 1.1986e-06 | 0.0029995 | 1.7303e-06 | -0.0001994 | 0.00039958 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -2.9816e-07 | 1.5e-07 | 0.0029996 | 3e-08 | -9.94e-05 | 5.0006e-05 | 1Hz10m3m | 1.1918e-07 | 0.0012472 | 2.9995e-07 | 0.0010183 | 1 | 3.4605e-06 | 0.0001 | 0.00080072
10 | m | 0.01 | 2 | 2 | Hz | 3 | m | 0.003 | 2.9996 | 0.0012496 | -0.00041994 | 0.0017962 | -0.00072187 | 0.00079806 | -0.00019054 | 0.0011144 | m | 0.001 | 0.0030004 | 1.4827e-06 | -2.294e-07 | 2.1139e-06 | 0.0030004 | 1.4827e-06 | -7.6457e-05 | 0.00070453 | 2Hz3m | 44 | 0.0030005 | 3e-08 | 7.0631e-08 | 1.5e-07 | 0.0030005 | 3e-08 | 2.354e-05 | 4.9992e-05 | 2Hz10m3m | 1.0342e-07 | 0.0024993 | 3.0003e-07 | 0.0035924 | 1 | 2.9655e-06 | 9.9997e-05 | 0.00141
10 | m | 0.01 | 5 | 5 | Hz | 3 | m | 0.003 | 2.9985 | 0.00029282 | 0.0017117 | 0.0003241 | -0.0011952 | 0.00063273 | 0.001711 | 0.00028168 | m | 0.001 | 0.0029997 | 6.972e-07 | 6.9002e-10 | 4.294e-07 | 0.0029997 | 6.972e-07 | 2.3003e-07 | 0.00014315 | 5Hz3m | 45 | 0.0029997 | 3e-08 | 3.0067e-07 | 1.5e-07 | 0.0029997 | 3e-08 | 0.00010023 | 5.0005e-05 | 5Hz10m3m | -2.6382e-08 | 0.00058565 | 2.9998e-07 | 0.00064819 | 0.99999 | 1.3944e-06 | 0.0001 | 0.00029063
10 | m | 0.01 | 10 | 10 | Hz | 3 | m | 0.003 | 3.0009 | 0.00096184 | -0.00049701 | 0.0016469 | 0.0005417 | 0.0012625 | -0.00014126 | 0.00073632 | m | 0.001 | 0.0030003 | 1.5871e-06 | -3.5575e-07 | 1.804e-06 | 0.0030003 | 1.5871e-06 | -0.00011857 | 0.00060128 | 10Hz3m | 46 | 0.0030002 | 3e-08 | -5.5728e-08 | 1.5e-07 | 0.0030002 | 3e-08 | -1.8575e-05 | 4.9997e-05 | 10Hz10m3m | -1.3483e-07 | 0.0019237 | 3.0003e-07 | 0.0032939 | 0.99996 | 3.1742e-06 | 9.9997e-05 | 0.0012036
10 | m | 0.01 | 20 | 20 | Hz | 3 | m | 0.003 | 2.9994 | 0.0019336 | 0.00031957 | 0.00071121 | -0.0013317 | 0.0019113 | -2.8267e-05 | 0.00065567 | m | 0.001 | 0.0030007 | 2.7188e-06 | 3.4784e-07 | 9.6733e-07 | 0.0030007 | 2.7188e-06 | 0.00011592 | 0.00032236 | 20Hz3m | 47 | 0.0030007 | 3e-08 | 6.4789e-07 | 1.5e-07 | 0.0030007 | 3e-08 | 0.00021591 | 4.9988e-05 | 20Hz10m3m | -2.0254e-08 | 0.0038673 | 3.0005e-07 | 0.0014224 | 0.99999 | 5.4376e-06 | 9.9993e-05 | 0.00064666
10 | m | 0.01 | 50 | 50 | Hz | 3 | m | 0.003 | 2.9983 | 0.00095941 | 0.0017025 | 5.8567e-05 | -0.0015408 | 0.0013827 | 0.0013876 | 0.00087114 | m | 0.001 | 0.0029999 | 1.683e-06 | 3.1492e-07 | 8.731e-07 | 0.0029999 | 1.683e-06 | 0.00010498 | 0.00029105 | 50Hz3m | 48 | 0.0030001 | 3e-08 | 6.1494e-07 | 1.5e-07 | 0.0030001 | 3e-08 | 0.00020497 | 4.9998e-05 | 50Hz10m3m | 2.1137e-07 | 0.0019188 | 3.0003e-07 | 0.00011713 | 1.0001 | 3.366e-06 | 9.9998e-05 | 0.00058423
10 | m | 0.01 | 100 | 100 | Hz | 3 | m | 0.003 | 3.0006 | 0.0002945 | -0.0012575 | 0.001501 | 0.00054453 | 0.0010886 | -0.0018591 | 0.0011379 | m | 0.001 | 0.003 | 1.1278e-06 | 6.0161e-07 | 1.8836e-06 | 0.003 | 1.1278e-06 | 0.00020053 | 0.00062784 | 100Hz3m | 49 | 0.0030002 | 3e-08 | 9.0165e-07 | 1.5e-07 | 0.0030002 | 3e-08 | 0.00030053 | 4.9996e-05 | 100Hz10m3m | 1.7855e-07 | 0.000589 | 3.0004e-07 | 0.003002 | 1.0001 | 2.2555e-06 | 9.9994e-05 | 0.0012567
10 | m | 0.01 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0014 | 0.0017764 | -0.00015839 | 0.0019933 | 0.00057079 | 0.00068726 | -0.0015875 | 0.0019816 | m | 0.001 | 0.0030009 | 1.9047e-06 | 1.4291e-06 | 2.8107e-06 | 0.0030009 | 1.9047e-06 | 0.00047624 | 0.00093663 | 200Hz3m | 50 | 0.0030006 | 3e-08 | 1.7291e-06 | 1.5e-07 | 0.0030006 | 3e-08 | 0.00057624 | 4.9989e-05 | 200Hz10m3m | -2.0914e-07 | 0.0035527 | 2.9996e-07 | 0.0039867 | 0.99993 | 3.8094e-06 | 9.9999e-05 | 0.0018739
10 | m | 0.01 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0035 | 0.00051008 | 0.0041837 | 0.0010514 | -0.00021645 | 0.00090299 | 0.00074925 | 0.0014333 | m | 0.001 | 0.0030038 | 1.0371e-06 | 3.4345e-06 | 1.7776e-06 | 0.0030038 | 1.0371e-06 | 0.0011434 | 0.00059178 | 500Hz3m | 51 | 0.0030038 | 4.7434e-08 | 3.7346e-06 | 1.5e-07 | 0.0030038 | 4.7434e-08 | 0.0012433 | 4.9937e-05 | 500Hz10m3m | -2.1582e-09 | 0.0010202 | 3.0014e-07 | 0.0021028 | 1 | 2.0742e-06 | 9.9922e-05 | 0.0011846
10 | m | 0.01 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0113 | 0.0014968 | 0.008588 | 0.00082679 | 0.00091746 | 0.00027058 | 0.00096732 | 0.0014184 | m | 0.001 | 0.0030104 | 1.5211e-06 | 7.6207e-06 | 1.6418e-06 | 0.0030104 | 1.5211e-06 | 0.0025314 | 0.00054536 | 1kHz3m | 52 | 0.0030106 | 1.3416e-07 | 7.9216e-06 | 3e-07 | 0.0030106 | 1.3417e-07 | 0.0026312 | 9.9648e-05 | 1kHz10m3m | 1.7979e-07 | 0.0029936 | 3.0091e-07 | 0.0016536 | 1.0001 | 3.0422e-06 | 9.9798e-05 | 0.0010953
10 | m | 0.01 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0285 | 0.0015798 | 0.014815 | 0.00029984 | -7.418e-05 | 0.00022403 | -0.00064446 | 0.00043521 | m | 0.001 | 0.0030286 | 1.5956e-06 | 1.546e-05 | 5.2849e-07 | 0.0030286 | 1.5956e-06 | 0.0051046 | 0.00017452 | 2kHz3m | 53 | 0.0030284 | 3.7947e-07 | 1.5762e-05 | 6e-07 | 0.0030285 | 3.7948e-07 | 0.0052045 | 0.00019812 | 2kHz10m3m | -1.2113e-07 | 0.0031596 | 3.0201e-07 | 0.00059968 | 0.99996 | 3.1912e-06 | 9.9926e-05 | 0.00040135
10 | m | 0.01 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.1127 | 0.0012166 | 0.03782 | 0.00058536 | -0.0010714 | 0.0013896 | -0.00074211 | 0.0012556 | m | 0.001 | 0.0031138 | 1.8469e-06 | 3.8562e-05 | 1.3854e-06 | 0.003114 | 1.8468e-06 | 0.012384 | 0.00044491 | 5kHz3m | 54 | 0.0031137 | 1.5e-06 | 3.8872e-05 | 1.5e-06 | 0.003114 | 1.5e-06 | 0.012483 | 0.0004817 | 5kHz10m3m | -2.7568e-08 | 0.0024332 | 3.0949e-07 | 0.0011707 | 0.99999 | 3.6937e-06 | 9.9491e-05 | 0.0010118
10 | m | 0.01 | 0.01 | 10 | mHz | 10 | m | 0.01 | 9.9992 | 0.001119 | 0.0011947 | 0.00099155 | 0.00069329 | 0.00066258 | 0.00058215 | 0.00011518 | m | 0.001 | 0.0099985 | 1.3005e-06 | 6.1257e-07 | 9.9822e-07 | 0.0099985 | 1.3005e-06 | 6.1267e-05 | 9.9837e-05 | 10mHz10m | 55 | 0.0099982 | 1e-07 | 1.3593e-06 | 5e-07 | 0.0099982 | 1e-07 | 0.00013595 | 5.0009e-05 | 10mHz10m10m | -3.409e-07 | 0.002238 | 7.4669e-07 | 0.0019831 | 0.99997 | 2.6009e-06 | 7.4685e-05 | 0.00020584
10 | m | 0.01 | 0.02 | 20 | mHz | 10 | m | 0.01 | 9.9977 | 0.0010867 | -0.0016669 | 0.0011604 | -0.0012717 | 0.0012219 | -0.00076994 | 0.0004332 | m | 0.001 | 0.009999 | 1.6353e-06 | -8.9699e-07 | 1.2386e-06 | 0.009999 | 1.6353e-06 | -8.9708e-05 | 0.00012387 | 20mHz10m | 56 | 0.0099996 | 1e-07 | -1.2518e-06 | 5e-07 | 0.0099996 | 1e-07 | -0.00012518 | 5.0002e-05 | 20mHz10m10m | 6.5208e-07 | 0.0021735 | -3.5481e-07 | 0.0023208 | 1.0001 | 3.2705e-06 | -3.5476e-05 | 0.00025274
10 | m | 0.01 | 0.05 | 50 | mHz | 10 | m | 0.01 | 9.9986 | 0.0018576 | 3.6839e-05 | 0.00046915 | -0.0013336 | 0.00075463 | -0.00020444 | 0.0012277 | m | 0.001 | 0.0099999 | 2.0051e-06 | 2.4128e-07 | 1.3143e-06 | 0.0099999 | 2.0051e-06 | 2.4128e-05 | 0.00013143 | 50mHz10m | 57 | 0.0099998 | 1e-07 | 3.3603e-07 | 5e-07 | 0.0099998 | 1e-07 | 3.3604e-05 | 5.0001e-05 | 50mHz10m10m | -9.4711e-08 | 0.0037153 | 9.476e-08 | 0.00093829 | 0.99999 | 4.0101e-06 | 9.4764e-06 | 0.00026758
10 | m | 0.01 | 0.1 | 100 | mHz | 10 | m | 0.01 | 10.0019 | 0.0014068 | 0.0013924 | 0.0018226 | 0.00091859 | 0.00041205 | 0.0015735 | 0.00016116 | m | 0.001 | 0.010001 | 1.4659e-06 | -1.8109e-07 | 1.8297e-06 | 0.010001 | 1.4659e-06 | -1.8107e-05 | 0.00018295 | 100mHz10m | 58 | 0.01 | 1e-07 | 2.0039e-07 | 5e-07 | 0.01 | 1e-07 | 2.0038e-05 | 4.9998e-05 | 100mHz10m10m | -6.6623e-07 | 0.0028136 | 3.8148e-07 | 0.0036452 | 0.99993 | 2.9318e-06 | 3.8145e-05 | 0.0003693
10 | m | 0.01 | 0.2 | 200 | mHz | 10 | m | 0.01 | 10.001 | 0.00090385 | 0.00020671 | 0.0011818 | -0.0013683 | 0.0016666 | 0.00028803 | 0.00028043 | m | 0.001 | 0.010002 | 1.8959e-06 | -8.1318e-08 | 1.2146e-06 | 0.010002 | 1.8959e-06 | -8.1298e-06 | 0.00012143 | 200mHz10m | 59 | 0.010002 | 1e-07 | -7.1478e-07 | 5e-07 | 0.010002 | 1e-07 | -7.1463e-05 | 4.999e-05 | 200mHz10m10m | -2.8064e-07 | 0.0018077 | -6.3346e-07 | 0.0023635 | 0.99997 | 3.7919e-06 | -6.3333e-05 | 0.00024795
10 | m | 0.01 | 0.5 | 500 | mHz | 10 | m | 0.01 | 10.0013 | 0.00096118 | 0.00016015 | 0.0016359 | 0.00033725 | 0.0018569 | -0.0003021 | 0.0014514 | m | 0.001 | 0.010001 | 2.0909e-06 | 4.6225e-07 | 2.187e-06 | 0.010001 | 2.0909e-06 | 4.622e-05 | 0.00021868 | 500mHz10m | 60 | 0.01 | 1e-07 | 1.0089e-06 | 5e-07 | 0.01 | 1e-07 | 0.00010088 | 4.9999e-05 | 500mHz10m10m | -7.983e-07 | 0.0019224 | 5.4661e-07 | 0.0032719 | 0.99992 | 4.1819e-06 | 5.4664e-05 | 0.0004402
10 | m | 0.01 | 1 | 1 | Hz | 10 | m | 0.01 | 10.0023 | 0.0017828 | 0.00017142 | 0.00049746 | 0.0017633 | 0.00016327 | 0.0005053 | 0.0016281 | m | 0.001 | 0.010001 | 1.7903e-06 | -3.3388e-07 | 1.7024e-06 | 0.010001 | 1.7903e-06 | -3.3387e-05 | 0.00017023 | 1Hz10m | 61 | 0.01 | 1e-07 | -8.1803e-08 | 5e-07 | 0.01 | 1e-07 | -8.18e-06 | 4.9998e-05 | 1Hz10m10m | -1.9857e-07 | 0.0035657 | 2.5208e-07 | 0.00099493 | 0.99998 | 3.5806e-06 | 2.5206e-05 | 0.00034411
10 | m | 0.01 | 2 | 2 | Hz | 10 | m | 0.01 | 10 | 0.00066354 | 0.00097003 | 0.001017 | -0.00019825 | 0.00011699 | -5.3096e-05 | 0.00088997 | m | 0.001 | 0.01 | 6.7377e-07 | 1.0231e-06 | 1.3515e-06 | 0.01 | 6.7377e-07 | 0.00010231 | 0.00013514 | 2Hz10m | 62 | 0.01 | 1e-07 | 1.618e-06 | 5e-07 | 0.01 | 1e-07 | 0.00016179 | 4.9998e-05 | 2Hz10m10m | 2.4445e-07 | 0.0013271 | 5.9488e-07 | 0.0020341 | 1 | 1.3475e-06 | 5.9483e-05 | 0.00027487
10 | m | 0.01 | 5 | 5 | Hz | 10 | m | 0.01 | 9.9984 | 0.001961 | -0.0012961 | 0.0014434 | -0.00072685 | 0.00015132 | -0.00153 | 0.0013912 | m | 0.001 | 0.0099992 | 1.9668e-06 | 2.3383e-07 | 2.0047e-06 | 0.0099992 | 1.9668e-06 | 2.3385e-05 | 0.00020049 | 5Hz10m | 63 | 0.0099991 | 1e-07 | 1.3304e-07 | 5e-07 | 0.0099991 | 1e-07 | 1.3305e-05 | 5.0005e-05 | 5Hz10m10m | -9.645e-08 | 0.003922 | -1.008e-07 | 0.0028868 | 0.99999 | 3.9336e-06 | -1.008e-05 | 0.00040408
10 | m | 0.01 | 10 | 10 | Hz | 10 | m | 0.01 | 10 | 0.00075253 | 0.0010718 | 9.4092e-05 | -0.00012495 | 0.0015374 | -0.0010195 | 0.00035267 | m | 0.001 | 0.01 | 1.7117e-06 | 2.0913e-06 | 3.6501e-07 | 0.01 | 1.7117e-06 | 0.00020913 | 3.65e-05 | 10Hz10m | 64 | 0.0099999 | 1e-07 | 2.0509e-06 | 5e-07 | 0.0099999 | 1e-07 | 0.0002051 | 5e-05 | 10Hz10m10m | -1.6362e-07 | 0.0015051 | -4.0397e-08 | 0.00018818 | 0.99998 | 3.4234e-06 | -4.0363e-06 | 8.8483e-05
10 | m | 0.01 | 20 | 20 | Hz | 10 | m | 0.01 | 10.0011 | 0.00142 | -0.001624 | 0.00010592 | 0.00034565 | 0.00029759 | -0.0017181 | 0.00091385 | m | 0.001 | 0.010001 | 1.4508e-06 | 9.4184e-08 | 9.1997e-07 | 0.010001 | 1.4508e-06 | 9.4177e-06 | 9.199e-05 | 20Hz10m | 65 | 0.010002 | 1e-07 | 4.0912e-07 | 5e-07 | 0.010002 | 1e-07 | 4.0906e-05 | 4.9992e-05 | 20Hz10m10m | 8.368e-07 | 0.00284 | 3.1494e-07 | 0.00021184 | 1.0001 | 2.9017e-06 | 3.1488e-05 | 0.00019065
10 | m | 0.01 | 50 | 50 | Hz | 10 | m | 0.01 | 10.0009 | 0.0018199 | 0.0027847 | 0.00073664 | 0.0016221 | 0.00041024 | -0.0013201 | 0.0012851 | m | 0.001 | 0.0099993 | 1.8656e-06 | 4.1047e-06 | 1.4812e-06 | 0.0099993 | 1.8656e-06 | 0.0004105 | 0.00014813 | 50Hz10m | 66 | 0.0099996 | 1e-07 | 3.5276e-06 | 5e-07 | 0.0099996 | 1e-07 | 0.00035278 | 5.0002e-05 | 50Hz10m10m | 2.6326e-07 | 0.0036399 | -5.7711e-07 | 0.0014733 | 1 | 3.7312e-06 | -5.7724e-05 | 0.00030045
10 | m | 0.01 | 100 | 100 | Hz | 10 | m | 0.01 | 10.0022 | 0.0015015 | 0.0042867 | 0.00030578 | 0.00075318 | 0.00057836 | 0.00062671 | 0.0013317 | m | 0.001 | 0.010001 | 1.6091e-06 | 3.66e-06 | 1.3664e-06 | 0.010001 | 1.6091e-06 | 0.00036594 | 0.00013662 | 100Hz10m | 67 | 0.010001 | 1e-07 | 2.9473e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00029469 | 4.9993e-05 | 100Hz10m10m | -1.4533e-07 | 0.003003 | -7.1273e-07 | 0.00061156 | 0.99999 | 3.2181e-06 | -7.1258e-05 | 0.00027777
10 | m | 0.01 | 200 | 200 | Hz | 10 | m | 0.01 | 10.0034 | 0.00076063 | 0.0049094 | 0.0014807 | 0.00014979 | 0.00078092 | 0.0017759 | 0.0010162 | m | 0.001 | 0.010003 | 1.0901e-06 | 3.1335e-06 | 1.7959e-06 | 0.010003 | 1.0901e-06 | 0.00031325 | 0.00017953 | 200Hz10m | 68 | 0.010003 | 1e-07 | 4.0264e-06 | 5e-07 | 0.010003 | 1e-07 | 0.00040251 | 4.9984e-05 | 200Hz10m10m | -9.7041e-08 | 0.0015213 | 8.9286e-07 | 0.0029613 | 0.99999 | 2.1803e-06 | 8.9261e-05 | 0.00036252
10 | m | 0.01 | 500 | 500 | Hz | 10 | m | 0.01 | 10.0153 | 0.0014734 | 0.013254 | 0.00068201 | 0.0011434 | 0.0015996 | -0.00029725 | 0.001734 | m | 0.001 | 0.010014 | 2.1747e-06 | 1.3551e-05 | 1.8633e-06 | 0.010014 | 2.1747e-06 | 0.0013532 | 0.00018606 | 500Hz10m | 69 | 0.010014 | 1.5811e-07 | 1.3086e-05 | 5e-07 | 0.010014 | 1.5812e-07 | 0.0013068 | 4.993e-05 | 500Hz10m10m | -1.3647e-07 | 0.0029467 | -4.6502e-07 | 0.001364 | 0.99999 | 4.3495e-06 | -4.6419e-05 | 0.00037546
10 | m | 0.01 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.0355 | 0.0012302 | 0.027354 | 0.00034783 | 0.0010726 | 5.6206e-05 | 0.0012733 | 0.00085395 | m | 0.001 | 0.010034 | 1.2314e-06 | 2.608e-05 | 9.2207e-07 | 0.010034 | 1.2314e-06 | 0.0025991 | 9.1891e-05 | 1kHz10m | 70 | 0.010034 | 4.4721e-07 | 2.5733e-05 | 1e-06 | 0.010034 | 4.4722e-07 | 0.0025646 | 9.966e-05 | 1kHz10m10m | -4.2464e-07 | 0.0024603 | -3.4684e-07 | 0.00069567 | 0.99996 | 2.4629e-06 | -3.4456e-05 | 0.00020906
10 | m | 0.01 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.0988 | 0.0010647 | 0.054194 | 0.00024734 | 0.0016431 | 0.00088313 | 0.0014289 | 0.00094119 | m | 0.001 | 0.010097 | 1.3833e-06 | 5.2765e-05 | 9.7314e-07 | 0.010097 | 1.3833e-06 | 0.0052257 | 9.6378e-05 | 2kHz10m | 71 | 0.010096 | 1.2649e-06 | 5.3584e-05 | 2e-06 | 0.010096 | 1.2649e-06 | 0.0053073 | 0.00019809 | 2kHz10m10m | -9.1201e-07 | 0.0021293 | 8.1914e-07 | 0.00049468 | 0.99991 | 2.7666e-06 | 8.1602e-05 | 0.00027639
10 | m | 0.01 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.3792 | 0.0015435 | 0.12821 | 0.00090767 | 0.00051159 | 0.0016631 | -0.0019535 | 0.0019485 | m | 0.001 | 0.010379 | 2.269e-06 | 0.00013017 | 2.1495e-06 | 0.01038 | 2.269e-06 | 0.012541 | 0.0002071 | 5kHz10m | 72 | 0.010379 | 5e-06 | 0.00013086 | 5e-06 | 0.01038 | 5e-06 | 0.012608 | 0.00048171 | 5kHz10m10m | 1.7562e-07 | 0.0030869 | 6.9801e-07 | 0.0018153 | 1 | 4.5382e-06 | 6.703e-05 | 0.00063529
100 | m | 0.1 | 0.01 | 10 | mHz | 10 | m | 0.01 | 9.9995 | 7.7427e-05 | 0.0022905 | 0.001501 | 0.00099356 | 0.00071648 | 0.0019311 | 0.0014809 | m | 0.001 | 0.0099985 | 7.2065e-07 | 3.5942e-07 | 2.1086e-06 | 0.0099985 | 7.2065e-07 | 3.5948e-05 | 0.00021089 | 10mHz10m | 55 | 0.0099982 | 1e-07 | 1.3593e-06 | 5e-07 | 0.0099982 | 1e-07 | 0.00013595 | 5.0009e-05 | 10mHz100m10m | -3.4704e-07 | 0.00015485 | 9.9984e-07 | 0.003002 | 0.99997 | 1.4413e-06 | 0.0001 | 0.00042473
100 | m | 0.1 | 0.02 | 20 | mHz | 10 | m | 0.01 | 9.9996 | 0.0011894 | -0.0024466 | 0.00016886 | 0.00084733 | 0.00077569 | -0.00019505 | 0.0014755 | m | 0.001 | 0.0099988 | 1.4199e-06 | -2.2516e-06 | 1.4851e-06 | 0.0099988 | 1.4199e-06 | -0.00022518 | 0.00014853 | 20mHz10m | 56 | 0.0099996 | 1e-07 | -1.2518e-06 | 5e-07 | 0.0099996 | 1e-07 | -0.00012518 | 5.0002e-05 | 20mHz100m10m | 8.7605e-07 | 0.0023787 | 9.9977e-07 | 0.00033772 | 1.0001 | 2.8399e-06 | 0.0001 | 0.00030124
100 | m | 0.1 | 0.05 | 50 | mHz | 10 | m | 0.01 | 10.0015 | 0.00094798 | 5.6192e-05 | 0.0004613 | 0.00099129 | 0.00021144 | 0.00072019 | 0.0011243 | m | 0.001 | 0.010001 | 9.7127e-07 | -6.6399e-07 | 1.2152e-06 | 0.010001 | 9.7127e-07 | -6.6396e-05 | 0.00012152 | 50mHz10m | 57 | 0.0099998 | 1e-07 | 3.3603e-07 | 5e-07 | 0.0099998 | 1e-07 | 3.3604e-05 | 5.0001e-05 | 50mHz100m10m | -6.8791e-07 | 0.001896 | 1e-06 | 0.00092261 | 0.99993 | 1.9425e-06 | 0.0001 | 0.00024812
100 | m | 0.1 | 0.1 | 100 | mHz | 10 | m | 0.01 | 10.0011 | 0.0017836 | 2.9941e-05 | 0.00049793 | 0.00039792 | 0.00074036 | 0.00082961 | 0.00028194 | m | 0.001 | 0.010001 | 1.9311e-06 | -7.9967e-07 | 5.7221e-07 | 0.010001 | 1.9311e-06 | -7.9961e-05 | 5.7217e-05 | 100mHz10m | 58 | 0.01 | 1e-07 | 2.0039e-07 | 5e-07 | 0.01 | 1e-07 | 2.0038e-05 | 4.9998e-05 | 100mHz100m10m | -3.2234e-07 | 0.0035671 | 1.0001e-06 | 0.00099585 | 0.99997 | 3.8623e-06 | 9.9999e-05 | 0.00012488
100 | m | 0.1 | 0.2 | 200 | mHz | 10 | m | 0.01 | 9.9999 | 0.0013496 | -0.0011147 | 0.00083942 | -0.0013607 | 0.00045327 | 0.00060014 | 0.0017207 | m | 0.001 | 0.010001 | 1.4237e-06 | -1.7148e-06 | 1.9146e-06 | 0.010001 | 1.4237e-06 | -0.00017146 | 0.00019143 | 200mHz10m | 59 | 0.010002 | 1e-07 | -7.1478e-07 | 5e-07 | 0.010002 | 1e-07 | -7.1463e-05 | 4.999e-05 | 200mHz100m10m | 8.5582e-07 | 0.0026992 | 1e-06 | 0.0016788 | 1.0001 | 2.8474e-06 | 9.9996e-05 | 0.00038612
100 | m | 0.1 | 0.5 | 500 | mHz | 10 | m | 0.01 | 9.9996 | 0.00098542 | -0.0010248 | 0.0012319 | 0.00013784 | 9.7409e-05 | -0.0010337 | 0.00074171 | m | 0.001 | 0.0099995 | 9.9023e-07 | 8.8443e-09 | 1.438e-06 | 0.0099995 | 9.9023e-07 | 8.8447e-07 | 0.00014381 | 500mHz10m | 60 | 0.01 | 1e-07 | 1.0089e-06 | 5e-07 | 0.01 | 1e-07 | 0.00010088 | 4.9999e-05 | 500mHz100m10m | 6.6981e-07 | 0.0019708 | 1e-06 | 0.0024639 | 1.0001 | 1.9805e-06 | 0.0001 | 0.00029193
100 | m | 0.1 | 1 | 1 | Hz | 10 | m | 0.01 | 10.0001 | 0.00029873 | -0.0021098 | 0.0011943 | -0.0011099 | 0.00016933 | -0.0010278 | 0.00090674 | m | 0.001 | 0.010001 | 3.4339e-07 | -1.0819e-06 | 1.4995e-06 | 0.010001 | 3.4339e-07 | -0.00010818 | 0.00014993 | 1Hz10m | 61 | 0.01 | 1e-07 | -8.1803e-08 | 5e-07 | 0.01 | 1e-07 | -8.18e-06 | 4.9998e-05 | 1Hz100m10m | -8.6174e-07 | 0.00059747 | 1.0001e-06 | 0.0023885 | 0.99991 | 6.8677e-07 | 0.0001 | 0.000304
100 | m | 0.1 | 2 | 2 | Hz | 10 | m | 0.01 | 10.0014 | 0.0010947 | 0.001922 | 0.00023348 | 0.00030671 | 0.0019487 | 0.001304 | 0.0016312 | m | 0.001 | 0.010001 | 2.2351e-06 | 6.1801e-07 | 1.6478e-06 | 0.010001 | 2.2351e-06 | 6.1794e-05 | 0.00016476 | 2Hz10m | 62 | 0.01 | 1e-07 | 1.618e-06 | 5e-07 | 0.01 | 1e-07 | 0.00016179 | 4.9998e-05 | 2Hz100m10m | -7.3847e-07 | 0.0021895 | 9.9999e-07 | 0.00046696 | 0.99993 | 4.4702e-06 | 0.0001 | 0.0003333
100 | m | 0.1 | 5 | 5 | Hz | 10 | m | 0.01 | 10.0003 | 0.00031839 | 0.00030234 | 0.001624 | 0.00068279 | 0.0013924 | 0.0011693 | 0.0005565 | m | 0.001 | 0.0099996 | 1.4283e-06 | -8.6692e-07 | 1.7167e-06 | 0.0099996 | 1.4283e-06 | -8.6695e-05 | 0.00017167 | 5Hz10m | 63 | 0.0099991 | 1e-07 | 1.3304e-07 | 5e-07 | 0.0099991 | 1e-07 | 1.3305e-05 | 5.0005e-05 | 5Hz100m10m | -5.3364e-07 | 0.00063677 | 9.9995e-07 | 0.0032479 | 0.99995 | 2.8566e-06 | 0.0001 | 0.00034697
100 | m | 0.1 | 10 | 10 | Hz | 10 | m | 0.01 | 10.0004 | 0.0013296 | 0.00081918 | 0.0014541 | 0.0007647 | 0.000482 | -0.00023174 | 9.0251e-05 | m | 0.001 | 0.0099997 | 1.4143e-06 | 1.0509e-06 | 1.4569e-06 | 0.0099997 | 1.4143e-06 | 0.0001051 | 0.0001457 | 10Hz10m | 64 | 0.0099999 | 1e-07 | 2.0509e-06 | 5e-07 | 0.0099999 | 1e-07 | 0.0002051 | 5e-05 | 10Hz100m10m | 2.7438e-07 | 0.0026592 | 1e-06 | 0.0029082 | 1 | 2.8285e-06 | 0.0001 | 0.00029565
100 | m | 0.1 | 20 | 20 | Hz | 10 | m | 0.01 | 10.002 | 0.0020191 | -0.00010635 | 0.00024866 | 1.4904e-05 | 0.0016622 | 0.00048469 | 0.00066692 | m | 0.001 | 0.010002 | 2.6153e-06 | -5.9104e-07 | 7.1177e-07 | 0.010002 | 2.6153e-06 | -5.9092e-05 | 7.1163e-05 | 20Hz10m | 65 | 0.010002 | 1e-07 | 4.0912e-07 | 5e-07 | 0.010002 | 1e-07 | 4.0906e-05 | 4.9992e-05 | 20Hz100m10m | -3.6477e-07 | 0.0040383 | 1.0002e-06 | 0.00049732 | 0.99996 | 5.2306e-06 | 9.9998e-05 | 0.00015085
100 | m | 0.1 | 50 | 50 | Hz | 10 | m | 0.01 | 9.9995 | 0.0019979 | 0.0026937 | 0.00049232 | 0.00079996 | 0.00038769 | 0.00016624 | 0.00015106 | m | 0.001 | 0.0099987 | 2.0352e-06 | 2.5274e-06 | 5.1498e-07 | 0.0099987 | 2.0352e-06 | 0.00025278 | 5.1505e-05 | 50Hz10m | 66 | 0.0099996 | 1e-07 | 3.5276e-06 | 5e-07 | 0.0099996 | 1e-07 | 0.00035278 | 5.0002e-05 | 50Hz100m10m | 9.0546e-07 | 0.0039958 | 1.0002e-06 | 0.00098465 | 1.0001 | 4.0703e-06 | 0.0001 | 0.0001145
100 | m | 0.1 | 100 | 100 | Hz | 10 | m | 0.01 | 10.0029 | 0.0011593 | 0.0020466 | 0.0010383 | 0.0010153 | 0.00092237 | 9.9315e-05 | 0.0014547 | m | 0.001 | 0.010002 | 1.4814e-06 | 1.9472e-06 | 1.7873e-06 | 0.010002 | 1.4814e-06 | 0.00019469 | 0.00017869 | 100Hz10m | 67 | 0.010001 | 1e-07 | 2.9473e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00029469 | 4.9993e-05 | 100Hz100m10m | -5.4453e-07 | 0.0023185 | 1e-06 | 0.0020767 | 0.99995 | 2.9629e-06 | 9.9999e-05 | 0.00036087
100 | m | 0.1 | 200 | 200 | Hz | 10 | m | 0.01 | 10.004 | 0.00086234 | 0.0019313 | 0.0012612 | 0.0017357 | 0.0003184 | -0.0010945 | 7.7965e-05 | m | 0.001 | 0.010002 | 9.1924e-07 | 3.0258e-06 | 1.2636e-06 | 0.010002 | 9.1924e-07 | 0.00030252 | 0.00012633 | 200Hz10m | 68 | 0.010003 | 1e-07 | 4.0264e-06 | 5e-07 | 0.010003 | 1e-07 | 0.00040251 | 4.9984e-05 | 200Hz100m10m | 9.2281e-07 | 0.0017247 | 1.0005e-06 | 0.0025224 | 1.0001 | 1.8385e-06 | 9.9995e-05 | 0.00025756
100 | m | 0.1 | 500 | 500 | Hz | 10 | m | 0.01 | 10.0132 | 0.00065738 | 0.011123 | 3.1304e-05 | -0.00025131 | 0.00074993 | -0.00096089 | 0.00096886 | m | 0.001 | 0.010013 | 9.9726e-07 | 1.2084e-05 | 9.6937e-07 | 0.010013 | 9.9726e-07 | 0.0012068 | 9.6806e-05 | 500Hz10m | 69 | 0.010014 | 1.5811e-07 | 1.3086e-05 | 5e-07 | 0.010014 | 1.5812e-07 | 0.0013068 | 4.993e-05 | 500Hz100m10m | 5.4485e-07 | 0.0013148 | 1.002e-06 | 6.2609e-05 | 1.0001 | 1.9945e-06 | 9.9996e-05 | 0.00019995
100 | m | 0.1 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.0343 | 0.0019234 | 0.025505 | 0.00020497 | 0.00065352 | 0.00039516 | 0.00077563 | 0.0018276 | m | 0.001 | 0.010034 | 1.9636e-06 | 2.473e-05 | 1.8391e-06 | 0.010034 | 1.9636e-06 | 0.0024647 | 0.00018329 | 1kHz10m | 70 | 0.010034 | 4.4721e-07 | 2.5733e-05 | 1e-06 | 0.010034 | 4.4722e-07 | 0.0025646 | 9.966e-05 | 1kHz100m10m | 3.7851e-07 | 0.0038469 | 1.0036e-06 | 0.00040995 | 1 | 3.9272e-06 | 9.9925e-05 | 0.00037989
100 | m | 0.1 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.0954 | 0.0013397 | 0.053682 | 0.0011115 | -0.000202 | 0.0001451 | 0.0011094 | 0.0016553 | m | 0.001 | 0.010096 | 1.3475e-06 | 5.2573e-05 | 1.9939e-06 | 0.010096 | 1.3475e-06 | 0.0052074 | 0.0001975 | 2kHz10m | 71 | 0.010096 | 1.2649e-06 | 5.3584e-05 | 2e-06 | 0.010096 | 1.2649e-06 | 0.0053073 | 0.00019809 | 2kHz100m10m | 6.6578e-07 | 0.0026793 | 1.0116e-06 | 0.002223 | 1.0001 | 2.6951e-06 | 9.9849e-05 | 0.00044188
100 | m | 0.1 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.3776 | 0.00026271 | 0.13068 | 0.0016816 | -0.0013731 | 0.00032755 | 0.00085371 | 0.0019161 | m | 0.001 | 0.010379 | 4.1989e-07 | 0.00012983 | 2.5494e-06 | 0.01038 | 4.2107e-07 | 0.012508 | 0.00024559 | 5kHz10m | 72 | 0.010379 | 5e-06 | 0.00013086 | 5e-06 | 0.01038 | 5e-06 | 0.012608 | 0.00048171 | 5kHz100m10m | -8.1853e-08 | 0.00052545 | 1.0333e-06 | 0.0033632 | 0.99999 | 8.4373e-07 | 9.9643e-05 | 0.00068797
100 | m | 0.1 | 0.01 | 10 | mHz | 100 | m | 0.1 | 100.0169 | 0.0025489 | -0.004767 | 0.00099021 | -0.0010167 | 0.0019721 | 0.00087987 | 0.001669 | m | 0.001 | 0.10002 | 3.2227e-06 | -5.6469e-06 | 1.9406e-06 | 0.10002 | 3.2227e-06 | -5.6459e-05 | 1.9403e-05 | 10mHz100m | 73 | 0.10002 | 1e-06 | -1.1852e-05 | 5e-06 | 0.10002 | 1e-06 | -0.0001185 | 4.999e-05 | 10mHz100m100m | 2.2624e-06 | 0.0050978 | -6.2053e-06 | 0.0019804 | 1 | 6.4463e-06 | -6.204e-05 | 6.3284e-05
100 | m | 0.1 | 0.02 | 20 | mHz | 100 | m | 0.1 | 99.9952 | 0.00087679 | 0.0011948 | 0.0014492 | 0.00055937 | 0.00015488 | -0.00017249 | 0.0010774 | m | 0.001 | 0.099995 | 8.9037e-07 | 1.3673e-06 | 1.8058e-06 | 0.099995 | 8.9037e-07 | 1.3674e-05 | 1.8059e-05 | 20mHz100m | 74 | 0.099994 | 1e-06 | -2.9683e-06 | 5e-06 | 0.099994 | 1e-06 | -2.9685e-05 | 5.0003e-05 | 20mHz100m100m | -4.0604e-07 | 0.0017536 | -4.3356e-06 | 0.0028984 | 1 | 1.7835e-06 | -4.3359e-05 | 6.1683e-05
100 | m | 0.1 | 0.05 | 50 | mHz | 100 | m | 0.1 | 100.0049 | 0.001806 | -0.0069255 | 0.00099102 | -0.00089026 | 0.00070947 | -0.00096539 | 0.0012683 | m | 0.001 | 0.10001 | 1.9404e-06 | -5.9601e-06 | 1.6095e-06 | 0.10001 | 1.9404e-06 | -5.9597e-05 | 1.6094e-05 | 50mHz100m | 75 | 0.099999 | 1e-06 | 3.6376e-06 | 5e-06 | 0.099999 | 1e-06 | 3.6376e-05 | 5e-05 | 50mHz100m100m | -6.5693e-06 | 0.0036121 | 9.5977e-06 | 0.001982 | 0.99993 | 3.8821e-06 | 9.5973e-05 | 5.9466e-05
100 | m | 0.1 | 0.1 | 100 | mHz | 100 | m | 0.1 | 99.9949 | 0.0023907 | 0.00819 | 0.00056446 | 0.0018642 | 0.00035478 | 0.0016498 | 0.00033744 | m | 0.001 | 0.099993 | 2.4169e-06 | 6.5402e-06 | 6.5764e-07 | 0.099993 | 2.4169e-06 | 6.5407e-05 | 6.5768e-06 | 100mHz100m | 76 | 0.1 | 1e-06 | 9.7792e-07 | 5e-06 | 0.1 | 1e-06 | 9.779e-06 | 4.9999e-05 | 100mHz100m100m | 9.3522e-06 | 0.0047814 | -5.5623e-06 | 0.0011289 | 1.0001 | 4.8348e-06 | -5.5628e-05 | 5.17e-05
100 | m | 0.1 | 0.2 | 200 | mHz | 100 | m | 0.1 | 100.0139 | 0.0023693 | 0.0079098 | 0.0018566 | -0.0007022 | 0.0010815 | 0.0014372 | 6.8898e-05 | m | 0.001 | 0.10001 | 2.6044e-06 | 6.4727e-06 | 1.8579e-06 | 0.10001 | 2.6044e-06 | 6.4717e-05 | 1.8576e-05 | 200mHz100m | 77 | 0.10001 | 1e-06 | 3.3939e-07 | 5e-06 | 0.10001 | 1e-06 | 3.3936e-06 | 4.9996e-05 | 200mHz100m100m | -7.227e-06 | 0.0047385 | -6.1333e-06 | 0.0037133 | 0.99993 | 5.2098e-06 | -6.1324e-05 | 6.2289e-05
100 | m | 0.1 | 0.5 | 500 | mHz | 100 | m | 0.1 | 99.9976 | 0.0019416 | 0.002097 | 0.00019069 | -0.0017861 | 0.001916 | -0.0019902 | 0.0019472 | m | 0.001 | 0.099999 | 2.7278e-06 | 4.0872e-06 | 1.9565e-06 | 0.099999 | 2.7278e-06 | 4.0872e-05 | 1.9566e-05 | 500mHz100m | 78 | 0.099993 | 1e-06 | -5.7838e-06 | 5e-06 | 0.099993 | 1e-06 | -5.7842e-05 | 5.0003e-05 | 500mHz100m100m | -6.0135e-06 | 0.0038832 | -9.871e-06 | 0.00038142 | 0.99994 | 5.4565e-06 | -9.8714e-05 | 6.3495e-05
100 | m | 0.1 | 1 | 1 | Hz | 100 | m | 0.1 | 100.0268 | 0.0017941 | 0.020364 | 0.0019777 | 0.0018637 | 0.00093848 | -0.00053432 | 0.0016491 | m | 0.001 | 0.10002 | 2.0248e-06 | 2.0898e-05 | 2.575e-06 | 0.10002 | 2.0248e-06 | 0.00020893 | 2.5744e-05 | 1Hz100m | 79 | 0.10002 | 1e-06 | 2.386e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00023856 | 4.9991e-05 | 1Hz100m100m | -7.8289e-06 | 0.0035883 | 2.9615e-06 | 0.0039554 | 0.99992 | 4.0508e-06 | 2.9626e-05 | 7.1765e-05
100 | m | 0.1 | 2 | 2 | Hz | 100 | m | 0.1 | 99.9988 | 0.0022577 | 0.00038103 | 0.00021003 | -0.0014822 | 0.00080023 | -0.00036526 | 0.0010317 | m | 0.001 | 0.1 | 2.3953e-06 | 7.4628e-07 | 1.0529e-06 | 0.1 | 2.3953e-06 | 7.4628e-06 | 1.0529e-05 | 2Hz100m | 80 | 0.099996 | 1e-06 | 7.6805e-06 | 5e-06 | 0.099996 | 1e-06 | 7.6808e-05 | 5.0002e-05 | 2Hz100m100m | -4.1287e-06 | 0.0045153 | 6.9342e-06 | 0.00042009 | 0.99996 | 4.7916e-06 | 6.9345e-05 | 5.4255e-05
100 | m | 0.1 | 5 | 5 | Hz | 100 | m | 0.1 | 99.991 | 0.0026787 | -0.038146 | 0.0015761 | 0.0019665 | 0.0014606 | 0.0003128 | 0.0002961 | m | 0.001 | 0.099989 | 3.0511e-06 | -3.8459e-05 | 1.6037e-06 | 0.099989 | 3.0511e-06 | -0.00038463 | 1.6039e-05 | 5Hz100m | 81 | 0.099991 | 1e-06 | -2.8674e-05 | 5e-06 | 0.099991 | 1e-06 | -0.00028677 | 5.0005e-05 | 5Hz100m100m | 1.5396e-06 | 0.0053574 | 9.785e-06 | 0.0031523 | 1 | 6.1029e-06 | 9.7865e-05 | 5.9409e-05
100 | m | 0.1 | 10 | 10 | Hz | 100 | m | 0.1 | 100.0034 | 0.0019458 | 0.002755 | 0.00043208 | -0.0014083 | 0.00051392 | -0.00066068 | 0.0002005 | m | 0.001 | 0.1 | 2.0126e-06 | 3.4157e-06 | 4.7634e-07 | 0.1 | 2.0126e-06 | 3.4155e-05 | 4.7631e-06 | 10Hz100m | 82 | 0.10001 | 1e-06 | 4.2077e-06 | 5e-06 | 0.10001 | 1e-06 | 4.2075e-05 | 4.9997e-05 | 10Hz100m100m | 5.9265e-07 | 0.0038917 | 7.9204e-07 | 0.00086417 | 1 | 4.0264e-06 | 7.9198e-06 | 5.0897e-05
100 | m | 0.1 | 20 | 20 | Hz | 100 | m | 0.1 | 100.0068 | 0.0025412 | 0.013473 | 0.00063884 | 0.0019232 | 0.00022936 | -0.0017679 | 0.0008806 | m | 0.001 | 0.1 | 2.5515e-06 | 1.5241e-05 | 1.0879e-06 | 0.1 | 2.5515e-06 | 0.0001524 | 1.0879e-05 | 20Hz100m | 83 | 0.10001 | 1e-06 | 1.3861e-05 | 5e-06 | 0.10001 | 1e-06 | 0.0001386 | 4.9996e-05 | 20Hz100m100m | 3.2106e-06 | 0.0050824 | -1.3796e-06 | 0.0012777 | 1 | 5.104e-06 | -1.3799e-05 | 5.4525e-05
100 | m | 0.1 | 50 | 50 | Hz | 100 | m | 0.1 | 100.0094 | 0.0023521 | 0.0054763 | 0.0018458 | -0.00036259 | 0.00021626 | 0.00083974 | 0.0015906 | m | 0.001 | 0.10001 | 2.362e-06 | 4.6365e-06 | 2.4365e-06 | 0.10001 | 2.362e-06 | 4.6361e-05 | 2.4363e-05 | 50Hz100m | 84 | 0.10001 | 1e-06 | 8.2003e-06 | 5e-06 | 0.10001 | 1e-06 | 8.1996e-05 | 4.9996e-05 | 50Hz100m100m | -2.1731e-06 | 0.0047042 | 3.5637e-06 | 0.0036915 | 0.99998 | 4.7251e-06 | 3.5636e-05 | 6.9813e-05
100 | m | 0.1 | 100 | 100 | Hz | 100 | m | 0.1 | 100.0222 | 0.001318 | 0.027191 | 0.00076395 | -0.00025796 | 0.0013546 | -0.00081522 | 0.0011618 | m | 0.001 | 0.10002 | 1.89e-06 | 2.8006e-05 | 1.3905e-06 | 0.10002 | 1.89e-06 | 0.00028 | 1.3901e-05 | 100Hz100m | 85 | 0.10002 | 1e-06 | 2.9508e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00029503 | 4.9992e-05 | 100Hz100m100m | -5.7359e-06 | 0.0026361 | 1.5014e-06 | 0.0015279 | 0.99994 | 3.7813e-06 | 1.5027e-05 | 5.7203e-05
100 | m | 0.1 | 200 | 200 | Hz | 100 | m | 0.1 | 100.0264 | 0.001261 | 0.058084 | 0.0002816 | 0.00194 | 0.00063108 | -0.0013094 | 0.00041364 | m | 0.001 | 0.10002 | 1.4101e-06 | 5.9393e-05 | 5.004e-07 | 0.10002 | 1.4101e-06 | 0.00059379 | 5.0028e-06 | 200Hz100m | 86 | 0.10002 | 1e-06 | 5.5464e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00055451 | 4.9988e-05 | 200Hz100m100m | -9.6164e-07 | 0.002522 | -3.9289e-06 | 0.00056323 | 0.99999 | 2.8219e-06 | -3.9274e-05 | 5.098e-05
100 | m | 0.1 | 500 | 500 | Hz | 100 | m | 0.1 | 100.1195 | 0.0013686 | 0.12857 | 0.00064866 | -0.0019854 | 0.0019294 | 0.0011644 | 1.4323e-05 | m | 0.001 | 0.10012 | 2.3655e-06 | 0.00012741 | 6.4882e-07 | 0.10012 | 2.3655e-06 | 0.0012725 | 6.4804e-06 | 500Hz100m | 87 | 0.10011 | 1.5811e-06 | 0.00013456 | 5e-06 | 0.10011 | 1.5812e-06 | 0.0013441 | 4.9943e-05 | 500Hz100m100m | -7.2572e-06 | 0.0027372 | 7.1503e-06 | 0.0012973 | 0.99993 | 4.7336e-06 | 7.1514e-05 | 5.1597e-05
100 | m | 0.1 | 1000 | 1 | kHz | 100 | m | 0.1 | 100.3484 | 0.0013042 | 0.27117 | 0.0013827 | -0.0014687 | 0.00038228 | 3.3133e-05 | 0.0014631 | m | 0.001 | 0.10035 | 1.3591e-06 | 0.00027114 | 2.0131e-06 | 0.10035 | 1.3591e-06 | 0.0027019 | 2.006e-05 | 1kHz100m | 88 | 0.10035 | 4.4721e-06 | 0.00026124 | 1e-05 | 0.10035 | 4.4722e-06 | 0.0026034 | 9.9653e-05 | 1kHz100m100m | -2.3449e-06 | 0.0026084 | -9.8959e-06 | 0.0027654 | 0.99998 | 2.7549e-06 | -9.8552e-05 | 0.00010743
100 | m | 0.1 | 2000 | 2 | kHz | 100 | m | 0.1 | 100.954 | 0.00096081 | 0.51483 | 0.0011988 | -0.00099031 | 0.0016135 | 0.00049362 | 0.00044222 | m | 0.001 | 0.10095 | 1.8779e-06 | 0.00051433 | 1.2778e-06 | 0.10096 | 1.8779e-06 | 0.0050947 | 1.2657e-05 | 2kHz100m | 89 | 0.10095 | 1.2649e-05 | 0.00051005 | 2e-05 | 0.10095 | 1.2649e-05 | 0.0050526 | 0.00019812 | 2kHz100m100m | -8.3635e-06 | 0.0019217 | -4.2838e-06 | 0.0023977 | 0.99992 | 3.9669e-06 | -4.2013e-05 | 0.00019973
100 | m | 0.1 | 5000 | 5 | kHz | 100 | m | 0.1 | 103.7748 | 0.001786 | 1.3067 | 0.00086999 | -0.0008783 | 6.2878e-05 | 0.00052527 | 0.00054938 | m | 0.001 | 0.10378 | 1.7871e-06 | 0.0013062 | 1.0289e-06 | 0.10378 | 1.787e-06 | 0.012586 | 9.9157e-06 | 5kHz100m | 90 | 0.10379 | 5e-05 | 0.0013075 | 5e-05 | 0.10379 | 5e-05 | 0.012598 | 0.00048172 | 5kHz100m100m | 9.8692e-06 | 0.0035723 | 1.3424e-06 | 0.0017407 | 1.0001 | 6.3009e-06 | 1.1736e-05 | 0.00048213

## Ref Z
This is sheet containing calibration data of reference standards to be used for calibration of RLC meter.
f | f |  | Nominal Z |  | Nom Z | Rs | U(Rs) | Xs | U(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | tag
Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | -
0.01 | 10 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10mHz0m
0.02 | 20 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20mHz0m
0.05 | 50 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50mHz0m
0.1 | 100 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100mHz0m
0.2 | 200 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200mHz0m
0.5 | 500 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500mHz0m
1 | 1 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1Hz0m
2 | 2 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2Hz0m
5 | 5 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5Hz0m
10 | 10 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10Hz0m
20 | 20 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20Hz0m
50 | 50 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50Hz0m
100 | 100 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100Hz0m
200 | 200 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200Hz0m
500 | 500 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500Hz0m
1000 | 1 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1kHz0m
2000 | 2 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2kHz0m
5000 | 5 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5kHz0m
0.01 | 10 | mHz | 1 | m | 0.001 | 0.99989 | 1e-05 | -2.6315e-05 | 5e-05 | m | 0.001 | 0.00099989 | 1e-08 | -2.6315e-08 | 5e-08 | 10mHz1m
0.02 | 20 | mHz | 1 | m | 0.001 | 0.9999 | 1e-05 | -6.5946e-05 | 5e-05 | m | 0.001 | 0.0009999 | 1e-08 | -6.5946e-08 | 5e-08 | 20mHz1m
0.05 | 50 | mHz | 1 | m | 0.001 | 0.99987 | 1e-05 | 5.5443e-05 | 5e-05 | m | 0.001 | 0.00099987 | 1e-08 | 5.5443e-08 | 5e-08 | 50mHz1m
0.1 | 100 | mHz | 1 | m | 0.001 | 0.99994 | 1e-05 | -7.6078e-05 | 5e-05 | m | 0.001 | 0.00099994 | 1e-08 | -7.6078e-08 | 5e-08 | 100mHz1m
0.2 | 200 | mHz | 1 | m | 0.001 | 0.99993 | 1e-05 | -3.1994e-05 | 5e-05 | m | 0.001 | 0.00099993 | 1e-08 | -3.1994e-08 | 5e-08 | 200mHz1m
0.5 | 500 | mHz | 1 | m | 0.001 | 0.99991 | 1e-05 | 0.00015338 | 5e-05 | m | 0.001 | 0.00099991 | 1e-08 | 1.5338e-07 | 5e-08 | 500mHz1m
1 | 1 | Hz | 1 | m | 0.001 | 1 | 1e-05 | -2.0346e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | -2.0346e-08 | 5e-08 | 1Hz1m
2 | 2 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | 1.5098e-05 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 1.5098e-08 | 5e-08 | 2Hz1m
5 | 5 | Hz | 1 | m | 0.001 | 1 | 1e-05 | -1.9943e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | -1.9943e-08 | 5e-08 | 5Hz1m
10 | 10 | Hz | 1 | m | 0.001 | 0.99983 | 1e-05 | 0.00016366 | 5e-05 | m | 0.001 | 0.00099983 | 1e-08 | 1.6366e-07 | 5e-08 | 10Hz1m
20 | 20 | Hz | 1 | m | 0.001 | 0.99993 | 1e-05 | -3.5291e-05 | 5e-05 | m | 0.001 | 0.00099993 | 1e-08 | -3.5291e-08 | 5e-08 | 20Hz1m
50 | 50 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | 0.0002303 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 2.303e-07 | 5e-08 | 50Hz1m
100 | 100 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | 0.00030766 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 3.0766e-07 | 5e-08 | 100Hz1m
200 | 200 | Hz | 1 | m | 0.001 | 1.0004 | 1e-05 | 0.00052145 | 5e-05 | m | 0.001 | 0.0010004 | 1e-08 | 5.2145e-07 | 5e-08 | 200Hz1m
500 | 500 | Hz | 1 | m | 0.001 | 1.0011 | 1.5811e-05 | 0.0013242 | 5e-05 | m | 0.001 | 0.0010011 | 1.5811e-08 | 1.3242e-06 | 5e-08 | 500Hz1m
1000 | 1 | kHz | 1 | m | 0.001 | 1.0033 | 4.4721e-05 | 0.0024578 | 0.0001 | m | 0.001 | 0.0010033 | 4.4721e-08 | 2.4578e-06 | 1e-07 | 1kHz1m
2000 | 2 | kHz | 1 | m | 0.001 | 1.0095 | 0.00012649 | 0.0051647 | 0.0002 | m | 0.001 | 0.0010095 | 1.2649e-07 | 5.1647e-06 | 2e-07 | 2kHz1m
5000 | 5 | kHz | 1 | m | 0.001 | 1.0378 | 0.0005 | 0.012944 | 0.0005 | m | 0.001 | 0.0010378 | 5e-07 | 1.2944e-05 | 5e-07 | 5kHz1m
0.01 | 10 | mHz | 3 | m | 0.003 | 3 | 3e-05 | 0.00038588 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | 3.8588e-07 | 1.5e-07 | 10mHz3m
0.02 | 20 | mHz | 3 | m | 0.003 | 2.9998 | 3e-05 | -3.6688e-06 | 0.00015 | m | 0.001 | 0.0029998 | 3e-08 | -3.6688e-09 | 1.5e-07 | 20mHz3m
0.05 | 50 | mHz | 3 | m | 0.003 | 3.0008 | 3e-05 | -0.00057727 | 0.00015 | m | 0.001 | 0.0030008 | 3e-08 | -5.7727e-07 | 1.5e-07 | 50mHz3m
0.1 | 100 | mHz | 3 | m | 0.003 | 3 | 3e-05 | 0.00011575 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | 1.1575e-07 | 1.5e-07 | 100mHz3m
0.2 | 200 | mHz | 3 | m | 0.003 | 2.9997 | 3e-05 | -0.00016619 | 0.00015 | m | 0.001 | 0.0029997 | 3e-08 | -1.6619e-07 | 1.5e-07 | 200mHz3m
0.5 | 500 | mHz | 3 | m | 0.003 | 2.9998 | 3e-05 | -5.5365e-05 | 0.00015 | m | 0.001 | 0.0029998 | 3e-08 | -5.5365e-08 | 1.5e-07 | 500mHz3m
1 | 1 | Hz | 3 | m | 0.003 | 2.9996 | 3e-05 | -0.00029816 | 0.00015 | m | 0.001 | 0.0029996 | 3e-08 | -2.9816e-07 | 1.5e-07 | 1Hz3m
2 | 2 | Hz | 3 | m | 0.003 | 3.0005 | 3e-05 | 7.0631e-05 | 0.00015 | m | 0.001 | 0.0030005 | 3e-08 | 7.0631e-08 | 1.5e-07 | 2Hz3m
5 | 5 | Hz | 3 | m | 0.003 | 2.9997 | 3e-05 | 0.00030067 | 0.00015 | m | 0.001 | 0.0029997 | 3e-08 | 3.0067e-07 | 1.5e-07 | 5Hz3m
10 | 10 | Hz | 3 | m | 0.003 | 3.0002 | 3e-05 | -5.5728e-05 | 0.00015 | m | 0.001 | 0.0030002 | 3e-08 | -5.5728e-08 | 1.5e-07 | 10Hz3m
20 | 20 | Hz | 3 | m | 0.003 | 3.0007 | 3e-05 | 0.00064789 | 0.00015 | m | 0.001 | 0.0030007 | 3e-08 | 6.4789e-07 | 1.5e-07 | 20Hz3m
50 | 50 | Hz | 3 | m | 0.003 | 3.0001 | 3e-05 | 0.00061494 | 0.00015 | m | 0.001 | 0.0030001 | 3e-08 | 6.1494e-07 | 1.5e-07 | 50Hz3m
100 | 100 | Hz | 3 | m | 0.003 | 3.0002 | 3e-05 | 0.00090165 | 0.00015 | m | 0.001 | 0.0030002 | 3e-08 | 9.0165e-07 | 1.5e-07 | 100Hz3m
200 | 200 | Hz | 3 | m | 0.003 | 3.0006 | 3e-05 | 0.0017291 | 0.00015 | m | 0.001 | 0.0030006 | 3e-08 | 1.7291e-06 | 1.5e-07 | 200Hz3m
500 | 500 | Hz | 3 | m | 0.003 | 3.0038 | 4.7434e-05 | 0.0037346 | 0.00015 | m | 0.001 | 0.0030038 | 4.7434e-08 | 3.7346e-06 | 1.5e-07 | 500Hz3m
1000 | 1 | kHz | 3 | m | 0.003 | 3.0106 | 0.00013416 | 0.0079216 | 0.0003 | m | 0.001 | 0.0030106 | 1.3416e-07 | 7.9216e-06 | 3e-07 | 1kHz3m
2000 | 2 | kHz | 3 | m | 0.003 | 3.0284 | 0.00037947 | 0.015762 | 0.0006 | m | 0.001 | 0.0030284 | 3.7947e-07 | 1.5762e-05 | 6e-07 | 2kHz3m
5000 | 5 | kHz | 3 | m | 0.003 | 3.1137 | 0.0015 | 0.038872 | 0.0015 | m | 0.001 | 0.0031137 | 1.5e-06 | 3.8872e-05 | 1.5e-06 | 5kHz3m
0.01 | 10 | mHz | 10 | m | 0.01 | 9.9982 | 0.0001 | 0.0013593 | 0.0005 | m | 0.001 | 0.0099982 | 1e-07 | 1.3593e-06 | 5e-07 | 10mHz10m
0.02 | 20 | mHz | 10 | m | 0.01 | 9.9996 | 0.0001 | -0.0012518 | 0.0005 | m | 0.001 | 0.0099996 | 1e-07 | -1.2518e-06 | 5e-07 | 20mHz10m
0.05 | 50 | mHz | 10 | m | 0.01 | 9.9998 | 0.0001 | 0.00033603 | 0.0005 | m | 0.001 | 0.0099998 | 1e-07 | 3.3603e-07 | 5e-07 | 50mHz10m
0.1 | 100 | mHz | 10 | m | 0.01 | 10.0004 | 0.0001 | 0.00020039 | 0.0005 | m | 0.001 | 0.01 | 1e-07 | 2.0039e-07 | 5e-07 | 100mHz10m
0.2 | 200 | mHz | 10 | m | 0.01 | 10.0021 | 0.0001 | -0.00071478 | 0.0005 | m | 0.001 | 0.010002 | 1e-07 | -7.1478e-07 | 5e-07 | 200mHz10m
0.5 | 500 | mHz | 10 | m | 0.01 | 10.0002 | 0.0001 | 0.0010089 | 0.0005 | m | 0.001 | 0.01 | 1e-07 | 1.0089e-06 | 5e-07 | 500mHz10m
1 | 1 | Hz | 10 | m | 0.01 | 10.0004 | 0.0001 | -8.1803e-05 | 0.0005 | m | 0.001 | 0.01 | 1e-07 | -8.1803e-08 | 5e-07 | 1Hz10m
2 | 2 | Hz | 10 | m | 0.01 | 10.0004 | 0.0001 | 0.001618 | 0.0005 | m | 0.001 | 0.01 | 1e-07 | 1.618e-06 | 5e-07 | 2Hz10m
5 | 5 | Hz | 10 | m | 0.01 | 9.9991 | 0.0001 | 0.00013304 | 0.0005 | m | 0.001 | 0.0099991 | 1e-07 | 1.3304e-07 | 5e-07 | 5Hz10m
10 | 10 | Hz | 10 | m | 0.01 | 9.9999 | 0.0001 | 0.0020509 | 0.0005 | m | 0.001 | 0.0099999 | 1e-07 | 2.0509e-06 | 5e-07 | 10Hz10m
20 | 20 | Hz | 10 | m | 0.01 | 10.0016 | 0.0001 | 0.00040912 | 0.0005 | m | 0.001 | 0.010002 | 1e-07 | 4.0912e-07 | 5e-07 | 20Hz10m
50 | 50 | Hz | 10 | m | 0.01 | 9.9996 | 0.0001 | 0.0035276 | 0.0005 | m | 0.001 | 0.0099996 | 1e-07 | 3.5276e-06 | 5e-07 | 50Hz10m
100 | 100 | Hz | 10 | m | 0.01 | 10.0013 | 0.0001 | 0.0029473 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 2.9473e-06 | 5e-07 | 100Hz10m
200 | 200 | Hz | 10 | m | 0.01 | 10.0032 | 0.0001 | 0.0040264 | 0.0005 | m | 0.001 | 0.010003 | 1e-07 | 4.0264e-06 | 5e-07 | 200Hz10m
500 | 500 | Hz | 10 | m | 0.01 | 10.014 | 0.00015811 | 0.013086 | 0.0005 | m | 0.001 | 0.010014 | 1.5811e-07 | 1.3086e-05 | 5e-07 | 500Hz10m
1000 | 1 | kHz | 10 | m | 0.01 | 10.034 | 0.00044721 | 0.025733 | 0.001 | m | 0.001 | 0.010034 | 4.4721e-07 | 2.5733e-05 | 1e-06 | 1kHz10m
2000 | 2 | kHz | 10 | m | 0.01 | 10.0963 | 0.0012649 | 0.053584 | 0.002 | m | 0.001 | 0.010096 | 1.2649e-06 | 5.3584e-05 | 2e-06 | 2kHz10m
5000 | 5 | kHz | 10 | m | 0.01 | 10.3789 | 0.005 | 0.13086 | 0.005 | m | 0.001 | 0.010379 | 5e-06 | 0.00013086 | 5e-06 | 5kHz10m
0.01 | 10 | mHz | 100 | m | 0.1 | 100.0202 | 0.001 | -0.011852 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | -1.1852e-05 | 5e-06 | 10mHz100m
0.02 | 20 | mHz | 100 | m | 0.1 | 99.9942 | 0.001 | -0.0029683 | 0.005 | m | 0.001 | 0.099994 | 1e-06 | -2.9683e-06 | 5e-06 | 20mHz100m
0.05 | 50 | mHz | 100 | m | 0.1 | 99.9992 | 0.001 | 0.0036376 | 0.005 | m | 0.001 | 0.099999 | 1e-06 | 3.6376e-06 | 5e-06 | 50mHz100m
0.1 | 100 | mHz | 100 | m | 0.1 | 100.0024 | 0.001 | 0.00097792 | 0.005 | m | 0.001 | 0.1 | 1e-06 | 9.7792e-07 | 5e-06 | 100mHz100m
0.2 | 200 | mHz | 100 | m | 0.1 | 100.0074 | 0.001 | 0.00033939 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | 3.3939e-07 | 5e-06 | 200mHz100m
0.5 | 500 | mHz | 100 | m | 0.1 | 99.9934 | 0.001 | -0.0057838 | 0.005 | m | 0.001 | 0.099993 | 1e-06 | -5.7838e-06 | 5e-06 | 500mHz100m
1 | 1 | Hz | 100 | m | 0.1 | 100.0171 | 0.001 | 0.02386 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 2.386e-05 | 5e-06 | 1Hz100m
2 | 2 | Hz | 100 | m | 0.1 | 99.9962 | 0.001 | 0.0076805 | 0.005 | m | 0.001 | 0.099996 | 1e-06 | 7.6805e-06 | 5e-06 | 2Hz100m
5 | 5 | Hz | 100 | m | 0.1 | 99.9905 | 0.001 | -0.028674 | 0.005 | m | 0.001 | 0.099991 | 1e-06 | -2.8674e-05 | 5e-06 | 5Hz100m
10 | 10 | Hz | 100 | m | 0.1 | 100.0054 | 0.001 | 0.0042077 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | 4.2077e-06 | 5e-06 | 10Hz100m
20 | 20 | Hz | 100 | m | 0.1 | 100.0081 | 0.001 | 0.013861 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | 1.3861e-05 | 5e-06 | 20Hz100m
50 | 50 | Hz | 100 | m | 0.1 | 100.0076 | 0.001 | 0.0082003 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | 8.2003e-06 | 5e-06 | 50Hz100m
100 | 100 | Hz | 100 | m | 0.1 | 100.0167 | 0.001 | 0.029508 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 2.9508e-05 | 5e-06 | 100Hz100m
200 | 200 | Hz | 100 | m | 0.1 | 100.0235 | 0.001 | 0.055464 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 5.5464e-05 | 5e-06 | 200Hz100m
500 | 500 | Hz | 100 | m | 0.1 | 100.1142 | 0.0015811 | 0.13456 | 0.005 | m | 0.001 | 0.10011 | 1.5811e-06 | 0.00013456 | 5e-06 | 500Hz100m
1000 | 1 | kHz | 100 | m | 0.1 | 100.3475 | 0.0044721 | 0.26124 | 0.01 | m | 0.001 | 0.10035 | 4.4721e-06 | 0.00026124 | 1e-05 | 1kHz100m
2000 | 2 | kHz | 100 | m | 0.1 | 100.9466 | 0.012649 | 0.51005 | 0.02 | m | 0.001 | 0.10095 | 1.2649e-05 | 0.00051005 | 2e-05 | 2kHz100m
5000 | 5 | kHz | 100 | m | 0.1 | 103.7856 | 0.05 | 1.3075 | 0.05 | m | 0.001 | 0.10379 | 5e-05 | 0.0013075 | 5e-05 | 5kHz100m

## Ref Z list
Ref. Z list
¦Z¦ |  | ¦Z¦ | Z tag
Ohm |  | Ohm | -
0 | m | 0 | 0m
1 | m | 0.001 | 1m
3 | m | 0.003 | 3m
10 | m | 0.01 | 10m
100 | m | 0.1 | 100m

## Ranges
Range list
Range |  | Range Z | Min ¦Z¦ |  | Min Z | Max ¦Z¦ |  | Max Z
Ohm |  | Ohm | Ohm |  | Ohm | Ohm |  | Ohm
1 | m | 0.001 | 0 | m | 0 | 1 | m | 0.001
3 | m | 0.003 | 1 | m | 0.001 | 3 | m | 0.003
10 | m | 0.01 | 3 | m | 0.003 | 10 | m | 0.01
100 | m | 0.1 | 10 | m | 0.01 | 100 | m | 0.1