## Measurement
This sheet uses correction data to fix bridge errors.
[5 empty] | Measured Z |  |  |  |  | Corrected Z (results) | [5 empty] | Validation data (debuging) |  |  | Deviation of XLS calculated from Octave calculated: |  |  | Measured Z (basic units) | [25 empty] | Interpolated correction |  |  |  | corrected Z | [11 empty] | Interpolated correction |  |  |  | corrected Z
Range |  | f | f |  | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | Fixed Rs | Fixed U(Rs) | Fixed Xs | Fixed U(Xs) |  | Status | Valid Rs | Valid Xs |  | Dev Rs | Dev Xs |  | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | ¦Z¦ | phi | lower Z ID | lower Z tag | higher Z tag | lower Z | higher Z | lower tag | higher tag | lower ID | higher ID | low check | low cor(Rs) | low U(cor(Rs)) | low cor(Xs) | low U(cor(Xs)) | high check | high cor(Rs) | high U(cor(Rs)) | high cor(Xs) | high U(cor(Xs)) | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | Rs | U(Rs) | Xs | U(Xs) | low cor(gain) |  | low cor(phi) |  | hi cor(gain) |  | hi cor(phi) |  | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | ¦Z¦ | phi | Rs | U(Rs) | Xs | U(Xs)
Ohm |  | Hz | - |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm |  | - | Ohm | Ohm |  | Ohm | Ohm |  | - | Ohm | Ohm | Ohm | Ohm | Ohm | rad | - | - | - | Ohm | Ohm | - | - | - | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - |  | rad |  | - |  | rad |  | - | - | rad | rad | Ohm | rad | Ohm | Ohm | Ohm | Ohm
3 | m | 5000 | 5 | kHz | -0.28812 | 0 | -1.37 | 0 | m | -0.28801 | 0 | -1.37 | 0 | m | OK | -0.288 | -1.37 | m | -6.2931e-06 | -7.3811e-06 | m | 0.001 | -0.00028812 | 0 | -0.00137 | 0 | 0.0014 | -1.7781 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | -3.2349e-08 | 0.0038778 | 1.0472e-07 | 0.0061575 | 5kHz3m3m | 3.1454e-07 | 0.0080214 | 9.0519e-08 | 0.00038508 | 3.7022e-08 | 0.0080214 | 1.0188e-07 | 0.0050031 | -0.00028809 | 0.0080214 | -0.0013699 | 0.0050031 | 0.99997 | 0 | 9.7633e-05 | 0 | 1.0001 | 0 | 2.7996e-05 | 0 | 1 | 0 | 8.3707e-05 | 0 | 0.0014 | -1.778 | -0.00028801 | 0 | -0.00137 | 0
100 | m | 10 | 10 | Hz | -24.7996 | 0 | -4.528 | 0 | m | -24.8 | 0 | -4.53 | 0 | m | OK | -24.8 | -4.53 | m | -5.0802e-09 | 2.592e-08 | m | 0.001 | -0.0248 | 0 | -0.004528 | 0 | 0.02521 | -2.961 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | 3.7878e-07 | 0.0082107 | 1e-06 | 0.0030243 | 10Hz100m100m | -1.7212e-06 | 0.010095 | -3.9909e-06 | 0.0061855 | 2.39e-08 | 0.010095 | 1.5658e-07 | 0.0035585 | -0.0248 | 0.010095 | -0.0045278 | 0.0035585 | 1 | 0 | 9.9998e-05 | 0 | 0.99998 | 0 | -3.9911e-05 | 0 | 1 | 0 | 7.6354e-05 | 0 | 0.02521 | -2.9609 | -0.0248 | 0 | -0.00453 | 0
100 | m | 2000 | 2 | kHz | -49.1003 | 0 | -8.4284 | 0 | m | -49.1 | 0 | -8.4299 | 0 | m | OK | -49.1 | -8.43 | m | 8.215e-06 | 8.0489e-05 | m | 0.001 | -0.0491 | 0 | -0.0084284 | 0 | 0.049818 | -2.9716 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | 1.8479e-07 | 0.0048235 | 1.0154e-06 | 0.0050493 | 2kHz100m100m | -2.5564e-06 | 0.010205 | -5.671e-06 | 0.0090173 | -1.028e-06 | 0.010205 | -1.9429e-06 | 0.0068048 | -0.049101 | 0.010205 | -0.0084304 | 0.0068048 | 1 | 0 | 9.9687e-05 | 0 | 0.99997 | 0 | -5.566e-05 | 0 | 1 | 0 | 3.0957e-05 | 0 | 0.049818 | -2.9716 | -0.0491 | 0 | -0.0084299 | 0
1 | m | 50 | 50 | Hz | -0.18599 | 0 | -0.26 | 0 | m | -0.18601 | 0 | -0.25999 | 0 | m | OK | -0.186 | -0.26 | m | -1.1672e-05 | 7.29e-06 | m | 0.001 | -0.00018599 | 0 | -0.00026 | 0 | 0.00031968 | -2.1918 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | 0 | 0.00075911 | 0 | 0.0021036 | 50Hz1m1m | 3.9691e-09 | 0.00045661 | -6.315e-08 | 0.0078449 | 1.2688e-09 | 0.00045661 | -6.315e-08 | 0.003939 | -0.00018599 | 0.00045661 | -0.00026007 | 0.003939 | 1 | 0 | -6.3149e-05 | 0 | 1 | 0 | -6.3149e-05 | 0 | 1 | 0 | -6.3149e-05 | 0 | 0.00031968 | -2.1918 | -0.00018601 | 0 | -0.00025999 | 0
1 | m | 2 | 2 | Hz | 0.56199 | 0 | 0.11498 | 0 | m | 0.56201 | 0 | 0.11501 | 0 | m | OK | 0.562 | 0.115 | m | 7.1792e-06 | 1.1289e-05 | m | 0.001 | 0.00056199 | 0 | 0.00011498 | 0 | 0.00057363 | 0.20182 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.0060159 | 0 | 0.0056977 | 2Hz1m1m | 3.8011e-08 | 0.0078244 | 3.9332e-08 | 0.002428 | 2.1804e-08 | 0.0078244 | 3.9332e-08 | 0.0038221 | 0.00056201 | 0.0078244 | 0.00011502 | 0.0038221 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 0.00057365 | 0.20186 | 0.00056201 | 0 | 0.00011501 | 0
10 | m | 1000 | 1 | kHz | -9.0693 | 0 | -3.7904 | 0 | m | -9.07 | 0 | -3.79 | 0 | m | OK | -9.07 | -3.79 | m | -1.7864e-05 | 9.1957e-06 | m | 0.001 | -0.0090693 | 0 | -0.0037904 | 0 | 0.0098296 | -2.7457 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.7712e-07 | 0.0081916 | 3.0055e-07 | 0.0022801 | 1kHz10m10m | 5.1261e-07 | 0.0034662 | -7.0524e-07 | 0.0068801 | 4.9338e-07 | 0.0034662 | -6.8075e-07 | 0.0067681 | -0.0090688 | 0.0034662 | -0.0037911 | 0.0067681 | 0.99991 | 0 | 9.9528e-05 | 0 | 1.0001 | 0 | -7.0075e-05 | 0 | 1 | 0 | -6.5945e-05 | 0 | 0.00983 | -2.7458 | -0.00907 | 0 | -0.00379 | 0
3 | m | 0.5 | 500 | mHz | 0.64392 | 0 | 2.5701 | 0 | m | 0.644 | 0 | 2.57 | 0 | m | OK | 0.644 | 2.57 | m | -9.292e-09 | 8.7565e-09 | m | 0.001 | 0.00064392 | 0 | 0.0025701 | 0 | 0.0026496 | 1.3253 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | 8.5627e-09 | 0.008001 | 9.9999e-08 | 0.0090527 | 500mHz3m3m | -1.7973e-07 | 0.0052338 | -2.1531e-07 | 0.0055561 | -1.4674e-07 | 0.0052338 | -1.6006e-07 | 0.0061687 | 0.00064378 | 0.0052338 | 0.00257 | 0.0061687 | 1 | 0 | 0.0001 | 0 | 0.99994 | 0 | -7.1769e-05 | 0 | 0.99995 | 0 | -4.1673e-05 | 0 | 0.0026495 | 1.3253 | 0.000644 | 0 | 0.00257 | 0
1 | m | 0.05 | 50 | mHz | -0.26301 | 0 | 0.20401 | 0 | m | -0.26298 | 0 | 0.20398 | 0 | m | OK | -0.263 | 0.204 | m | 1.5064e-05 | -1.5957e-05 | m | 0.001 | -0.00026301 | 0 | 0.00020401 | 0 | 0.00033285 | 2.4819 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.0049811 | 0 | 0.0050777 | 50mHz1m1m | -9.7658e-08 | 0.00076464 | 1.519e-08 | 0.008821 | -3.2506e-08 | 0.00076464 | 1.519e-08 | 0.0063237 | -0.00026304 | 0.00076464 | 0.00020402 | 0.0063237 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.00033282 | 2.4819 | -0.00026298 | 0 | 0.00020398 | 0
10 | m | 1000 | 1 | kHz | 7.64 | 0 | -0.11191 | 0 | m | 7.64 | 0 | -0.11201 | 0 | m | OK | 7.64 | -0.112 | m | 7.7017e-06 | -9.3902e-06 | m | 0.001 | 0.00764 | 0 | -0.00011191 | 0 | 0.0076408 | -0.014647 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.7712e-07 | 0.0081916 | 3.0055e-07 | 0.0022801 | 1kHz10m10m | 5.1261e-07 | 0.0034662 | -7.0524e-07 | 0.0068801 | 2.4645e-07 | 0.0034662 | -3.6626e-07 | 0.0053298 | 0.0076402 | 0.0034662 | -0.00011228 | 0.0053298 | 0.99991 | 0 | 9.9528e-05 | 0 | 1.0001 | 0 | -7.0075e-05 | 0 | 1 | 0 | -1.2914e-05 | 0 | 0.0076408 | -0.01466 | 0.00764 | 0 | -0.00011201 | 0
3 | m | 0.2 | 200 | mHz | -2.0098 | 0 | 0.63604 | 0 | m | -2.01 | 0 | 0.636 | 0 | m | OK | -2.01 | 0.636 | m | -5.169e-09 | -1.4515e-08 | m | 0.001 | -0.0020098 | 0 | 0.00063604 | 0 | 0.0021081 | 2.8351 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | 8.0669e-08 | 0.0027077 | 9.999e-08 | 0.002112 | 200mHz3m3m | 2.3152e-07 | 0.0088393 | 4.0846e-09 | 0.0047677 | 1.6424e-07 | 0.0088393 | 4.6856e-08 | 0.0035834 | -0.0020096 | 0.0088393 | 0.00063609 | 0.0035834 | 1.0001 | 0 | 9.9999e-05 | 0 | 1.0001 | 0 | 1.3572e-06 | 0 | 1.0001 | 0 | 4.5349e-05 | 0 | 0.0021082 | 2.8351 | -0.00201 | 0 | 0.000636 | 0
100 | m | 10 | 10 | Hz | 1.6318 | 0 | 52.0993 | 0 | m | 1.63 | 0 | 52.1 | 0 | m | OK | 1.63 | 52.1 | m | -5.5996e-08 | 6.3948e-08 | m | 0.001 | 0.0016318 | 0 | 0.052099 | 0 | 0.052125 | 1.5395 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | 3.7878e-07 | 0.0082107 | 1e-06 | 0.0030243 | 10Hz100m100m | -1.7212e-06 | 0.010095 | -3.9909e-06 | 0.0061855 | -6.0411e-07 | 0.010095 | -1.336e-06 | 0.0045039 | 0.0016312 | 0.010095 | 0.052098 | 0.0045039 | 1 | 0 | 9.9998e-05 | 0 | 0.99998 | 0 | -3.9911e-05 | 0 | 1 | 0 | 3.4513e-05 | 0 | 0.052125 | 1.5395 | 0.00163 | 0 | 0.0521 | 0
1 | m | 0.2 | 200 | mHz | 0.49101 | 0 | 0.31197 | 0 | m | 0.49099 | 0 | 0.31202 | 0 | m | OK | 0.491 | 0.312 | m | -8.7195e-06 | 2.0429e-05 | m | 0.001 | 0.00049101 | 0 | 0.00031197 | 0 | 0.00058174 | 0.56602 | 1 | 0m | 1m | 0 | 0.001 | 200mHz1m0m | 200mHz1m1m | 5 | 23 | 200mHz1m0m | 0 | 0.0073611 | 0 | 0.0042415 | 200mHz1m1m | 1.4791e-08 | 0.0044457 | 9.0079e-08 | 0.0099916 | 8.6046e-09 | 0.0044457 | 9.0079e-08 | 0.0075865 | 0.00049102 | 0.0044457 | 0.00031206 | 0.0075865 | 1 | 0 | 9.0079e-05 | 0 | 1 | 0 | 9.0079e-05 | 0 | 1 | 0 | 9.0079e-05 | 0 | 0.00058175 | 0.56611 | 0.00049099 | 0 | 0.00031202 | 0
10 | m | 10 | 10 | Hz | -2.8203 | 0 | -4.7201 | 0 | m | -2.82 | 0 | -4.72 | 0 | m | OK | -2.82 | -4.72 | m | -9.5714e-08 | -9.8996e-09 | m | 0.001 | -0.0028203 | 0 | -0.0047201 | 0 | 0.0054985 | -2.1094 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | -7.1505e-08 | 0.0049286 | -1.0912e-08 | 0.0054811 | -0.0028204 | 0.0049286 | -0.0047201 | 0.0054811 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 0.99996 | 0 | 4.3926e-05 | 0 | 0.0054983 | -2.1093 | -0.00282 | 0 | -0.00472 | 0
1 | m | 20 | 20 | Hz | -0.53498 | 0 | 0.156 | 0 | m | -0.53501 | 0 | 0.156 | 0 | m | OK | -0.535 | 0.156 | m | -1.2225e-05 | 1.6207e-06 | m | 0.001 | -0.00053498 | 0 | 0.000156 | 0 | 0.00055726 | 2.8579 | 1 | 0m | 1m | 0 | 0.001 | 20Hz1m0m | 20Hz1m1m | 11 | 29 | 20Hz1m0m | 0 | 0.0036421 | 0 | 0.008204 | 20Hz1m1m | 4.9405e-08 | 0.0045511 | 7.5608e-09 | 5.0763e-05 | 2.7531e-08 | 0.0045511 | 7.5608e-09 | 0.0036605 | -0.00053496 | 0.0045511 | 0.00015601 | 0.0036605 | 1 | 0 | 7.5635e-06 | 0 | 1 | 0 | 7.5635e-06 | 0 | 1 | 0 | 7.5635e-06 | 0 | 0.00055729 | 2.8579 | -0.00053501 | 0 | 0.000156 | 0
10 | m | 10 | 10 | Hz | -2.0304 | 0 | -2.7 | 0 | m | -2.03 | 0 | -2.7 | 0 | m | OK | -2.03 | -2.7 | m | -2.1333e-08 | 8.8896e-10 | m | 0.001 | -0.0020304 | 0 | -0.0027 | 0 | 0.0033782 | -2.2156 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | -2e-07 | 0.0049286 | 2.5293e-07 | 0.0075309 | -0.0020306 | 0.0049286 | -0.0026997 | 0.0075309 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 0.99993 | 0 | 9.1508e-05 | 0 | 0.003378 | -2.2155 | -0.00203 | 0 | -0.0027 | 0
10 | m | 100 | 100 | Hz | -6.0097 | 0 | 2.7902 | 0 | m | -6.01 | 0 | 2.79 | 0 | m | OK | -6.01 | 2.79 | m | -1.8948e-08 | 1.671e-07 | m | 0.001 | -0.0060097 | 0 | 0.0027902 | 0 | 0.0066258 | 2.7069 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 3.2865e-07 | 0.0013885 | 8.9527e-08 | 0.0056347 | -0.0060094 | 0.0013885 | 0.0027903 | 0.0056347 | 1 | 0 | 9.9996e-05 | 0 | 1.0001 | 0 | -1.0637e-05 | 0 | 1 | 0 | 4.2691e-05 | 0 | 0.006626 | 2.707 | -0.00601 | 0 | 0.00279 | 0
3 | m | 100 | 100 | Hz | -1.9597 | 0 | 0.9241 | 0 | m | -1.96 | 0 | 0.924 | 0 | m | OK | -1.96 | 0.924 | m | -8.8227e-09 | 9.9136e-09 | m | 0.001 | -0.0019597 | 0 | 0.0009241 | 0 | 0.0021667 | 2.701 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | 6.7713e-08 | 0.0095207 | 1.0003e-07 | 0.0097997 | 100Hz3m3m | 2.9949e-07 | 0.0054632 | 2.5697e-07 | 0.0010516 | 2.0292e-07 | 0.0054632 | 1.9158e-07 | 0.0046965 | -0.0019595 | 0.0054632 | 0.00092429 | 0.0046965 | 1.0001 | 0 | 9.9997e-05 | 0 | 1.0001 | 0 | 8.565e-05 | 0 | 1.0001 | 0 | 9.1627e-05 | 0 | 0.0021669 | 2.7011 | -0.00196 | 0 | 0.000924 | 0
1 | m | 5000 | 5 | kHz | 0.0907 | 0 | 0.0393 | 0 | m | 0.090698 | 0 | 0.039301 | 0 | m | OK | 0.0907 | 0.0393 | m | -2.0094e-06 | 6.0412e-07 | m | 0.001 | 9.07e-05 | 0 | 3.93e-05 | 0 | 9.8848e-05 | 0.40888 | 1 | 0m | 1m | 0 | 0.001 | 5kHz1m0m | 5kHz1m1m | 18 | 36 | 5kHz1m0m | 0 | 0.0075446 | 0 | 0.0045896 | 5kHz1m1m | -1.9193e-08 | 0.0023865 | 1.6155e-08 | 0.0095804 | -1.8972e-09 | 0.0023865 | 1.6155e-08 | 0.0050829 | 9.0698e-05 | 0.0023865 | 3.9316e-05 | 0.0050829 | 0.99998 | 0 | 1.5079e-05 | 0 | 0.99998 | 0 | 1.5079e-05 | 0 | 0.99998 | 0 | 1.5079e-05 | 0 | 9.8847e-05 | 0.40889 | 9.0698e-05 | 0 | 3.9301e-05 | 0
10 | m | 50 | 50 | Hz | -6.3604 | 0 | 5.8899 | 0 | m | -6.36 | 0 | 5.89 | 0 | m | OK | -6.36 | 5.89 | m | 3.3347e-07 | 2.5213e-07 | m | 0.001 | -0.0063604 | 0 | 0.0058899 | 0 | 0.0086686 | 2.3946 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | 3.0977e-09 | 0.0015823 | 3.0001e-07 | 0.00091715 | 50Hz10m10m | -3.1031e-07 | 0.006568 | -7.6553e-07 | 0.0090646 | -2.507e-07 | 0.006568 | -5.6287e-07 | 0.007515 | -0.0063607 | 0.006568 | 0.0058893 | 0.007515 | 1 | 0 | 9.9987e-05 | 0 | 0.99997 | 0 | -7.6535e-05 | 0 | 0.99998 | 0 | -4.2962e-05 | 0 | 0.0086684 | 2.3945 | -0.00636 | 0 | 0.00589 | 0
1 | m | 2 | 2 | Hz | 0.59898 | 0 | -0.22601 | 0 | m | 0.59901 | 0 | -0.22599 | 0 | m | OK | 0.599 | -0.226 | m | 1.139e-05 | 5.3867e-06 | m | 0.001 | 0.00059898 | 0 | -0.00022601 | 0 | 0.0006402 | -0.36081 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.0060159 | 0 | 0.0056977 | 2Hz1m1m | 3.8011e-08 | 0.0078244 | 3.9332e-08 | 0.002428 | 2.4335e-08 | 0.0078244 | 3.9332e-08 | 0.0036044 | 0.000599 | 0.0078244 | -0.00022597 | 0.0036044 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 0.00064023 | -0.36077 | 0.00059901 | 0 | -0.00022599 | 0
10 | m | 100 | 100 | Hz | -1.6203 | 0 | -3.1698 | 0 | m | -1.62 | 0 | -3.17 | 0 | m | OK | -1.62 | -3.17 | m | -2.8923e-08 | -2.0962e-09 | m | 0.001 | -0.0016203 | 0 | -0.0031698 | 0 | 0.0035599 | -2.0433 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 5.2648e-08 | 0.0013885 | 2.6751e-07 | 0.005363 | -0.0016202 | 0.0013885 | -0.0031696 | 0.005363 | 1 | 0 | 9.9996e-05 | 0 | 1.0001 | 0 | -1.0637e-05 | 0 | 1 | 0 | 9.1146e-05 | 0 | 0.00356 | -2.0432 | -0.00162 | 0 | -0.00317 | 0
10 | m | 0.2 | 200 | mHz | 0.49636 | 0 | 4.11 | 0 | m | 0.496 | 0 | 4.11 | 0 | m | OK | 0.496 | 4.11 | m | 2.5391e-09 | 2.0535e-09 | m | 0.001 | 0.00049636 | 0 | 0.00411 | 0 | 0.0041399 | 1.4506 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | -6.5943e-08 | 0.0010701 | 2.9999e-07 | 0.00054156 | 200mHz10m10m | 2.3191e-07 | 0.0030196 | 1.3784e-07 | 0.0069261 | -1.744e-08 | 0.0030196 | 2.7358e-07 | 0.0015812 | 0.00049634 | 0.0030196 | 0.0041103 | 0.0015812 | 0.99998 | 0 | 0.0001 | 0 | 1 | 0 | 1.3786e-05 | 0 | 0.99999 | 0 | 8.5964e-05 | 0 | 0.0041398 | 1.4507 | 0.000496 | 0 | 0.00411 | 0
100 | m | 20 | 20 | Hz | 20.0966 | 0 | 84.702 | 0 | m | 20.1 | 0 | 84.7 | 0 | m | OK | 20.1 | 84.7 | m | 4.8369e-07 | -2.8424e-08 | m | 0.001 | 0.020097 | 0 | 0.084702 | 0 | 0.087053 | 1.3378 | 4 | 10m | 100m | 0.01 | 0.1 | 20Hz100m10m | 20Hz100m100m | 119 | 137 | 20Hz100m10m | -3.5692e-07 | 0.0065402 | 9.9995e-07 | 0.0041199 | 20Hz100m100m | -8.9075e-07 | 0.0087139 | -6.7344e-06 | 0.00094046 | -8.1396e-07 | 0.0087139 | -5.6218e-06 | 0.0013978 | 0.020096 | 0.0087139 | 0.084696 | 0.0013978 | 0.99996 | 0 | 0.0001 | 0 | 0.99999 | 0 | -6.734e-05 | 0 | 0.99999 | 0 | -4.3268e-05 | 0 | 0.087052 | 1.3378 | 0.0201 | 0 | 0.0847 | 0
3 | m | 2 | 2 | Hz | -0.79003 | 0 | 2.15 | 0 | m | -0.79 | 0 | 2.15 | 0 | m | OK | -0.79 | 2.15 | m | -3.7802e-09 | 1.6976e-09 | m | 0.001 | -0.00079003 | 0 | 0.00215 | 0 | 0.0022905 | 1.9229 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | -4.7398e-08 | 0.0012746 | 1.0001e-07 | 0.0051494 | 2Hz3m3m | 1.3089e-07 | 0.005742 | -2.5906e-07 | 0.0087132 | 6.7642e-08 | 0.005742 | -1.3169e-07 | 0.007449 | -0.00078997 | 0.005742 | 0.0021498 | 0.007449 | 0.99995 | 0 | 0.0001 | 0 | 1 | 0 | -8.6361e-05 | 0 | 1 | 0 | -2.0251e-05 | 0 | 0.0022905 | 1.9229 | -0.00079 | 0 | 0.00215 | 0
1 | m | 2000 | 2 | kHz | -0.032398 | 0 | -0.29299 | 0 | m | -0.032404 | 0 | -0.29302 | 0 | m | OK | -0.0324 | -0.293 | m | -4.1413e-06 | -2.0363e-05 | m | 0.001 | -3.2398e-05 | 0 | -0.00029299 | 0 | 0.00029478 | -1.6809 | 1 | 0m | 1m | 0 | 0.001 | 2kHz1m0m | 2kHz1m1m | 17 | 35 | 2kHz1m0m | 0 | 0.0082269 | 0 | 0.0083167 | 2kHz1m1m | 1.0061e-07 | 0.0054645 | -9.0811e-09 | 0.0030182 | 2.9657e-08 | 0.0054645 | -9.0811e-09 | 0.0067548 | -3.2369e-05 | 0.0054645 | -0.000293 | 0.0067548 | 1.0001 | 0 | -8.9663e-06 | 0 | 1.0001 | 0 | -8.9663e-06 | 0 | 1.0001 | 0 | -8.9663e-06 | 0 | 0.00029481 | -1.6809 | -3.2404e-05 | 0 | -0.00029302 | 0
10 | m | 10 | 10 | Hz | 4.53 | 0 | 6.4801 | 0 | m | 4.53 | 0 | 6.48 | 0 | m | OK | 4.53 | 6.48 | m | 2.2963e-07 | -3.1449e-09 | m | 0.001 | 0.00453 | 0 | 0.0064801 | 0 | 0.0079065 | 0.9607 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | 7.4434e-08 | 0.0049286 | -3.1056e-07 | 0.0031532 | 0.00453 | 0.0049286 | 0.0064798 | 0.0031532 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 0.99999 | 0 | -1.0114e-05 | 0 | 0.0079064 | 0.96069 | 0.00453 | 0 | 0.00648 | 0
100 | m | 0.2 | 200 | mHz | -41.8975 | 0 | 2.2717 | 0 | m | -41.9 | 0 | 2.27 | 0 | m | OK | -41.9 | 2.27 | m | 1.1025e-08 | 8.1479e-08 | m | 0.001 | -0.041897 | 0 | 0.0022717 | 0 | 0.041959 | 3.0874 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | 6.4711e-07 | 0.0036835 | 9.9995e-07 | 0.0072304 | 200mHz100m100m | 4.5167e-06 | 0.0021714 | -5.5869e-06 | 0.0082059 | 2.0212e-06 | 0.0021714 | -1.339e-06 | 0.0075768 | -0.041895 | 0.0021714 | 0.0022704 | 0.0075768 | 1.0001 | 0 | 9.9997e-05 | 0 | 1 | 0 | -5.5871e-05 | 0 | 1.0001 | 0 | 4.4648e-05 | 0 | 0.041961 | 3.0875 | -0.0419 | 0 | 0.00227 | 0
10 | m | 100 | 100 | Hz | 2.8596 | 0 | -5.01 | 0 | m | 2.86 | 0 | -5.01 | 0 | m | OK | 2.86 | -5.01 | m | -5.6942e-08 | -1.0458e-07 | m | 0.001 | 0.0028596 | 0 | -0.00501 | 0 | 0.0057687 | -1.0521 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 2.5149e-07 | 0.0013885 | 1.3928e-07 | 0.0055588 | 0.0028599 | 0.0013885 | -0.0050099 | 0.0055588 | 1 | 0 | 9.9996e-05 | 0 | 1.0001 | 0 | -1.0637e-05 | 0 | 1 | 0 | 5.6237e-05 | 0 | 0.0057689 | -1.0521 | 0.00286 | 0 | -0.00501 | 0
3 | m | 0.5 | 500 | mHz | -2.6702 | 0 | 1.0799 | 0 | m | -2.67 | 0 | 1.08 | 0 | m | OK | -2.67 | 1.08 | m | -1.084e-08 | -1.0302e-08 | m | 0.001 | -0.0026702 | 0 | 0.0010799 | 0 | 0.0028803 | 2.7573 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | 8.5627e-09 | 0.008001 | 9.9999e-08 | 0.0090527 | 500mHz3m3m | -1.7973e-07 | 0.0052338 | -2.1531e-07 | 0.0055561 | -1.6846e-07 | 0.0052338 | -1.9644e-07 | 0.0057654 | -0.0026704 | 0.0052338 | 0.0010797 | 0.0057654 | 1 | 0 | 0.0001 | 0 | 0.99994 | 0 | -7.1769e-05 | 0 | 0.99994 | 0 | -6.149e-05 | 0 | 0.0028802 | 2.7572 | -0.00267 | 0 | 0.00108 | 0
3 | m | 5 | 5 | Hz | 0.39011 | 0 | 2.1498 | 0 | m | 0.39 | 0 | 2.15 | 0 | m | OK | 0.39 | 2.15 | m | 3.6989e-09 | -1.7023e-09 | m | 0.001 | 0.00039011 | 0 | 0.0021498 | 0 | 0.0021849 | 1.3913 | 2 | 1m | 3m | 0.001 | 0.003 | 5Hz3m1m | 5Hz3m3m | 45 | 63 | 5Hz3m1m | 8.8834e-08 | 0.0045761 | 9.9985e-08 | 0.002973 | 5Hz3m3m | 2.0952e-07 | 0.0076576 | 1.1353e-07 | 0.0066287 | 1.6033e-07 | 0.0076576 | 1.0801e-07 | 0.0051388 | 0.00039027 | 0.0076576 | 0.0021499 | 0.0051388 | 1.0001 | 0 | 0.0001 | 0 | 1.0001 | 0 | 3.7855e-05 | 0 | 1.0001 | 0 | 6.3182e-05 | 0 | 0.0021851 | 1.3914 | 0.00039 | 0 | 0.00215 | 0
100 | m | 1000 | 1 | kHz | 40.2958 | 0 | -26.0989 | 0 | m | 40.3 | 0 | -26.1 | 0 | m | OK | 40.3 | -26.1 | m | -1.3465e-05 | -2.5366e-05 | m | 0.001 | 0.040296 | 0 | -0.026099 | 0 | 0.048009 | -0.57474 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | 8.1145e-07 | 0.0088059 | 1.0054e-06 | 0.0024854 | 1kHz100m100m | 9.6046e-06 | 0.0045809 | -7.0617e-06 | 0.0068005 | 4.525e-06 | 0.0045809 | -2.4015e-06 | 0.0043078 | 0.0403 | 0.0045809 | -0.026101 | 0.0043078 | 1.0001 | 0 | 9.9876e-05 | 0 | 1.0001 | 0 | -7.0183e-05 | 0 | 1.0001 | 0 | 2.8056e-05 | 0 | 0.048014 | -0.57472 | 0.0403 | 0 | -0.0261 | 0
10 | m | 0.05 | 50 | mHz | -7.8903 | 0 | 3.2602 | 0 | m | -7.89 | 0 | 3.26 | 0 | m | OK | -7.89 | 3.26 | m | 5.3512e-09 | 1.0189e-07 | m | 0.001 | -0.0078903 | 0 | 0.0032602 | 0 | 0.0085373 | 2.7498 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -1.966e-07 | 0.0016505 | 3e-07 | 0.00092106 | 50mHz10m10m | -2.6057e-07 | 0.0029905 | -1.9484e-07 | 0.0036795 | -2.472e-07 | 0.0029905 | -9.1434e-08 | 0.0031031 | -0.0078905 | 0.0029905 | 0.0032601 | 0.0031031 | 0.99993 | 0 | 9.9994e-05 | 0 | 0.99997 | 0 | -1.9476e-05 | 0 | 0.99997 | 0 | 5.4889e-06 | 0 | 0.008537 | 2.7498 | -0.00789 | 0 | 0.00326 | 0
10 | m | 2000 | 2 | kHz | -0.68243 | 0 | -5.5901 | 0 | m | -0.68201 | 0 | -5.59 | 0 | m | OK | -0.682 | -5.59 | m | -6.1694e-06 | -2.8112e-06 | m | 0.001 | -0.00068243 | 0 | -0.0055901 | 0 | 0.0056316 | -1.6923 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -9.077e-08 | 0.0076693 | 3.0092e-07 | 0.0045108 | -0.00068252 | 0.0076693 | -0.0055898 | 0.0045108 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99997 | 0 | 7.2994e-05 | 0 | 0.0056315 | -1.6922 | -0.00068201 | 0 | -0.00559 | 0
100 | m | 50 | 50 | Hz | -32.203 | 0 | -72.6024 | 0 | m | -32.2 | 0 | -72.6 | 0 | m | OK | -32.2 | -72.6 | m | 8.5731e-07 | -3.5717e-07 | m | 0.001 | -0.032203 | 0 | -0.072602 | 0 | 0.079424 | -1.9883 | 4 | 10m | 100m | 0.01 | 0.1 | 50Hz100m10m | 50Hz100m100m | 120 | 138 | 50Hz100m10m | -4.4594e-07 | 0.005803 | 1.0003e-06 | 0.0070726 | 50Hz100m100m | -4.2272e-06 | 0.0099288 | -1.5433e-07 | 0.002197 | -3.3627e-06 | 0.0099288 | 1.0964e-07 | 0.0033117 | -0.032206 | 0.0099288 | -0.072602 | 0.0033117 | 0.99996 | 0 | 9.9998e-05 | 0 | 0.99996 | 0 | -1.5405e-06 | 0 | 0.99996 | 0 | 2.1674e-05 | 0 | 0.07942 | -1.9883 | -0.0322 | 0 | -0.0726 | 0
3 | m | 200 | 200 | Hz | -2.2098 | 0 | -0.94584 | 0 | m | -2.21 | 0 | -0.946 | 0 | m | OK | -2.21 | -0.946 | m | -5.5706e-08 | 1.3644e-07 | m | 0.001 | -0.0022098 | 0 | -0.00094584 | 0 | 0.0024037 | -2.7372 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | 9.9186e-08 | 0.0011654 | 1.0001e-07 | 0.0083568 | 200Hz3m3m | 2.9261e-07 | 0.0029843 | 8.5449e-09 | 0.0015311 | 2.3494e-07 | 0.0029843 | 3.5815e-08 | 0.0035661 | -0.0022096 | 0.0029843 | -0.0009458 | 0.0035661 | 1.0001 | 0 | 9.9974e-05 | 0 | 1.0001 | 0 | 2.8448e-06 | 0 | 1.0001 | 0 | 3.1803e-05 | 0 | 0.002404 | -2.7371 | -0.00221 | 0 | -0.000946 | 0
100 | m | 0.1 | 100 | mHz | -29.3967 | 0 | 44.7001 | 0 | m | -29.4 | 0 | 44.7 | 0 | m | OK | -29.4 | 44.7 | m | -1.8651e-07 | 9.4046e-08 | m | 0.001 | -0.029397 | 0 | 0.0447 | 0 | 0.0535 | 2.1525 | 4 | 10m | 100m | 0.01 | 0.1 | 100mHz100m10m | 100mHz100m100m | 112 | 130 | 100mHz100m10m | 7.2738e-07 | 0.00023073 | 9.9998e-07 | 0.0027026 | 100mHz100m100m | -1.0097e-06 | 0.0014513 | 3.4827e-07 | 0.0073655 | -1.1222e-07 | 0.0014513 | 6.8499e-07 | 0.0049563 | -0.029397 | 0.0014513 | 0.044701 | 0.0049563 | 1.0001 | 0 | 0.0001 | 0 | 0.99999 | 0 | 3.4841e-06 | 0 | 1 | 0 | 5.335e-05 | 0 | 0.053502 | 2.1526 | -0.0294 | 0 | 0.0447 | 0
1 | m | 0.5 | 500 | mHz | -0.65496 | 0 | 0.66001 | 0 | m | -0.655 | 0 | 0.66 | 0 | m | OK | -0.655 | 0.66 | m | -3.2063e-06 | -1.0319e-06 | m | 0.001 | -0.00065496 | 0 | 0.00066001 | 0 | 0.00092983 | 2.3523 | 1 | 0m | 1m | 0 | 0.001 | 500mHz1m0m | 500mHz1m1m | 6 | 24 | 500mHz1m0m | 0 | 0.00065229 | 0 | 0.0096122 | 500mHz1m1m | 2.3394e-08 | 0.0025222 | 4.6022e-08 | 0.0036456 | 2.1752e-08 | 0.0025222 | 4.6022e-08 | 0.0040643 | -0.00065494 | 0.0025222 | 0.00066006 | 0.0040643 | 1 | 0 | 4.6024e-05 | 0 | 1 | 0 | 4.6024e-05 | 0 | 1 | 0 | 4.6024e-05 | 0 | 0.00092985 | 2.3524 | -0.000655 | 0 | 0.00066 | 0
10 | m | 200 | 200 | Hz | 3.7098 | 0 | -3.7001 | 0 | m | 3.71 | 0 | -3.7 | 0 | m | OK | 3.71 | -3.7 | m | -2.6666e-08 | -4.3746e-07 | m | 0.001 | 0.0037098 | 0 | -0.0037001 | 0 | 0.0052396 | -0.7841 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -7.8893e-08 | 0.0054354 | 3.0006e-07 | 0.0020919 | 200Hz10m10m | 9.884e-07 | 0.0064562 | -5.1026e-07 | 0.0043182 | 2.6258e-07 | 0.0064562 | 4.0806e-08 | 0.0028042 | 0.00371 | 0.0064562 | -0.0037001 | 0.0028042 | 0.99997 | 0 | 9.9966e-05 | 0 | 1.0001 | 0 | -5.1006e-05 | 0 | 1 | 0 | 5.1664e-05 | 0 | 0.0052397 | -0.78405 | 0.00371 | 0 | -0.0037 | 0
10 | m | 0.02 | 20 | mHz | 4.4502 | 0 | 3.0095 | 0 | m | 4.45 | 0 | 3.01 | 0 | m | OK | 4.45 | 3.01 | m | -4.5621e-10 | 3.8092e-10 | m | 0.001 | 0.0044502 | 0 | 0.0030095 | 0 | 0.0053723 | 0.59462 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 1.1878e-07 | 0.0030642 | 4.9609e-07 | 0.0069003 | 0.0044504 | 0.0030642 | 0.00301 | 0.0069003 | 1 | 0 | 0.0001 | 0 | 1 | 0 | 8.7863e-05 | 0 | 1 | 0 | 9.5887e-05 | 0 | 0.0053724 | 0.59471 | 0.00445 | 0 | 0.00301 | 0
10 | m | 50 | 50 | Hz | 0.40874 | 0 | -4.7101 | 0 | m | 0.409 | 0 | -4.71 | 0 | m | OK | 0.409 | -4.71 | m | -1.1145e-07 | 9.5009e-09 | m | 0.001 | 0.00040874 | 0 | -0.0047101 | 0 | 0.0047278 | -1.4842 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | 3.0977e-09 | 0.0015823 | 3.0001e-07 | 0.00091715 | 50Hz10m10m | -3.1031e-07 | 0.006568 | -7.6553e-07 | 0.0090646 | -7.4258e-08 | 0.006568 | 3.7012e-08 | 0.0029281 | 0.00040866 | 0.006568 | -0.00471 | 0.0029281 | 1 | 0 | 9.9987e-05 | 0 | 0.99997 | 0 | -7.6535e-05 | 0 | 0.99999 | 0 | 5.6418e-05 | 0 | 0.0047277 | -1.4842 | 0.000409 | 0 | -0.00471 | 0
100 | m | 20 | 20 | Hz | 7.9484 | 0 | -28.1014 | 0 | m | 7.95 | 0 | -28.1 | 0 | m | OK | 7.95 | -28.1 | m | -1.666e-08 | -1.1552e-08 | m | 0.001 | 0.0079484 | 0 | -0.028101 | 0 | 0.029204 | -1.2951 | 4 | 10m | 100m | 0.01 | 0.1 | 20Hz100m10m | 20Hz100m100m | 119 | 137 | 20Hz100m10m | -3.5692e-07 | 0.0065402 | 9.9995e-07 | 0.0041199 | 20Hz100m100m | -8.9075e-07 | 0.0087139 | -6.7344e-06 | 0.00094046 | -4.7083e-07 | 0.0087139 | -6.5036e-07 | 0.0034415 | 0.007948 | 0.0087139 | -0.028102 | 0.0034415 | 0.99996 | 0 | 0.0001 | 0 | 0.99999 | 0 | -6.734e-05 | 0 | 0.99997 | 0 | 6.4294e-05 | 0 | 0.029203 | -1.2951 | 0.00795 | 0 | -0.0281 | 0
10 | m | 2 | 2 | Hz | 1.2295 | 0 | -4.4398 | 0 | m | 1.23 | 0 | -4.44 | 0 | m | OK | 1.23 | -4.44 | m | -6.0209e-09 | 9.9599e-11 | m | 0.001 | 0.0012295 | 0 | -0.0044398 | 0 | 0.0046069 | -1.3006 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.9667e-07 | 0.0080028 | 2.9999e-07 | 0.0092189 | 2Hz10m10m | 4.1052e-07 | 0.0037012 | 3.5445e-07 | 0.0031449 | 2.4576e-07 | 0.0037012 | 3.1249e-07 | 0.0078245 | 0.0012298 | 0.0037012 | -0.0044395 | 0.0078245 | 1.0001 | 0 | 0.0001 | 0 | 1 | 0 | 3.5443e-05 | 0 | 1.0001 | 0 | 8.518e-05 | 0 | 0.0046072 | -1.3005 | 0.00123 | 0 | -0.00444 | 0
3 | m | 5000 | 5 | kHz | 0.11587 | 0 | -1.79 | 0 | m | 0.11599 | 0 | -1.79 | 0 | m | OK | 0.116 | -1.79 | m | -6.6977e-06 | -1.445e-05 | m | 0.001 | 0.00011587 | 0 | -0.00179 | 0 | 0.0017937 | -1.5062 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | -3.2349e-08 | 0.0038778 | 1.0472e-07 | 0.0061575 | 5kHz3m3m | 3.1454e-07 | 0.0080214 | 9.0519e-08 | 0.00038508 | 1.0532e-07 | 0.0080214 | 9.9082e-08 | 0.0038666 | 0.00011597 | 0.0080214 | -0.0017899 | 0.0038666 | 0.99997 | 0 | 9.7633e-05 | 0 | 1.0001 | 0 | 2.7996e-05 | 0 | 1 | 0 | 6.9997e-05 | 0 | 0.0017938 | -1.5061 | 0.00011599 | 0 | -0.00179 | 0
3 | m | 50 | 50 | Hz | 0.070674 | 0 | -1.2299 | 0 | m | 0.0708 | 0 | -1.23 | 0 | m | OK | 0.0708 | -1.23 | m | -7.7057e-10 | 8.0596e-10 | m | 0.001 | 7.0674e-05 | 0 | -0.0012299 | 0 | 0.0012319 | -1.5134 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | 8.0951e-08 | 0.0021467 | 9.9997e-08 | 0.0018435 | 50Hz3m3m | 1.8269e-07 | 0.0072902 | 2.5164e-07 | 0.0041084 | 9.275e-08 | 0.0072902 | 1.1758e-07 | 0.0021061 | 7.0766e-05 | 0.0072902 | -0.0012298 | 0.0021061 | 1.0001 | 0 | 0.0001 | 0 | 1.0001 | 0 | 8.3873e-05 | 0 | 1.0001 | 0 | 9.813e-05 | 0 | 0.001232 | -1.5133 | 7.08e-05 | 0 | -0.00123 | 0
10 | m | 5000 | 5 | kHz | 0.51272 | 0 | 8.1001 | 0 | m | 0.51313 | 0 | 8.1 | 0 | m | OK | 0.513 | 8.1 | m | 0.00012531 | -3.3402e-05 | m | 0.001 | 0.00051272 | 0 | 0.0081001 | 0 | 0.0081163 | 1.5076 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | 3.8348e-08 | 0.0077777 | 3.0329e-07 | 0.00081221 | 5kHz10m10m | -2.9561e-07 | 0.004558 | -1.122e-06 | 0.0059978 | -2.0575e-07 | 0.004558 | -7.3844e-07 | 0.0046024 | 0.00051251 | 0.004558 | 0.0080994 | 0.0046024 | 1 | 0 | 9.4207e-05 | 0 | 0.99997 | 0 | -0.00010453 | 0 | 0.99998 | 0 | -5.1051e-05 | 0 | 0.0081162 | 1.5075 | 0.00051313 | 0 | 0.0081 | 0
10 | m | 0.1 | 100 | mHz | 6.88 | 0 | 0.11071 | 0 | m | 6.88 | 0 | 0.111 | 0 | m | OK | 6.88 | 0.111 | m | 1.2024e-08 | 1.7202e-08 | m | 0.001 | 0.00688 | 0 | 0.00011071 | 0 | 0.0068809 | 0.01609 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | -2.3891e-07 | 0.0043614 | 3.0005e-07 | 0.0066178 | 100mHz10m10m | 6.6571e-07 | 0.00029851 | -3.3573e-08 | 0.0037274 | 2.6262e-07 | 0.00029851 | 1.1509e-07 | 0.0050153 | 0.0068803 | 0.00029851 | 0.00011082 | 0.0050153 | 0.99992 | 0 | 0.00010001 | 0 | 1.0001 | 0 | -3.3607e-06 | 0 | 1 | 0 | 4.27e-05 | 0 | 0.0068809 | 0.016132 | 0.00688 | 0 | 0.000111 | 0
10 | m | 0.1 | 100 | mHz | -5.7202 | 0 | -0.49267 | 0 | m | -5.72 | 0 | -0.493 | 0 | m | OK | -5.72 | -0.493 | m | 1.0223e-08 | -2.5545e-08 | m | 0.001 | -0.0057202 | 0 | -0.00049267 | 0 | 0.0057413 | -3.0557 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | -2.3891e-07 | 0.0043614 | 3.0005e-07 | 0.0066178 | 100mHz10m10m | 6.6571e-07 | 0.00029851 | -3.3573e-08 | 0.0037274 | 1.1535e-07 | 0.00029851 | 1.694e-07 | 0.0054859 | -0.00572 | 0.00029851 | -0.0004925 | 0.0054859 | 0.99992 | 0 | 0.00010001 | 0 | 1.0001 | 0 | -3.3607e-06 | 0 | 0.99998 | 0 | 5.9528e-05 | 0 | 0.0057412 | -3.0556 | -0.00572 | 0 | -0.000493 | 0
10 | m | 10 | 10 | Hz | -9.0399 | 0 | -1.2103 | 0 | m | -9.04 | 0 | -1.21 | 0 | m | OK | -9.04 | -1.21 | m | -2.0206e-07 | 2.4054e-07 | m | 0.001 | -0.0090399 | 0 | -0.0012103 | 0 | 0.0091205 | -3.0085 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | 1.4801e-07 | 0.0049286 | -4.6163e-07 | 0.0019795 | -0.0090397 | 0.0049286 | -0.0012108 | 0.0019795 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 1 | 0 | -3.736e-05 | 0 | 0.0091206 | -3.0085 | -0.00904 | 0 | -0.00121 | 0
10 | m | 0.02 | 20 | mHz | 3.1105 | 0 | 5.9296 | 0 | m | 3.11 | 0 | 5.93 | 0 | m | OK | 3.11 | 5.93 | m | -1.8688e-09 | -8.3494e-10 | m | 0.001 | 0.0031105 | 0 | 0.0059296 | 0 | 0.0066959 | 1.0877 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 1.8525e-07 | 0.0030642 | 6.0549e-07 | 0.0072215 | 0.0031107 | 0.0030642 | 0.0059302 | 0.0072215 | 1 | 0 | 0.0001 | 0 | 1 | 0 | 8.7863e-05 | 0 | 1 | 0 | 9.3592e-05 | 0 | 0.006696 | 1.0878 | 0.00311 | 0 | 0.00593 | 0
3 | m | 0.1 | 100 | mHz | 1.18 | 0 | 1.7099 | 0 | m | 1.18 | 0 | 1.71 | 0 | m | OK | 1.18 | 1.71 | m | -3.3078e-08 | 3.8601e-08 | m | 0.001 | 0.00118 | 0 | 0.0017099 | 0 | 0.0020775 | 0.96675 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 5.5717e-08 | 0.0071683 | 1e-07 | 0.0063429 | 100mHz3m3m | 5.8282e-08 | 0.0018787 | -1.3465e-07 | 0.0076313 | 5.7099e-08 | 0.0018787 | -2.6421e-08 | 0.0070371 | 0.0011801 | 0.0018787 | 0.0017099 | 0.0070371 | 1.0001 | 0 | 9.9996e-05 | 0 | 1 | 0 | -4.4888e-05 | 0 | 1 | 0 | 2.1937e-05 | 0 | 0.0020776 | 0.96677 | 0.00118 | 0 | 0.00171 | 0
10 | m | 10 | 10 | Hz | 2.6604 | 0 | 2.3299 | 0 | m | 2.66 | 0 | 2.33 | 0 | m | OK | 2.66 | 2.33 | m | 2.4715e-08 | -5.9152e-09 | m | 0.001 | 0.0026604 | 0 | 0.0023299 | 0 | 0.0035364 | 0.71927 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | -1.9042e-07 | 0.0049286 | 2.3324e-07 | 0.007378 | 0.0026602 | 0.0049286 | 0.0023302 | 0.007378 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 0.99993 | 0 | 8.7958e-05 | 0 | 0.0035362 | 0.71936 | 0.00266 | 0 | 0.00233 | 0
3 | m | 0.1 | 100 | mHz | -0.90602 | 0 | -0.81088 | 0 | m | -0.906 | 0 | -0.811 | 0 | m | OK | -0.906 | -0.811 | m | 7.0909e-10 | -1.6669e-09 | m | 0.001 | -0.00090602 | 0 | -0.00081088 | 0 | 0.0012159 | -2.4116 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 5.5717e-08 | 0.0071683 | 1e-07 | 0.0063429 | 100mHz3m3m | 5.8282e-08 | 0.0018787 | -1.3465e-07 | 0.0076313 | 5.5994e-08 | 0.0018787 | 7.467e-08 | 0.006482 | -0.00090597 | 0.0018787 | -0.00081081 | 0.006482 | 1.0001 | 0 | 9.9996e-05 | 0 | 1 | 0 | -4.4888e-05 | 0 | 1.0001 | 0 | 8.4356e-05 | 0 | 0.001216 | -2.4115 | -0.000906 | 0 | -0.000811 | 0
3 | m | 50 | 50 | Hz | 0.76479 | 0 | -1.65 | 0 | m | 0.765 | 0 | -1.65 | 0 | m | OK | 0.765 | -1.65 | m | -4.4562e-09 | 2.0828e-09 | m | 0.001 | 0.00076479 | 0 | -0.00165 | 0 | 0.0018186 | -1.1368 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | 8.0951e-08 | 0.0021467 | 9.9997e-08 | 0.0018435 | 50Hz3m3m | 1.8269e-07 | 0.0072902 | 2.5164e-07 | 0.0041084 | 1.2259e-07 | 0.0072902 | 1.6206e-07 | 0.0027705 | 0.00076491 | 0.0072902 | -0.0016498 | 0.0027705 | 1.0001 | 0 | 0.0001 | 0 | 1.0001 | 0 | 8.3873e-05 | 0 | 1.0001 | 0 | 9.3399e-05 | 0 | 0.0018187 | -1.1367 | 0.000765 | 0 | -0.00165 | 0
10 | m | 20 | 20 | Hz | -6.2302 | 0 | 7.529 | 0 | m | -6.23 | 0 | 7.53 | 0 | m | OK | -6.23 | 7.53 | m | -6.7064e-08 | -1.2095e-07 | m | 0.001 | -0.0062302 | 0 | 0.007529 | 0 | 0.0097725 | 2.2621 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | -3.5033e-08 | 0.003442 | 3e-07 | 0.0039013 | 20Hz10m10m | 6.3819e-07 | 0.0012733 | -8.6252e-07 | 0.00022773 | 6.1631e-07 | 0.0012733 | -8.2474e-07 | 0.00034711 | -0.0062296 | 0.0012733 | 0.0075282 | 0.00034711 | 0.99999 | 0 | 0.0001 | 0 | 1.0001 | 0 | -8.6269e-05 | 0 | 1.0001 | 0 | -8.0216e-05 | 0 | 0.0097731 | 2.262 | -0.00623 | 0 | 0.00753 | 0
10 | m | 2 | 2 | Hz | 2.2503 | 0 | 6.2095 | 0 | m | 2.25 | 0 | 6.21 | 0 | m | OK | 2.25 | 6.21 | m | 1.6064e-08 | -1.2598e-08 | m | 0.001 | 0.0022503 | 0 | 0.0062095 | 0 | 0.0066047 | 1.2231 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.9667e-07 | 0.0080028 | 2.9999e-07 | 0.0092189 | 2Hz10m10m | 4.1052e-07 | 0.0037012 | 3.5445e-07 | 0.0031449 | 3.068e-07 | 0.0037012 | 3.2803e-07 | 0.0060911 | 0.0022506 | 0.0037012 | 0.0062098 | 0.0060911 | 1.0001 | 0 | 0.0001 | 0 | 1 | 0 | 3.5443e-05 | 0 | 1.0001 | 0 | 6.6756e-05 | 0 | 0.006605 | 1.2232 | 0.00225 | 0 | 0.00621 | 0
100 | m | 10 | 10 | Hz | -11.0993 | 0 | 66.4999 | 0 | m | -11.1 | 0 | 66.5 | 0 | m | OK | -11.1 | 66.5 | m | -1.7356e-07 | 8.1829e-08 | m | 0.001 | -0.011099 | 0 | 0.0665 | 0 | 0.06742 | 1.7362 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | 3.7878e-07 | 0.0082107 | 1e-06 | 0.0030243 | 10Hz100m100m | -1.7212e-06 | 0.010095 | -3.9909e-06 | 0.0061855 | -9.6098e-07 | 0.010095 | -2.1842e-06 | 0.0050411 | -0.0111 | 0.010095 | 0.066498 | 0.0050411 | 1 | 0 | 9.9998e-05 | 0 | 0.99998 | 0 | -3.9911e-05 | 0 | 1 | 0 | 1.0736e-05 | 0 | 0.06742 | 1.7362 | -0.0111 | 0 | 0.0665 | 0
100 | m | 0.5 | 500 | mHz | -49.3987 | 0 | -83.2083 | 0 | m | -49.4 | 0 | -83.2 | 0 | m | OK | -49.4 | -83.2 | m | -7.4627e-07 | 3.4744e-07 | m | 0.001 | -0.049399 | 0 | -0.083208 | 0 | 0.096767 | -2.1066 | 4 | 10m | 100m | 0.01 | 0.1 | 500mHz100m10m | 500mHz100m100m | 114 | 132 | 500mHz100m10m | -8.2605e-07 | 0.00017609 | 1.0003e-06 | 0.0005888 | 500mHz100m100m | -6.6614e-06 | 0.0086963 | -6.1008e-06 | 0.00085803 | -6.4518e-06 | 0.0086963 | -5.8457e-06 | 0.00084836 | -0.049405 | 0.0086963 | -0.083214 | 0.00084836 | 0.99992 | 0 | 9.9998e-05 | 0 | 0.99993 | 0 | -6.0999e-05 | 0 | 0.99993 | 0 | -5.5215e-05 | 0 | 0.096761 | -2.1066 | -0.0494 | 0 | -0.0832 | 0
10 | m | 2000 | 2 | kHz | -1.9104 | 0 | -3.91 | 0 | m | -1.91 | 0 | -3.91 | 0 | m | OK | -1.91 | -3.91 | m | -4.0301e-06 | -3.8205e-07 | m | 0.001 | -0.0019104 | 0 | -0.00391 | 0 | 0.0043517 | -2.0253 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -1.1285e-07 | 0.0076693 | 3.0232e-07 | 0.0041829 | -0.0019105 | 0.0076693 | -0.0039097 | 0.0041829 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99996 | 0 | 8.5862e-05 | 0 | 0.0043516 | -2.0252 | -0.00191 | 0 | -0.00391 | 0
3 | m | 5000 | 5 | kHz | 0.83311 | 0 | 1.3899 | 0 | m | 0.83301 | 0 | 1.39 | 0 | m | OK | 0.833 | 1.39 | m | 1.12e-05 | 6.5949e-06 | m | 0.001 | 0.00083311 | 0 | 0.0013899 | 0 | 0.0016205 | 1.0308 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | -3.2349e-08 | 0.0038778 | 1.0472e-07 | 0.0061575 | 5kHz3m3m | 3.1454e-07 | 0.0080214 | 9.0519e-08 | 0.00038508 | 7.5271e-08 | 0.0080214 | 1.0031e-07 | 0.0043666 | 0.00083318 | 0.0080214 | 0.00139 | 0.0043666 | 0.99997 | 0 | 9.7633e-05 | 0 | 1.0001 | 0 | 2.7996e-05 | 0 | 1 | 0 | 7.6029e-05 | 0 | 0.0016205 | 1.0309 | 0.00083301 | 0 | 0.00139 | 0
10 | m | 5000 | 5 | kHz | -0.1183 | 0 | -3.0199 | 0 | m | -0.11802 | 0 | -3.02 | 0 | m | OK | -0.118 | -3.02 | m | -1.7486e-05 | 4.2109e-06 | m | 0.001 | -0.0001183 | 0 | -0.0030199 | 0 | 0.0030223 | -1.6099 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | 3.8348e-08 | 0.0077777 | 3.0329e-07 | 0.00081221 | 5kHz10m10m | -2.9561e-07 | 0.004558 | -1.122e-06 | 0.0059978 | 3.7286e-08 | 0.004558 | 2.9876e-07 | 0.00082871 | -0.00011826 | 0.004558 | -0.0030197 | 0.00082871 | 1 | 0 | 9.4207e-05 | 0 | 0.99997 | 0 | -0.00010453 | 0 | 1 | 0 | 9.3575e-05 | 0 | 0.0030223 | -1.6099 | -0.00011802 | 0 | -0.00302 | 0
10 | m | 2000 | 2 | kHz | -0.84273 | 0 | 3.1902 | 0 | m | -0.843 | 0 | 3.19 | 0 | m | OK | -0.843 | 3.19 | m | 1.6095e-06 | 1.6809e-06 | m | 0.001 | -0.00084273 | 0 | 0.0031902 | 0 | 0.0032996 | 1.8291 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -1.31e-07 | 0.0076693 | 3.0348e-07 | 0.0039134 | -0.00084286 | 0.0076693 | 0.0031905 | 0.0039134 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99996 | 0 | 9.644e-05 | 0 | 0.0032995 | 1.8292 | -0.000843 | 0 | 0.00319 | 0
10 | m | 2000 | 2 | kHz | 7.74 | 0 | -2.7804 | 0 | m | 7.74 | 0 | -2.78 | 0 | m | OK | 7.74 | -2.78 | m | 2.9291e-06 | -1.4158e-05 | m | 0.001 | 0.00774 | 0 | -0.0027804 | 0 | 0.0082242 | -0.34487 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -4.6047e-08 | 0.0076693 | 2.9806e-07 | 0.005175 | 0.0077399 | 0.0076693 | -0.0027801 | 0.005175 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99999 | 0 | 4.6928e-05 | 0 | 0.0082241 | -0.34482 | 0.00774 | 0 | -0.00278 | 0
1 | m | 100 | 100 | Hz | -0.080924 | 0 | 0.76896 | 0 | m | -0.080893 | 0 | 0.76901 | 0 | m | OK | -0.0809 | 0.769 | m | 7.1953e-06 | 1.0633e-05 | m | 0.001 | -8.0924e-05 | 0 | 0.00076896 | 0 | 0.00077321 | 1.6756 | 1 | 0m | 1m | 0 | 0.001 | 100Hz1m0m | 100Hz1m1m | 13 | 31 | 100Hz1m0m | 0 | 0.0047626 | 0 | 0.0076039 | 100Hz1m1m | 5.5951e-08 | 0.0047466 | -4.7099e-08 | 0.0084662 | 4.3262e-08 | 0.0047466 | -4.7099e-08 | 0.0082707 | -8.0881e-05 | 0.0047466 | 0.00076892 | 0.0082707 | 1.0001 | 0 | -4.7081e-05 | 0 | 1.0001 | 0 | -4.7081e-05 | 0 | 1.0001 | 0 | -4.7081e-05 | 0 | 0.00077325 | 1.6756 | -8.0893e-05 | 0 | 0.00076901 | 0
1 | m | 1000 | 1 | kHz | -0.54698 | 0 | 0.16799 | 0 | m | -0.54702 | 0 | 0.16801 | 0 | m | OK | -0.547 | 0.168 | m | -1.6891e-05 | 5.1487e-06 | m | 0.001 | -0.00054698 | 0 | 0.00016799 | 0 | 0.00057219 | 2.8436 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 0 | 0.0093653 | 0 | 0.0072173 | 1kHz1m1m | 7.1975e-08 | 0.0046398 | 1.7424e-10 | 0.0016346 | 4.1184e-08 | 0.0046398 | 1.7424e-10 | 0.0040229 | -0.00054694 | 0.0046398 | 0.00016799 | 0.0040229 | 1.0001 | 0 | 1.5054e-07 | 0 | 1.0001 | 0 | 1.5054e-07 | 0 | 1.0001 | 0 | 1.5054e-07 | 0 | 0.00057224 | 2.8436 | -0.00054702 | 0 | 0.00016801 | 0
10 | m | 0.02 | 20 | mHz | -2.6303 | 0 | -3.2697 | 0 | m | -2.63 | 0 | -3.27 | 0 | m | OK | -2.63 | -3.27 | m | -1.0654e-10 | -3.5089e-10 | m | 0.001 | -0.0026303 | 0 | -0.0032697 | 0 | 0.0041964 | -2.2482 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 5.9726e-08 | 0.0030642 | 3.9889e-07 | 0.006615 | -0.0026302 | 0.0030642 | -0.0032693 | 0.006615 | 1 | 0 | 0.0001 | 0 | 1 | 0 | 8.7863e-05 | 0 | 1 | 0 | 9.7926e-05 | 0 | 0.0041964 | -2.2481 | -0.00263 | 0 | -0.00327 | 0
1 | m | 50 | 50 | Hz | 0.168 | 0 | -0.188 | 0 | m | 0.16799 | 0 | -0.18801 | 0 | m | OK | 0.168 | -0.188 | m | -8.381e-06 | -8.4914e-06 | m | 0.001 | 0.000168 | 0 | -0.000188 | 0 | 0.00025213 | -0.8415 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | 0 | 0.00075911 | 0 | 0.0021036 | 50Hz1m1m | 3.9691e-09 | 0.00045661 | -6.315e-08 | 0.0078449 | 1.0007e-09 | 0.00045661 | -6.315e-08 | 0.0035511 | 0.000168 | 0.00045661 | -0.00018806 | 0.0035511 | 1 | 0 | -6.3149e-05 | 0 | 1 | 0 | -6.3149e-05 | 0 | 1 | 0 | -6.3149e-05 | 0 | 0.00025213 | -0.84157 | 0.00016799 | 0 | -0.00018801 | 0
10 | m | 2000 | 2 | kHz | -4.3398 | 0 | 4.5104 | 0 | m | -4.34 | 0 | 4.51 | 0 | m | OK | -4.34 | 4.51 | m | 2.1179e-06 | 8.1027e-06 | m | 0.001 | -0.0043398 | 0 | 0.0045104 | 0 | 0.0062592 | 2.3369 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -7.9944e-08 | 0.0076693 | 3.0022e-07 | 0.0046716 | -0.0043399 | 0.0076693 | 0.0045107 | 0.0046716 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99998 | 0 | 6.6685e-05 | 0 | 0.0062591 | 2.337 | -0.00434 | 0 | 0.00451 | 0
10 | m | 1000 | 1 | kHz | 7.12 | 0 | -3.1499 | 0 | m | 7.12 | 0 | -3.15 | 0 | m | OK | 7.12 | -3.15 | m | 3.5447e-06 | -1.2101e-05 | m | 0.001 | 0.00712 | 0 | -0.0031499 | 0 | 0.0077856 | -0.41651 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.7712e-07 | 0.0081916 | 3.0055e-07 | 0.0022801 | 1kHz10m10m | 5.1261e-07 | 0.0034662 | -7.0524e-07 | 0.0068801 | 2.6279e-07 | 0.0034662 | -3.8707e-07 | 0.0054249 | 0.0071203 | 0.0034662 | -0.0031503 | 0.0054249 | 0.99991 | 0 | 9.9528e-05 | 0 | 1.0001 | 0 | -7.0075e-05 | 0 | 1 | 0 | -1.6424e-05 | 0 | 0.0077857 | -0.41653 | 0.00712 | 0 | -0.00315 | 0
10 | m | 1 | 1 | Hz | 8.5204 | 0 | 3.8595 | 0 | m | 8.52 | 0 | 3.86 | 0 | m | OK | 8.52 | 3.86 | m | -7.7358e-10 | 6.398e-08 | m | 0.001 | 0.0085204 | 0 | 0.0038595 | 0 | 0.0093538 | 0.42533 | 3 | 3m | 10m | 0.003 | 0.01 | 1Hz10m3m | 1Hz10m10m | 79 | 97 | 1Hz10m3m | -1.7506e-09 | 0.0065386 | 2.9997e-07 | 0.0034264 | 1Hz10m10m | -1.8985e-07 | 0.0039514 | 5.762e-07 | 0.00089748 | -1.7248e-07 | 0.0039514 | 5.507e-07 | 0.0011309 | 0.0085202 | 0.0039514 | 0.0038601 | 0.0011309 | 1 | 0 | 0.0001 | 0 | 0.99998 | 0 | 5.7626e-05 | 0 | 0.99998 | 0 | 6.1538e-05 | 0 | 0.0093536 | 0.42539 | 0.00852 | 0 | 0.00386 | 0
10 | m | 100 | 100 | Hz | -3.3397 | 0 | 6.2999 | 0 | m | -3.34 | 0 | 6.3 | 0 | m | OK | -3.34 | 6.3 | m | 1.0317e-07 | 1.7211e-07 | m | 0.001 | -0.0033397 | 0 | 0.0062999 | 0 | 0.0071303 | 2.0582 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 3.7407e-07 | 0.0013885 | 6.0237e-08 | 0.0056795 | -0.0033393 | 0.0013885 | 0.0062999 | 0.0056795 | 1 | 0 | 9.9996e-05 | 0 | 1.0001 | 0 | -1.0637e-05 | 0 | 1 | 0 | 3.4717e-05 | 0 | 0.0071306 | 2.0583 | -0.00334 | 0 | 0.0063 | 0
10 | m | 500 | 500 | Hz | 5.8595 | 0 | 3.1297 | 0 | m | 5.86 | 0 | 3.13 | 0 | m | OK | 5.86 | 3.13 | m | 8.7742e-07 | -2.373e-06 | m | 0.001 | 0.0058595 | 0 | 0.0031297 | 0 | 0.006643 | 0.49057 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 2.7795e-07 | 0.0014136 | 3.0022e-07 | 0.002703 | 500Hz10m10m | 6.7384e-07 | 0.0094192 | -8.365e-07 | 0.00078689 | 4.8399e-07 | 0.0094192 | -2.9137e-07 | 0.0017058 | 0.00586 | 0.0094192 | 0.0031294 | 0.0017058 | 1.0001 | 0 | 9.9828e-05 | 0 | 1.0001 | 0 | -8.3468e-05 | 0 | 1.0001 | 0 | 4.4356e-06 | 0 | 0.0066435 | 0.49058 | 0.00586 | 0 | 0.00313 | 0
1 | m | 2 | 2 | Hz | 0.0383 | 0 | 0.129 | 0 | m | 0.038297 | 0 | 0.12901 | 0 | m | OK | 0.0383 | 0.129 | m | -3.1317e-06 | 5.5474e-06 | m | 0.001 | 3.83e-05 | 0 | 0.000129 | 0 | 0.00013456 | 1.2822 | 1 | 0m | 1m | 0 | 0.001 | 2Hz1m0m | 2Hz1m1m | 8 | 26 | 2Hz1m0m | 0 | 0.0060159 | 0 | 0.0056977 | 2Hz1m1m | 3.8011e-08 | 0.0078244 | 3.9332e-08 | 0.002428 | 5.115e-09 | 0.0078244 | 3.9332e-08 | 0.0052577 | 3.8306e-05 | 0.0078244 | 0.00012904 | 0.0052577 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 1 | 0 | 3.9336e-05 | 0 | 0.00013457 | 1.2822 | 3.8297e-05 | 0 | 0.00012901 | 0
3 | m | 200 | 200 | Hz | 2.6498 | 0 | 1.0499 | 0 | m | 2.65 | 0 | 1.05 | 0 | m | OK | 2.65 | 1.05 | m | 7.2826e-08 | -1.9324e-07 | m | 0.001 | 0.0026498 | 0 | 0.0010499 | 0 | 0.0028502 | 0.37724 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | 9.9186e-08 | 0.0011654 | 1.0001e-07 | 0.0083568 | 200Hz3m3m | 2.9261e-07 | 0.0029843 | 8.5449e-09 | 0.0015311 | 2.7812e-07 | 0.0029843 | 1.5398e-08 | 0.0020425 | 0.00265 | 0.0029843 | 0.0010499 | 0.0020425 | 1.0001 | 0 | 9.9974e-05 | 0 | 1.0001 | 0 | 2.8448e-06 | 0 | 1.0001 | 0 | 1.0122e-05 | 0 | 0.0028504 | 0.37725 | 0.00265 | 0 | 0.00105 | 0
10 | m | 2000 | 2 | kHz | -1.4696 | 0 | 6.2502 | 0 | m | -1.47 | 0 | 6.25 | 0 | m | OK | -1.47 | 6.25 | m | 6.2654e-06 | 6.1974e-06 | m | 0.001 | -0.0014696 | 0 | 0.0062502 | 0 | 0.0064207 | 1.8017 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | -7.7158e-08 | 0.0076693 | 3.0005e-07 | 0.004713 | -0.0014697 | 0.0076693 | 0.0062505 | 0.004713 | 0.99996 | 0 | 9.9453e-05 | 0 | 1 | 0 | 2.9074e-05 | 0 | 0.99998 | 0 | 6.5061e-05 | 0 | 0.0064205 | 1.8018 | -0.00147 | 0 | 0.00625 | 0
1 | m | 0.05 | 50 | mHz | -0.11801 | 0 | -0.51803 | 0 | m | -0.11799 | 0 | -0.51798 | 0 | m | OK | -0.118 | -0.518 | m | 9.0918e-06 | 2.2866e-05 | m | 0.001 | -0.00011801 | 0 | -0.00051803 | 0 | 0.0005313 | -1.7948 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.0049811 | 0 | 0.0050777 | 50mHz1m1m | -9.7658e-08 | 0.00076464 | 1.519e-08 | 0.008821 | -5.1885e-08 | 0.00076464 | 1.519e-08 | 0.0070665 | -0.00011806 | 0.00076464 | -0.00051801 | 0.0070665 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.9999 | 0 | 1.5204e-05 | 0 | 0.00053125 | -1.7948 | -0.00011799 | 0 | -0.00051798 | 0
3 | m | 0.02 | 20 | mHz | 0.38806 | 0 | -2.3799 | 0 | m | 0.388 | 0 | -2.38 | 0 | m | OK | 0.388 | -2.38 | m | -3.7519e-10 | -1.6617e-10 | m | 0.001 | 0.00038806 | 0 | -0.0023799 | 0 | 0.0024114 | -1.4092 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | 3.9555e-08 | 0.0033415 | 9.9995e-08 | 0.0060072 | 20mHz3m3m | 6.8657e-08 | 0.0033998 | -2.544e-07 | 0.0013912 | 6.0092e-08 | 0.0033998 | -1.5009e-07 | 0.0027498 | 0.00038812 | 0.0033998 | -0.0023801 | 0.0027498 | 1 | 0 | 0.0001 | 0 | 1 | 0 | -8.4806e-05 | 0 | 1 | 0 | -3.0412e-05 | 0 | 0.0024114 | -1.4092 | 0.000388 | 0 | -0.00238 | 0
3 | m | 0.02 | 20 | mHz | 1.3101 | 0 | -2.1999 | 0 | m | 1.31 | 0 | -2.2 | 0 | m | OK | 1.31 | -2.2 | m | -1.5473e-09 | -8.7681e-10 | m | 0.001 | 0.0013101 | 0 | -0.0021999 | 0 | 0.0025604 | -1.0337 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | 3.9555e-08 | 0.0033415 | 9.9995e-08 | 0.0060072 | 20mHz3m3m | 6.8657e-08 | 0.0033998 | -2.544e-07 | 0.0013912 | 6.2261e-08 | 0.0033998 | -1.7651e-07 | 0.0024058 | 0.0013101 | 0.0033998 | -0.0022001 | 0.0024058 | 1 | 0 | 0.0001 | 0 | 1 | 0 | -8.4806e-05 | 0 | 1 | 0 | -4.4187e-05 | 0 | 0.0025605 | -1.0337 | 0.00131 | 0 | -0.0022 | 0
1 | m | 0.5 | 500 | mHz | 0.304 | 0 | 0.33999 | 0 | m | 0.304 | 0 | 0.34001 | 0 | m | OK | 0.304 | 0.34 | m | -4.6437e-06 | 1.1936e-05 | m | 0.001 | 0.000304 | 0 | 0.00033999 | 0 | 0.00045608 | 0.84122 | 1 | 0m | 1m | 0 | 0.001 | 500mHz1m0m | 500mHz1m1m | 6 | 24 | 500mHz1m0m | 0 | 0.00065229 | 0 | 0.0096122 | 500mHz1m1m | 2.3394e-08 | 0.0025222 | 4.6022e-08 | 0.0036456 | 1.0669e-08 | 0.0025222 | 4.6022e-08 | 0.0068909 | 0.00030401 | 0.0025222 | 0.00034004 | 0.0068909 | 1 | 0 | 4.6024e-05 | 0 | 1 | 0 | 4.6024e-05 | 0 | 1 | 0 | 4.6024e-05 | 0 | 0.00045609 | 0.84127 | 0.000304 | 0 | 0.00034001 | 0
10 | m | 0.05 | 50 | mHz | 1.8899 | 0 | -6.9303 | 0 | m | 1.89 | 0 | -6.93 | 0 | m | OK | 1.89 | -6.93 | m | -6.2539e-08 | -4.4498e-08 | m | 0.001 | 0.0018899 | 0 | -0.0069303 | 0 | 0.0071834 | -1.3046 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -1.966e-07 | 0.0016505 | 3e-07 | 0.00092106 | 50mHz10m10m | -2.6057e-07 | 0.0029905 | -1.9484e-07 | 0.0036795 | -2.3483e-07 | 0.0029905 | 4.2701e-09 | 0.0025696 | 0.0018896 | 0.0029905 | -0.0069303 | 0.0025696 | 0.99993 | 0 | 9.9994e-05 | 0 | 0.99997 | 0 | -1.9476e-05 | 0 | 0.99996 | 0 | 2.8595e-05 | 0 | 0.0071831 | -1.3045 | 0.00189 | 0 | -0.00693 | 0
1 | m | 20 | 20 | Hz | -0.88896 | 0 | 0.102 | 0 | m | -0.889 | 0 | 0.102 | 0 | m | OK | -0.889 | 0.102 | m | -4.7042e-06 | -1.7734e-07 | m | 0.001 | -0.00088896 | 0 | 0.000102 | 0 | 0.00089479 | 3.0273 | 1 | 0m | 1m | 0 | 0.001 | 20Hz1m0m | 20Hz1m1m | 11 | 29 | 20Hz1m0m | 0 | 0.0036421 | 0 | 0.008204 | 20Hz1m1m | 4.9405e-08 | 0.0045511 | 7.5608e-09 | 5.0763e-05 | 4.4207e-08 | 0.0045511 | 7.5608e-09 | 0.00090854 | -0.00088892 | 0.0045511 | 0.00010201 | 0.00090854 | 1 | 0 | 7.5635e-06 | 0 | 1 | 0 | 7.5635e-06 | 0 | 1 | 0 | 7.5635e-06 | 0 | 0.00089484 | 3.0274 | -0.000889 | 0 | 0.000102 | 0
1 | m | 100 | 100 | Hz | -0.014024 | 0 | 0.72597 | 0 | m | -0.013991 | 0 | 0.72601 | 0 | m | OK | -0.014 | 0.726 | m | 9.1579e-06 | 1.1315e-05 | m | 0.001 | -1.4024e-05 | 0 | 0.00072597 | 0 | 0.00072611 | 1.5901 | 1 | 0m | 1m | 0 | 0.001 | 100Hz1m0m | 100Hz1m1m | 13 | 31 | 100Hz1m0m | 0 | 0.0047626 | 0 | 0.0076039 | 100Hz1m1m | 5.5951e-08 | 0.0047466 | -4.7099e-08 | 0.0084662 | 4.0626e-08 | 0.0047466 | -4.7099e-08 | 0.00823 | -1.3984e-05 | 0.0047466 | 0.00072592 | 0.00823 | 1.0001 | 0 | -4.7081e-05 | 0 | 1.0001 | 0 | -4.7081e-05 | 0 | 1.0001 | 0 | -4.7081e-05 | 0 | 0.00072615 | 1.5901 | -1.3991e-05 | 0 | 0.00072601 | 0
10 | m | 100 | 100 | Hz | -4.1601 | 0 | -4.9997 | 0 | m | -4.16 | 0 | -5 | 0 | m | OK | -4.16 | -5 | m | -1.5903e-07 | 2.3298e-08 | m | 0.001 | -0.0041601 | 0 | -0.0049997 | 0 | 0.0065041 | -2.2648 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 3.1769e-07 | 0.0013885 | 9.6594e-08 | 0.005624 | -0.0041598 | 0.0013885 | -0.0049996 | 0.005624 | 1 | 0 | 9.9996e-05 | 0 | 1.0001 | 0 | -1.0637e-05 | 0 | 1 | 0 | 4.4615e-05 | 0 | 0.0065043 | -2.2647 | -0.00416 | 0 | -0.005 | 0
10 | m | 10 | 10 | Hz | -3.4404 | 0 | -2.0799 | 0 | m | -3.44 | 0 | -2.08 | 0 | m | OK | -3.44 | -2.08 | m | -3.6481e-08 | 1.4935e-08 | m | 0.001 | -0.0034404 | 0 | -0.0020799 | 0 | 0.0040202 | -2.5978 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | -1.611e-07 | 0.0049286 | 1.7304e-07 | 0.0069103 | -0.0034405 | 0.0049286 | -0.0020797 | 0.0069103 | 0.99993 | 0 | 9.9996e-05 | 0 | 1 | 0 | -5.7097e-05 | 0 | 0.99994 | 0 | 7.7101e-05 | 0 | 0.00402 | -2.5978 | -0.00344 | 0 | -0.00208 | 0
3 | m | 0.01 | 10 | mHz | 1.5599 | 0 | -1.2301 | 0 | m | 1.56 | 0 | -1.23 | 0 | m | OK | 1.56 | -1.23 | m | 3.7715e-09 | 9.2086e-10 | m | 0.001 | 0.0015599 | 0 | -0.0012301 | 0 | 0.0019866 | -0.66774 | 2 | 1m | 3m | 0.001 | 0.003 | 10mHz3m1m | 10mHz3m3m | 37 | 55 | 10mHz3m1m | -6.9054e-08 | 0.0078687 | 9.9999e-08 | 0.0068365 | 10mHz3m3m | 2.1001e-07 | 0.002327 | 1.2308e-07 | 0.0047587 | 6.8606e-08 | 0.002327 | 1.1139e-07 | 0.0058115 | 0.00156 | 0.002327 | -0.00123 | 0.0058115 | 0.99993 | 0 | 0.0001 | 0 | 1.0001 | 0 | 4.1025e-05 | 0 | 1 | 0 | 7.0909e-05 | 0 | 0.0019866 | -0.66767 | 0.00156 | 0 | -0.00123 | 0
3 | m | 20 | 20 | Hz | -2.0502 | 0 | -1.9701 | 0 | m | -2.05 | 0 | -1.97 | 0 | m | OK | -2.05 | -1.97 | m | -1.1206e-08 | -2.3834e-08 | m | 0.001 | -0.0020502 | 0 | -0.0019701 | 0 | 0.0028433 | -2.3761 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | 7.7175e-08 | 0.0035201 | 9.9992e-08 | 0.00318 | 20Hz3m3m | -2.3188e-07 | 0.0072946 | 7.8386e-08 | 0.0075897 | -2.0767e-07 | 0.0072946 | 8.0079e-08 | 0.0072443 | -0.0020504 | 0.0072946 | -0.00197 | 0.0072443 | 1.0001 | 0 | 9.9998e-05 | 0 | 0.99992 | 0 | 2.6133e-05 | 0 | 0.99993 | 0 | 3.1919e-05 | 0 | 0.0028431 | -2.3761 | -0.00205 | 0 | -0.00197 | 0
10 | m | 0.01 | 10 | mHz | -2.9296 | 0 | 3.6302 | 0 | m | -2.93 | 0 | 3.63 | 0 | m | OK | -2.93 | 3.63 | m | -3.721e-10 | -1.0896e-08 | m | 0.001 | -0.0029296 | 0 | 0.0036302 | 0 | 0.0046648 | 2.2498 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | 1.8571e-08 | 0.0029007 | 3e-07 | 0.0072065 | 10mHz10m10m | 9.212e-07 | 0.0099497 | 5.767e-07 | 0.0069975 | 2.3325e-07 | 0.0099497 | 3.6581e-07 | 0.0071568 | -0.0029294 | 0.0099497 | 0.0036305 | 0.0071568 | 1 | 0 | 0.0001 | 0 | 1.0001 | 0 | 5.7696e-05 | 0 | 1 | 0 | 8.9939e-05 | 0 | 0.004665 | 2.2499 | -0.00293 | 0 | 0.00363 | 0
10 | m | 0.02 | 20 | mHz | -5.9097 | 0 | 2.0605 | 0 | m | -5.91 | 0 | 2.06 | 0 | m | OK | -5.91 | 2.06 | m | 6.1838e-10 | -1.2882e-09 | m | 0.001 | -0.0059097 | 0 | 0.0020605 | 0 | 0.0062586 | 2.8061 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 1.6329e-07 | 0.0030642 | 5.6934e-07 | 0.0071154 | -0.0059095 | 0.0030642 | 0.0020611 | 0.0071154 | 1 | 0 | 0.0001 | 0 | 1 | 0 | 8.7863e-05 | 0 | 1 | 0 | 9.435e-05 | 0 | 0.0062587 | 2.8062 | -0.00591 | 0 | 0.00206 | 0
3 | m | 10 | 10 | Hz | -1.1399 | 0 | 1.04 | 0 | m | -1.14 | 0 | 1.04 | 0 | m | OK | -1.14 | 1.04 | m | -7.34e-10 | -8.4241e-10 | m | 0.001 | -0.0011399 | 0 | 0.00104 | 0 | 0.0015431 | 2.402 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 3.2828e-08 | 0.0025658 | 9.9991e-08 | 0.0091467 | 10Hz3m3m | 1.117e-07 | 0.007815 | -1.6997e-07 | 0.0066531 | 5.4244e-08 | 0.007815 | 2.6688e-08 | 0.0084696 | -0.0011398 | 0.007815 | 0.0010401 | 0.0084696 | 1 | 0 | 0.00010001 | 0 | 1 | 0 | -5.6662e-05 | 0 | 1 | 0 | 5.7466e-05 | 0 | 0.0015431 | 2.402 | -0.00114 | 0 | 0.00104 | 0
1 | m | 2000 | 2 | kHz | 0.30798 | 0 | 0.36898 | 0 | m | 0.30802 | 0 | 0.36902 | 0 | m | OK | 0.308 | 0.369 | m | 1.7816e-05 | 1.7793e-05 | m | 0.001 | 0.00030798 | 0 | 0.00036898 | 0 | 0.00048063 | 0.87526 | 1 | 0m | 1m | 0 | 0.001 | 2kHz1m0m | 2kHz1m1m | 17 | 35 | 2kHz1m0m | 0 | 0.0082269 | 0 | 0.0083167 | 2kHz1m1m | 1.0061e-07 | 0.0054645 | -9.0811e-09 | 0.0030182 | 4.8355e-08 | 0.0054645 | -9.0811e-09 | 0.0057701 | 0.00030803 | 0.0054645 | 0.00036898 | 0.0057701 | 1.0001 | 0 | -8.9663e-06 | 0 | 1.0001 | 0 | -8.9663e-06 | 0 | 1.0001 | 0 | -8.9663e-06 | 0 | 0.00048068 | 0.87525 | 0.00030802 | 0 | 0.00036902 | 0
3 | m | 0.2 | 200 | mHz | -0.73004 | 0 | -1.3998 | 0 | m | -0.73 | 0 | -1.4 | 0 | m | OK | -0.73 | -1.4 | m | 4.8014e-09 | -2.7275e-09 | m | 0.001 | -0.00073004 | 0 | -0.0013998 | 0 | 0.0015788 | -2.0515 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | 8.0669e-08 | 0.0027077 | 9.999e-08 | 0.002112 | 200mHz3m3m | 2.3152e-07 | 0.0088393 | 4.0846e-09 | 0.0047677 | 1.2432e-07 | 0.0088393 | 7.2237e-08 | 0.0028805 | -0.00072992 | 0.0088393 | -0.0013998 | 0.0028805 | 1.0001 | 0 | 9.9999e-05 | 0 | 1.0001 | 0 | 1.3572e-06 | 0 | 1.0001 | 0 | 7.1454e-05 | 0 | 0.0015789 | -2.0514 | -0.00073 | 0 | -0.0014 | 0
10 | m | 1000 | 1 | kHz | -6.2701 | 0 | 2.8501 | 0 | m | -6.27 | 0 | 2.85 | 0 | m | OK | -6.27 | 2.85 | m | -2.6832e-06 | 9.4937e-06 | m | 0.001 | -0.0062701 | 0 | 0.0028501 | 0 | 0.0068874 | 2.715 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.7712e-07 | 0.0081916 | 3.0055e-07 | 0.0022801 | 1kHz10m10m | 5.1261e-07 | 0.0034662 | -7.0524e-07 | 0.0068801 | 1.6146e-07 | 0.0034662 | -2.5801e-07 | 0.0048347 | -0.0062699 | 0.0034662 | 0.0028498 | 0.0048347 | 0.99991 | 0 | 9.9528e-05 | 0 | 1.0001 | 0 | -7.0075e-05 | 0 | 0.99999 | 0 | 5.3393e-06 | 0 | 0.0068873 | 2.715 | -0.00627 | 0 | 0.00285 | 0
10 | m | 200 | 200 | Hz | 1.6797 | 0 | -5.26 | 0 | m | 1.68 | 0 | -5.26 | 0 | m | OK | 1.68 | -5.26 | m | -2.5059e-07 | -4.1953e-07 | m | 0.001 | 0.0016797 | 0 | -0.00526 | 0 | 0.0055217 | -1.2617 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -7.8893e-08 | 0.0054354 | 3.0006e-07 | 0.0020919 | 200Hz10m10m | 9.884e-07 | 0.0064562 | -5.1026e-07 | 0.0043182 | 3.0559e-07 | 0.0064562 | 8.1526e-09 | 0.0028939 | 0.00168 | 0.0064562 | -0.00526 | 0.0028939 | 0.99997 | 0 | 9.9966e-05 | 0 | 1.0001 | 0 | -5.1006e-05 | 0 | 1 | 0 | 4.558e-05 | 0 | 0.0055218 | -1.2616 | 0.00168 | 0 | -0.00526 | 0
10 | m | 0.02 | 20 | mHz | -3.33 | 0 | -0.53167 | 0 | m | -3.33 | 0 | -0.532 | 0 | m | OK | -3.33 | -0.532 | m | -4.0181e-10 | 7.6461e-13 | m | 0.001 | -0.00333 | 0 | -0.00053167 | 0 | 0.0033722 | -2.9833 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 1.8337e-08 | 0.0030642 | 3.3077e-07 | 0.006415 | -0.00333 | 0.0030642 | -0.00053134 | 0.006415 | 1 | 0 | 0.0001 | 0 | 1 | 0 | 8.7863e-05 | 0 | 1 | 0 | 9.9355e-05 | 0 | 0.0033722 | -2.9832 | -0.00333 | 0 | -0.000532 | 0
3 | m | 0.02 | 20 | mHz | 0.86122 | 0 | -2.8699 | 0 | m | 0.861 | 0 | -2.87 | 0 | m | OK | 0.861 | -2.87 | m | -6.6129e-09 | -1.3517e-09 | m | 0.001 | 0.00086122 | 0 | -0.0028699 | 0 | 0.0029963 | -1.2793 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | 3.9555e-08 | 0.0033415 | 9.9995e-08 | 0.0060072 | 20mHz3m3m | 6.8657e-08 | 0.0033998 | -2.544e-07 | 0.0013912 | 6.8604e-08 | 0.0033998 | -2.5374e-07 | 0.0013998 | 0.00086129 | 0.0033998 | -0.0028701 | 0.0013998 | 1 | 0 | 0.0001 | 0 | 1 | 0 | -8.4806e-05 | 0 | 1 | 0 | -8.4464e-05 | 0 | 0.0029964 | -1.2793 | 0.000861 | 0 | -0.00287 | 0
10 | m | 50 | 50 | Hz | 3.9805 | 0 | -8 | 0 | m | 3.98 | 0 | -8 | 0 | m | OK | 3.98 | -8 | m | -4.2789e-07 | -1.252e-07 | m | 0.001 | 0.0039805 | 0 | -0.008 | 0 | 0.0089356 | -1.1091 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | 3.0977e-09 | 0.0015823 | 3.0001e-07 | 0.00091715 | 50Hz10m10m | -3.1031e-07 | 0.006568 | -7.6553e-07 | 0.0090646 | -2.6265e-07 | 0.006568 | -6.035e-07 | 0.0078257 | 0.0039802 | 0.006568 | -0.0080006 | 0.0078257 | 1 | 0 | 9.9987e-05 | 0 | 0.99997 | 0 | -7.6535e-05 | 0 | 0.99997 | 0 | -4.9693e-05 | 0 | 0.0089353 | -1.1092 | 0.00398 | 0 | -0.008 | 0

## Ranges
Range list
Range |  | Range Z | Min ¦Z¦ |  | Min Z | Max ¦Z¦ |  | Max Z
Ohm |  | Ohm | Ohm |  | Ohm | Ohm |  | Ohm
1 | m | 0.001 | 0 | m | 0 | 1 | m | 0.001
3 | m | 0.003 | 1 | m | 0.001 | 3 | m | 0.003
10 | m | 0.01 | 3 | m | 0.003 | 10 | m | 0.01
100 | m | 0.1 | 10 | m | 0.01 | 100 | m | 0.1

## Cal Data
Calibration data: measured values of reference standards.
[9 empty] | Measured calibration Z | [11 empty] | Reference Z (auto search from Ref Z sheet) | [7 empty] | Absolute correction
Range |  | Range Z | f | f |  | Nominal Z |  | Nom Z | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | phi | raf tag | ref ID | Rs | U(Rs) | Xs | U(Xs) | Zr | phir | tag | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | cor(gain) | cor(phi)
Ohm |  | Ohm | Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | rad | - | - | Ohm | Ohm | Ohm | Ohm | Ohm | rad | - | Ohm | Ohm | Ohm | Ohm | - | rad
1 | m | 0.001 | 0.01 | 10 | mHz | 0 | m | 0 | 0 | 0.0026794 | 0 | 0.0020396 | m | 0.001 | 0 | 2.6794e-06 | 0 | 2.0396e-06 | 0 | #DIV/0! | 10mHz0m | 1 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 10mHz1m0m | 0 | 0.0053587 | 0 | 0.0040791 | 1 | -2.3802e-05
1 | m | 0.001 | 0.02 | 20 | mHz | 0 | m | 0 | 0 | 0.00055261 | 0 | 0.0028746 | m | 0.001 | 0 | 5.5261e-07 | 0 | 2.8746e-06 | 0 | #DIV/0! | 20mHz0m | 2 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 20mHz1m0m | 0 | 0.0011052 | 0 | 0.0057493 | 0.9999 | 2.6633e-05
1 | m | 0.001 | 0.05 | 50 | mHz | 0 | m | 0 | 0 | 0.0024906 | 0 | 0.0025389 | m | 0.001 | 0 | 2.4906e-06 | 0 | 2.5389e-06 | 0 | #DIV/0! | 50mHz0m | 3 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 50mHz1m0m | 0 | 0.0049811 | 0 | 0.0050777 | 0.9999 | 1.5204e-05
1 | m | 0.001 | 0.1 | 100 | mHz | 0 | m | 0 | 0 | 0.0043544 | 0 | 0.002004 | m | 0.001 | 0 | 4.3544e-06 | 0 | 2.004e-06 | 0 | #DIV/0! | 100mHz0m | 4 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 100mHz1m0m | 0 | 0.0087088 | 0 | 0.004008 | 1 | 5.6635e-05
1 | m | 0.001 | 0.2 | 200 | mHz | 0 | m | 0 | 0 | 0.0036805 | 0 | 0.0021207 | m | 0.001 | 0 | 3.6805e-06 | 0 | 2.1207e-06 | 0 | #DIV/0! | 200mHz0m | 5 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 200mHz1m0m | 0 | 0.0073611 | 0 | 0.0042415 | 1 | 9.0079e-05
1 | m | 0.001 | 0.5 | 500 | mHz | 0 | m | 0 | 0 | 0.00032614 | 0 | 0.0048061 | m | 0.001 | 0 | 3.2614e-07 | 0 | 4.8061e-06 | 0 | #DIV/0! | 500mHz0m | 6 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 500mHz1m0m | 0 | 0.00065229 | 0 | 0.0096122 | 1 | 4.6024e-05
1 | m | 0.001 | 1 | 1 | Hz | 0 | m | 0 | 0 | 0.0034869 | 0 | 0.0043614 | m | 0.001 | 0 | 3.4869e-06 | 0 | 4.3614e-06 | 0 | #DIV/0! | 1Hz0m | 7 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 1Hz1m0m | 0 | 0.0069739 | 0 | 0.0087229 | 1 | -3.6545e-05
1 | m | 0.001 | 2 | 2 | Hz | 0 | m | 0 | 0 | 0.0030079 | 0 | 0.0028489 | m | 0.001 | 0 | 3.0079e-06 | 0 | 2.8489e-06 | 0 | #DIV/0! | 2Hz0m | 8 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 2Hz1m0m | 0 | 0.0060159 | 0 | 0.0056977 | 1 | 3.9336e-05
1 | m | 0.001 | 5 | 5 | Hz | 0 | m | 0 | 0 | 0.0021737 | 0 | 0.0028956 | m | 0.001 | 0 | 2.1737e-06 | 0 | 2.8956e-06 | 0 | #DIV/0! | 5Hz0m | 9 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 5Hz1m0m | 0 | 0.0043474 | 0 | 0.0057912 | 0.99999 | 5.5339e-05
1 | m | 0.001 | 10 | 10 | Hz | 0 | m | 0 | 0 | 0.0027504 | 0 | 0.00053447 | m | 0.001 | 0 | 2.7504e-06 | 0 | 5.3447e-07 | 0 | #DIV/0! | 10Hz0m | 10 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 10Hz1m0m | 0 | 0.0055008 | 0 | 0.0010689 | 0.99997 | 9.3993e-05
1 | m | 0.001 | 20 | 20 | Hz | 0 | m | 0 | 0 | 0.001821 | 0 | 0.004102 | m | 0.001 | 0 | 1.821e-06 | 0 | 4.102e-06 | 0 | #DIV/0! | 20Hz0m | 11 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 20Hz1m0m | 0 | 0.0036421 | 0 | 0.008204 | 1 | 7.5635e-06
1 | m | 0.001 | 50 | 50 | Hz | 0 | m | 0 | 0 | 0.00037955 | 0 | 0.0010518 | m | 0.001 | 0 | 3.7955e-07 | 0 | 1.0518e-06 | 0 | #DIV/0! | 50Hz0m | 12 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 50Hz1m0m | 0 | 0.00075911 | 0 | 0.0021036 | 1 | -6.3149e-05
1 | m | 0.001 | 100 | 100 | Hz | 0 | m | 0 | 0 | 0.0023813 | 0 | 0.0038019 | m | 0.001 | 0 | 2.3813e-06 | 0 | 3.8019e-06 | 0 | #DIV/0! | 100Hz0m | 13 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 100Hz1m0m | 0 | 0.0047626 | 0 | 0.0076039 | 1.0001 | -4.7081e-05
1 | m | 0.001 | 200 | 200 | Hz | 0 | m | 0 | 0 | 0.00099181 | 0 | 0.0013264 | m | 0.001 | 0 | 9.9181e-07 | 0 | 1.3264e-06 | 0 | #DIV/0! | 200Hz0m | 14 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 200Hz1m0m | 0 | 0.0019836 | 0 | 0.0026527 | 0.99999 | 2.0172e-05
1 | m | 0.001 | 500 | 500 | Hz | 0 | m | 0 | 0 | 0.0012807 | 0 | 0.0039369 | m | 0.001 | 0 | 1.2807e-06 | 0 | 3.9369e-06 | 0 | #DIV/0! | 500Hz0m | 15 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 500Hz1m0m | 0 | 0.0025615 | 0 | 0.0078738 | 1 | 8.0235e-05
1 | m | 0.001 | 1000 | 1 | kHz | 0 | m | 0 | 0 | 0.0046827 | 0 | 0.0036087 | m | 0.001 | 0 | 4.6827e-06 | 0 | 3.6087e-06 | 0 | #DIV/0! | 1kHz0m | 16 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 1kHz1m0m | 0 | 0.0093653 | 0 | 0.0072173 | 1.0001 | 1.5054e-07
1 | m | 0.001 | 2000 | 2 | kHz | 0 | m | 0 | 0 | 0.0041134 | 0 | 0.0041583 | m | 0.001 | 0 | 4.1134e-06 | 0 | 4.1583e-06 | 0 | #DIV/0! | 2kHz0m | 17 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 2kHz1m0m | 0 | 0.0082269 | 0 | 0.0083167 | 1.0001 | -8.9663e-06
1 | m | 0.001 | 5000 | 5 | kHz | 0 | m | 0 | 0 | 0.0037723 | 0 | 0.0022948 | m | 0.001 | 0 | 3.7723e-06 | 0 | 2.2948e-06 | 0 | #DIV/0! | 5kHz0m | 18 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | 5kHz1m0m | 0 | 0.0075446 | 0 | 0.0045896 | 0.99998 | 1.5079e-05
1 | m | 0.001 | 0.01 | 10 | mHz | 1 | m | 0.001 | 0.99994 | 0.0010156 | 9.1594e-05 | 0.0018761 | m | 0.001 | 0.00099994 | 1.0156e-06 | 9.1594e-08 | 1.8761e-06 | 0.00099994 | 9.1599e-05 | 10mHz1m | 19 | 0.00099995 | 1e-08 | 6.7794e-08 | 5e-08 | 0.00099995 | 6.7797e-05 | 10mHz1m1m | 9.0865e-09 | 0.0020313 | -2.38e-08 | 0.0037522 | 1 | -2.3802e-05
1 | m | 0.001 | 0.02 | 20 | mHz | 1 | m | 0.001 | 1 | 0.0042586 | -5.3338e-05 | 0.0031283 | m | 0.001 | 0.001 | 4.2586e-06 | -5.3338e-08 | 3.1283e-06 | 0.001 | -5.3336e-05 | 20mHz1m | 20 | 0.00099995 | 1e-08 | -2.6701e-08 | 5e-08 | 0.00099995 | -2.6703e-05 | 20mHz1m1m | -9.7918e-08 | 0.0085172 | 2.6637e-08 | 0.0062565 | 0.9999 | 2.6633e-05
1 | m | 0.001 | 0.05 | 50 | mHz | 1 | m | 0.001 | 1.0001 | 0.00038232 | 0.00014737 | 0.0044105 | m | 0.001 | 0.0010001 | 3.8232e-07 | 1.4737e-07 | 4.4105e-06 | 0.0010001 | 0.00014735 | 50mHz1m | 21 | 0.001 | 1e-08 | 1.6256e-07 | 5e-08 | 0.001 | 0.00016255 | 50mHz1m1m | -9.7658e-08 | 0.00076464 | 1.519e-08 | 0.008821 | 0.9999 | 1.5204e-05
1 | m | 0.001 | 0.1 | 100 | mHz | 1 | m | 0.001 | 1.0001 | 0.002685 | 1.0567e-05 | 0.0043335 | m | 0.001 | 0.0010001 | 2.685e-06 | 1.0567e-08 | 4.3335e-06 | 0.0010001 | 1.0566e-05 | 100mHz1m | 22 | 0.0010001 | 1e-08 | 6.7205e-08 | 5e-08 | 0.0010001 | 6.7202e-05 | 100mHz1m1m | -4.6628e-09 | 0.00537 | 5.6638e-08 | 0.008667 | 1 | 5.6635e-05
1 | m | 0.001 | 0.2 | 200 | mHz | 1 | m | 0.001 | 1 | 0.0022228 | -0.00012114 | 0.0049958 | m | 0.001 | 0.001 | 2.2228e-06 | -1.2114e-07 | 4.9958e-06 | 0.001 | -0.00012114 | 200mHz1m | 23 | 0.001 | 1e-08 | -3.1063e-08 | 5e-08 | 0.001 | -3.1062e-05 | 200mHz1m1m | 1.4791e-08 | 0.0044457 | 9.0079e-08 | 0.0099916 | 1 | 9.0079e-05
1 | m | 0.001 | 0.5 | 500 | mHz | 1 | m | 0.001 | 0.99997 | 0.0012611 | -7.9394e-05 | 0.0018228 | m | 0.001 | 0.00099997 | 1.2611e-06 | -7.9394e-08 | 1.8228e-06 | 0.00099997 | -7.9396e-05 | 500mHz1m | 24 | 0.00099999 | 1e-08 | -3.3372e-08 | 5e-08 | 0.00099999 | -3.3372e-05 | 500mHz1m1m | 2.3394e-08 | 0.0025222 | 4.6022e-08 | 0.0036456 | 1 | 4.6024e-05
1 | m | 0.001 | 1 | 1 | Hz | 1 | m | 0.001 | 1 | 0.0021979 | -1.3184e-05 | 0.0019064 | m | 0.001 | 0.001 | 2.1979e-06 | -1.3184e-08 | 1.9064e-06 | 0.001 | -1.3184e-05 | 1Hz1m | 25 | 0.001 | 1e-08 | -4.9729e-08 | 5e-08 | 0.001 | -4.9728e-05 | 1Hz1m1m | 2.4541e-08 | 0.0043959 | -3.6546e-08 | 0.0038128 | 1 | -3.6545e-05
1 | m | 0.001 | 2 | 2 | Hz | 1 | m | 0.001 | 0.99994 | 0.0039122 | -7.0189e-05 | 0.001214 | m | 0.001 | 0.00099994 | 3.9122e-06 | -7.0189e-08 | 1.214e-06 | 0.00099994 | -7.0193e-05 | 2Hz1m | 26 | 0.00099998 | 1e-08 | -3.0857e-08 | 5e-08 | 0.00099998 | -3.0857e-05 | 2Hz1m1m | 3.8011e-08 | 0.0078244 | 3.9332e-08 | 0.002428 | 1 | 3.9336e-05
1 | m | 0.001 | 5 | 5 | Hz | 1 | m | 0.001 | 1 | 0.0030529 | -0.00013099 | 0.0022757 | m | 0.001 | 0.001 | 3.0529e-06 | -1.3099e-07 | 2.2757e-06 | 0.001 | -0.00013099 | 5Hz1m | 27 | 0.001 | 1e-08 | -7.5652e-08 | 5e-08 | 0.001 | -7.5651e-05 | 5Hz1m1m | -9.0123e-09 | 0.0061058 | 5.5341e-08 | 0.0045514 | 0.99999 | 5.5339e-05
1 | m | 0.001 | 10 | 10 | Hz | 1 | m | 0.001 | 0.99995 | 0.0024436 | -0.00021545 | 0.0020256 | m | 0.001 | 0.00099995 | 2.4436e-06 | -2.1545e-07 | 2.0256e-06 | 0.00099995 | -0.00021546 | 10Hz1m | 28 | 0.00099993 | 1e-08 | -1.2146e-07 | 5e-08 | 0.00099993 | -0.00012147 | 10Hz1m1m | -2.7396e-08 | 0.0048871 | 9.3992e-08 | 0.0040511 | 0.99997 | 9.3993e-05
1 | m | 0.001 | 20 | 20 | Hz | 1 | m | 0.001 | 1 | 0.0022755 | -6.2705e-05 | 2.5381e-05 | m | 0.001 | 0.001 | 2.2755e-06 | -6.2705e-08 | 2.5381e-08 | 0.001 | -6.2705e-05 | 20Hz1m | 29 | 0.0010001 | 1e-08 | -5.5145e-08 | 5e-08 | 0.0010001 | -5.5141e-05 | 20Hz1m1m | 4.9405e-08 | 0.0045511 | 7.5608e-09 | 5.0763e-05 | 1 | 7.5635e-06
1 | m | 0.001 | 50 | 50 | Hz | 1 | m | 0.001 | 1 | 0.0002283 | 0.00011237 | 0.0039224 | m | 0.001 | 0.001 | 2.283e-07 | 1.1237e-07 | 3.9224e-06 | 0.001 | 0.00011237 | 50Hz1m | 30 | 0.001 | 1e-08 | 4.9221e-08 | 5e-08 | 0.001 | 4.9221e-05 | 50Hz1m1m | 3.9691e-09 | 0.00045661 | -6.315e-08 | 0.0078449 | 1 | -6.3149e-05
1 | m | 0.001 | 100 | 100 | Hz | 1 | m | 0.001 | 1.0004 | 0.0023733 | 3.8828e-05 | 0.0042331 | m | 0.001 | 0.0010004 | 2.3733e-06 | 3.8828e-08 | 4.2331e-06 | 0.0010004 | 3.8813e-05 | 100Hz1m | 31 | 0.0010004 | 1e-08 | -8.2709e-09 | 5e-08 | 0.0010004 | -8.2674e-06 | 100Hz1m1m | 5.5951e-08 | 0.0047466 | -4.7099e-08 | 0.0084662 | 1.0001 | -4.7081e-05
1 | m | 0.001 | 200 | 200 | Hz | 1 | m | 0.001 | 1.0005 | 0.0048292 | -5.8484e-05 | 0.0043026 | m | 0.001 | 0.0010005 | 4.8292e-06 | -5.8484e-08 | 4.3026e-06 | 0.0010005 | -5.8452e-05 | 200Hz1m | 32 | 0.0010005 | 1e-08 | -3.8301e-08 | 5e-08 | 0.0010005 | -3.828e-05 | 200Hz1m1m | -8.2665e-09 | 0.0096585 | 2.0183e-08 | 0.0086052 | 0.99999 | 2.0172e-05
1 | m | 0.001 | 500 | 500 | Hz | 1 | m | 0.001 | 1.0022 | 0.00011421 | -0.00010304 | 0.0016591 | m | 0.001 | 0.0010022 | 1.1421e-07 | -1.0304e-07 | 1.6591e-06 | 0.0010022 | -0.00010281 | 500Hz1m | 33 | 0.0010023 | 1.5811e-08 | -2.2623e-08 | 5e-08 | 0.0010023 | -2.2572e-05 | 500Hz1m1m | 1.1707e-08 | 0.00022843 | 8.0415e-08 | 0.0033182 | 1 | 8.0235e-05
1 | m | 0.001 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0064 | 0.0023199 | 0.00031789 | 0.00081729 | m | 0.001 | 0.0010064 | 2.3199e-06 | 3.1789e-07 | 8.1729e-07 | 0.0010064 | 0.00031588 | 1kHz1m | 34 | 0.0010064 | 4.4721e-08 | 3.1807e-07 | 1e-07 | 0.0010064 | 0.00031603 | 1kHz1m1m | 7.1975e-08 | 0.0046398 | 1.7424e-10 | 0.0016346 | 1.0001 | 1.5054e-07
1 | m | 0.001 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0182 | 0.0027323 | 0.00050326 | 0.0015091 | m | 0.001 | 0.0010182 | 2.7323e-06 | 5.0326e-07 | 1.5091e-06 | 0.0010182 | 0.00049424 | 2kHz1m | 35 | 0.0010183 | 1.2649e-07 | 4.9418e-07 | 2e-07 | 0.0010183 | 0.00048528 | 2kHz1m1m | 1.0061e-07 | 0.0054645 | -9.0811e-09 | 0.0030182 | 1.0001 | -8.9663e-06
1 | m | 0.001 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.073 | 0.0011933 | 0.001359 | 0.0047902 | m | 0.001 | 0.001073 | 1.1933e-06 | 1.359e-06 | 4.7902e-06 | 0.001073 | 0.0012666 | 5kHz1m | 36 | 0.0010729 | 5e-07 | 1.3752e-06 | 5e-07 | 0.0010729 | 0.0012817 | 5kHz1m1m | -1.9193e-08 | 0.0023865 | 1.6155e-08 | 0.0095804 | 0.99998 | 1.5079e-05
3 | m | 0.003 | 0.01 | 10 | mHz | 1 | m | 0.001 | 1 | 0.0039344 | -3.2205e-05 | 0.0034182 | m | 0.001 | 0.001 | 3.9344e-06 | -3.2205e-08 | 3.4182e-06 | 0.001 | -3.2204e-05 | 10mHz1m | 19 | 0.00099995 | 1e-08 | 6.7794e-08 | 5e-08 | 0.00099995 | 6.7797e-05 | 10mHz3m1m | -6.9054e-08 | 0.0078687 | 9.9999e-08 | 0.0068365 | 0.99993 | 0.0001
3 | m | 0.003 | 0.02 | 20 | mHz | 1 | m | 0.001 | 0.99991 | 0.0016708 | -0.0001267 | 0.0030036 | m | 0.001 | 0.00099991 | 1.6708e-06 | -1.267e-07 | 3.0036e-06 | 0.00099991 | -0.00012671 | 20mHz1m | 20 | 0.00099995 | 1e-08 | -2.6701e-08 | 5e-08 | 0.00099995 | -2.6703e-05 | 20mHz3m1m | 3.9555e-08 | 0.0033415 | 9.9995e-08 | 0.0060072 | 1 | 0.0001
3 | m | 0.003 | 0.05 | 50 | mHz | 1 | m | 0.001 | 0.99997 | 0.0037259 | 6.2552e-05 | 0.001424 | m | 0.001 | 0.00099997 | 3.7259e-06 | 6.2552e-08 | 1.424e-06 | 0.00099997 | 6.2553e-05 | 50mHz1m | 21 | 0.001 | 1e-08 | 1.6256e-07 | 5e-08 | 0.001 | 0.00016255 | 50mHz3m1m | 4.9e-08 | 0.0074519 | 1e-07 | 0.0028481 | 1 | 9.9998e-05
3 | m | 0.003 | 0.1 | 100 | mHz | 1 | m | 0.001 | 1 | 0.0035841 | -3.2794e-05 | 0.0031715 | m | 0.001 | 0.001 | 3.5841e-06 | -3.2794e-08 | 3.1715e-06 | 0.001 | -3.2794e-05 | 100mHz1m | 22 | 0.0010001 | 1e-08 | 6.7205e-08 | 5e-08 | 0.0010001 | 6.7202e-05 | 100mHz3m1m | 5.5717e-08 | 0.0071683 | 1e-07 | 0.0063429 | 1.0001 | 9.9996e-05
3 | m | 0.003 | 0.2 | 200 | mHz | 1 | m | 0.001 | 0.99993 | 0.0013539 | -0.00013105 | 0.001056 | m | 0.001 | 0.00099993 | 1.3539e-06 | -1.3105e-07 | 1.056e-06 | 0.00099993 | -0.00013106 | 200mHz1m | 23 | 0.001 | 1e-08 | -3.1063e-08 | 5e-08 | 0.001 | -3.1062e-05 | 200mHz3m1m | 8.0669e-08 | 0.0027077 | 9.999e-08 | 0.002112 | 1.0001 | 9.9999e-05
3 | m | 0.003 | 0.5 | 500 | mHz | 1 | m | 0.001 | 0.99999 | 0.0040005 | -0.00013337 | 0.0045264 | m | 0.001 | 0.00099999 | 4.0005e-06 | -1.3337e-07 | 4.5264e-06 | 0.00099999 | -0.00013337 | 500mHz1m | 24 | 0.00099999 | 1e-08 | -3.3372e-08 | 5e-08 | 0.00099999 | -3.3372e-05 | 500mHz3m1m | 8.5627e-09 | 0.008001 | 9.9999e-08 | 0.0090527 | 1 | 0.0001
3 | m | 0.003 | 1 | 1 | Hz | 1 | m | 0.001 | 0.99993 | 0.0043435 | -0.00014972 | 0.00045009 | m | 0.001 | 0.00099993 | 4.3435e-06 | -1.4972e-07 | 4.5009e-07 | 0.00099993 | -0.00014973 | 1Hz1m | 25 | 0.001 | 1e-08 | -4.9729e-08 | 5e-08 | 0.001 | -4.9728e-05 | 1Hz3m1m | 9.1737e-08 | 0.008687 | 9.9987e-08 | 0.00090018 | 1.0001 | 9.9999e-05
3 | m | 0.003 | 2 | 2 | Hz | 1 | m | 0.001 | 1 | 0.00063728 | -0.00013086 | 0.0025747 | m | 0.001 | 0.001 | 6.3728e-07 | -1.3086e-07 | 2.5747e-06 | 0.001 | -0.00013086 | 2Hz1m | 26 | 0.00099998 | 1e-08 | -3.0857e-08 | 5e-08 | 0.00099998 | -3.0857e-05 | 2Hz3m1m | -4.7398e-08 | 0.0012746 | 1.0001e-07 | 0.0051494 | 0.99995 | 0.0001
3 | m | 0.003 | 5 | 5 | Hz | 1 | m | 0.001 | 0.99992 | 0.0022881 | -0.00017564 | 0.0014865 | m | 0.001 | 0.00099992 | 2.2881e-06 | -1.7564e-07 | 1.4865e-06 | 0.00099992 | -0.00017565 | 5Hz1m | 27 | 0.001 | 1e-08 | -7.5652e-08 | 5e-08 | 0.001 | -7.5651e-05 | 5Hz3m1m | 8.8834e-08 | 0.0045761 | 9.9985e-08 | 0.002973 | 1.0001 | 0.0001
3 | m | 0.003 | 10 | 10 | Hz | 1 | m | 0.001 | 0.99989 | 0.0012829 | -0.00022145 | 0.0045733 | m | 0.001 | 0.00099989 | 1.2829e-06 | -2.2145e-07 | 4.5733e-06 | 0.00099989 | -0.00022148 | 10Hz1m | 28 | 0.00099993 | 1e-08 | -1.2146e-07 | 5e-08 | 0.00099993 | -0.00012147 | 10Hz3m1m | 3.2828e-08 | 0.0025658 | 9.9991e-08 | 0.0091467 | 1 | 0.00010001
3 | m | 0.003 | 20 | 20 | Hz | 1 | m | 0.001 | 0.99998 | 0.00176 | -0.00015514 | 0.00159 | m | 0.001 | 0.00099998 | 1.76e-06 | -1.5514e-07 | 1.59e-06 | 0.00099998 | -0.00015514 | 20Hz1m | 29 | 0.0010001 | 1e-08 | -5.5145e-08 | 5e-08 | 0.0010001 | -5.5141e-05 | 20Hz3m1m | 7.7175e-08 | 0.0035201 | 9.9992e-08 | 0.00318 | 1.0001 | 9.9998e-05
3 | m | 0.003 | 50 | 50 | Hz | 1 | m | 0.001 | 0.99994 | 0.0010734 | -5.0776e-05 | 0.00092173 | m | 0.001 | 0.00099994 | 1.0734e-06 | -5.0776e-08 | 9.2173e-07 | 0.00099994 | -5.0779e-05 | 50Hz1m | 30 | 0.001 | 1e-08 | 4.9221e-08 | 5e-08 | 0.001 | 4.9221e-05 | 50Hz3m1m | 8.0951e-08 | 0.0021467 | 9.9997e-08 | 0.0018435 | 1.0001 | 0.0001
3 | m | 0.003 | 100 | 100 | Hz | 1 | m | 0.001 | 1.0004 | 0.0047603 | -0.0001083 | 0.0048999 | m | 0.001 | 0.0010004 | 4.7603e-06 | -1.083e-07 | 4.8999e-06 | 0.0010004 | -0.00010826 | 100Hz1m | 31 | 0.0010004 | 1e-08 | -8.2709e-09 | 5e-08 | 0.0010004 | -8.2674e-06 | 100Hz3m1m | 6.7713e-08 | 0.0095207 | 1.0003e-07 | 0.0097997 | 1.0001 | 9.9997e-05
3 | m | 0.003 | 200 | 200 | Hz | 1 | m | 0.001 | 1.0004 | 0.00058269 | -0.00013831 | 0.0041784 | m | 0.001 | 0.0010004 | 5.8269e-07 | -1.3831e-07 | 4.1784e-06 | 0.0010004 | -0.00013825 | 200Hz1m | 32 | 0.0010005 | 1e-08 | -3.8301e-08 | 5e-08 | 0.0010005 | -3.828e-05 | 200Hz3m1m | 9.9186e-08 | 0.0011654 | 1.0001e-07 | 0.0083568 | 1.0001 | 9.9974e-05
3 | m | 0.003 | 500 | 500 | Hz | 1 | m | 0.001 | 1.0023 | 0.0023126 | -0.00012268 | 0.0035637 | m | 0.001 | 0.0010023 | 2.3126e-06 | -1.2268e-07 | 3.5637e-06 | 0.0010023 | -0.00012239 | 500Hz1m | 33 | 0.0010023 | 1.5811e-08 | -2.2623e-08 | 5e-08 | 0.0010023 | -2.2572e-05 | 500Hz3m1m | -9.5708e-08 | 0.0046252 | 1.0006e-07 | 0.0071273 | 0.9999 | 9.982e-05
3 | m | 0.003 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0065 | 0.0032233 | 0.00021753 | 0.001035 | m | 0.001 | 0.0010065 | 3.2233e-06 | 2.1753e-07 | 1.035e-06 | 0.0010065 | 0.00021613 | 1kHz1m | 34 | 0.0010064 | 4.4721e-08 | 3.1807e-07 | 1e-07 | 0.0010064 | 0.00031603 | 1kHz3m1m | -4.0985e-08 | 0.0064465 | 1.0054e-07 | 0.00207 | 0.99996 | 9.9906e-05
3 | m | 0.003 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0184 | 0.0038509 | 0.00039344 | 0.00028994 | m | 0.001 | 0.0010184 | 3.8509e-06 | 3.9344e-07 | 2.8994e-07 | 0.0010184 | 0.00038632 | 2kHz1m | 35 | 0.0010183 | 1.2649e-07 | 4.9418e-07 | 2e-07 | 0.0010183 | 0.00048528 | 2kHz3m1m | -7.0423e-08 | 0.0077018 | 1.0074e-07 | 0.00057988 | 0.99993 | 9.8952e-05
3 | m | 0.003 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.073 | 0.0019389 | 0.0012705 | 0.0030787 | m | 0.001 | 0.001073 | 1.9389e-06 | 1.2705e-06 | 3.0787e-06 | 0.001073 | 0.0011841 | 5kHz1m | 36 | 0.0010729 | 5e-07 | 1.3752e-06 | 5e-07 | 0.0010729 | 0.0012817 | 5kHz3m1m | -3.2349e-08 | 0.0038778 | 1.0472e-07 | 0.0061575 | 0.99997 | 9.7633e-05
3 | m | 0.003 | 0.01 | 10 | mHz | 3 | m | 0.003 | 2.9997 | 0.0011635 | 0.00015277 | 0.0023793 | m | 0.001 | 0.0029997 | 1.1635e-06 | 1.5277e-07 | 2.3793e-06 | 0.0029997 | 5.0927e-05 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 2.7585e-07 | 1.5e-07 | 0.0029999 | 9.1951e-05 | 10mHz3m3m | 2.1001e-07 | 0.002327 | 1.2308e-07 | 0.0047587 | 1.0001 | 4.1025e-05
3 | m | 0.003 | 0.02 | 20 | mHz | 3 | m | 0.003 | 3 | 0.0016999 | 0.0010039 | 0.00069562 | m | 0.001 | 0.003 | 1.6999e-06 | 1.0039e-06 | 6.9562e-07 | 0.003 | 0.00033463 | 20mHz3m | 38 | 0.003 | 3e-08 | 7.4948e-07 | 1.5e-07 | 0.003 | 0.00024982 | 20mHz3m3m | 6.8657e-08 | 0.0033998 | -2.544e-07 | 0.0013912 | 1 | -8.4806e-05
3 | m | 0.003 | 0.05 | 50 | mHz | 3 | m | 0.003 | 3.0002 | 0.00027546 | 0.00075017 | 0.0005635 | m | 0.001 | 0.0030002 | 2.7546e-07 | 7.5017e-07 | 5.635e-07 | 0.0030002 | 0.00025004 | 50mHz3m | 39 | 0.0030004 | 3e-08 | 6.3046e-07 | 1.5e-07 | 0.0030004 | 0.00021013 | 50mHz3m3m | 1.4521e-07 | 0.00055092 | -1.1971e-07 | 0.001127 | 1 | -3.9909e-05
3 | m | 0.003 | 0.1 | 100 | mHz | 3 | m | 0.003 | 2.9993 | 0.00093933 | -0.0007061 | 0.0038157 | m | 0.001 | 0.0029993 | 9.3933e-07 | -7.061e-07 | 3.8157e-06 | 0.0029993 | -0.00023542 | 100mHz3m | 40 | 0.0029993 | 3e-08 | -8.4074e-07 | 1.5e-07 | 0.0029993 | -0.00028031 | 100mHz3m3m | 5.8282e-08 | 0.0018787 | -1.3465e-07 | 0.0076313 | 1 | -4.4888e-05
3 | m | 0.003 | 0.2 | 200 | mHz | 3 | m | 0.003 | 2.9995 | 0.0044196 | 0.00017353 | 0.0023839 | m | 0.001 | 0.0029995 | 4.4196e-06 | 1.7353e-07 | 2.3839e-06 | 0.0029995 | 5.7855e-05 | 200mHz3m | 41 | 0.0029997 | 3e-08 | 1.7762e-07 | 1.5e-07 | 0.0029997 | 5.9212e-05 | 200mHz3m3m | 2.3152e-07 | 0.0088393 | 4.0846e-09 | 0.0047677 | 1.0001 | 1.3572e-06
3 | m | 0.003 | 0.5 | 500 | mHz | 3 | m | 0.003 | 3.0001 | 0.0026169 | 6.381e-05 | 0.0027781 | m | 0.001 | 0.0030001 | 2.6169e-06 | 6.381e-08 | 2.7781e-06 | 0.0030001 | 2.1269e-05 | 500mHz3m | 42 | 0.0029999 | 3e-08 | -1.5149e-07 | 1.5e-07 | 0.0029999 | -5.0499e-05 | 500mHz3m3m | -1.7973e-07 | 0.0052338 | -2.1531e-07 | 0.0055561 | 0.99994 | -7.1769e-05
3 | m | 0.003 | 1 | 1 | Hz | 3 | m | 0.003 | 2.9995 | 6.0651e-05 | -6.0791e-05 | 0.0024472 | m | 0.001 | 0.0029995 | 6.0651e-08 | -6.0791e-08 | 2.4472e-06 | 0.0029995 | -2.0267e-05 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -1.3225e-07 | 1.5e-07 | 0.0029996 | -4.4088e-05 | 1Hz3m3m | 1.324e-07 | 0.0001213 | -7.1455e-08 | 0.0048944 | 1 | -2.382e-05
3 | m | 0.003 | 2 | 2 | Hz | 3 | m | 0.003 | 2.9999 | 0.002871 | 0.00046232 | 0.0043566 | m | 0.001 | 0.0029999 | 2.871e-06 | 4.6232e-07 | 4.3566e-06 | 0.0029999 | 0.00015411 | 2Hz3m | 44 | 0.003 | 3e-08 | 2.0325e-07 | 1.5e-07 | 0.003 | 6.7752e-05 | 2Hz3m3m | 1.3089e-07 | 0.005742 | -2.5906e-07 | 0.0087132 | 1 | -8.6361e-05
3 | m | 0.003 | 5 | 5 | Hz | 3 | m | 0.003 | 2.9999 | 0.0038288 | -0.00056605 | 0.0033143 | m | 0.001 | 0.0029999 | 3.8288e-06 | -5.6605e-07 | 3.3143e-06 | 0.0029999 | -0.00018869 | 5Hz3m | 45 | 0.0030001 | 3e-08 | -4.5252e-07 | 1.5e-07 | 0.0030001 | -0.00015083 | 5Hz3m3m | 2.0952e-07 | 0.0076576 | 1.1353e-07 | 0.0066287 | 1.0001 | 3.7855e-05
3 | m | 0.003 | 10 | 10 | Hz | 3 | m | 0.003 | 3.0001 | 0.0039075 | 0.0006437 | 0.0033266 | m | 0.001 | 0.0030001 | 3.9075e-06 | 6.437e-07 | 3.3266e-06 | 0.0030001 | 0.00021456 | 10Hz3m | 46 | 0.0030002 | 3e-08 | 4.7373e-07 | 1.5e-07 | 0.0030002 | 0.0001579 | 10Hz3m3m | 1.117e-07 | 0.007815 | -1.6997e-07 | 0.0066531 | 1 | -5.6662e-05
3 | m | 0.003 | 20 | 20 | Hz | 3 | m | 0.003 | 3.0001 | 0.0036473 | 0.0001163 | 0.0037949 | m | 0.001 | 0.0030001 | 3.6473e-06 | 1.163e-07 | 3.7949e-06 | 0.0030001 | 3.8765e-05 | 20Hz3m | 47 | 0.0029999 | 3e-08 | 1.9468e-07 | 1.5e-07 | 0.0029999 | 6.4897e-05 | 20Hz3m3m | -2.3188e-07 | 0.0072946 | 7.8386e-08 | 0.0075897 | 0.99992 | 2.6133e-05
3 | m | 0.003 | 50 | 50 | Hz | 3 | m | 0.003 | 3.0003 | 0.0036451 | -0.00031334 | 0.0020542 | m | 0.001 | 0.0030003 | 3.6451e-06 | -3.1334e-07 | 2.0542e-06 | 0.0030003 | -0.00010443 | 50Hz3m | 48 | 0.0030005 | 3e-08 | -6.1695e-08 | 1.5e-07 | 0.0030005 | -2.0562e-05 | 50Hz3m3m | 1.8269e-07 | 0.0072902 | 2.5164e-07 | 0.0041084 | 1.0001 | 8.3873e-05
3 | m | 0.003 | 100 | 100 | Hz | 3 | m | 0.003 | 3 | 0.0027316 | -7.9294e-06 | 0.00052581 | m | 0.001 | 0.003 | 2.7316e-06 | -7.9294e-09 | 5.2581e-07 | 0.003 | -2.6432e-06 | 100Hz3m | 49 | 0.0030003 | 3e-08 | 2.4904e-07 | 1.5e-07 | 0.0030003 | 8.3006e-05 | 100Hz3m3m | 2.9949e-07 | 0.0054632 | 2.5697e-07 | 0.0010516 | 1.0001 | 8.565e-05
3 | m | 0.003 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0013 | 0.0014921 | 6.1394e-05 | 0.00076557 | m | 0.001 | 0.0030013 | 1.4921e-06 | 6.1394e-08 | 7.6557e-07 | 0.0030013 | 2.0456e-05 | 200Hz3m | 50 | 0.0030016 | 3e-08 | 6.9938e-08 | 1.5e-07 | 0.0030016 | 2.3301e-05 | 200Hz3m3m | 2.9261e-07 | 0.0029843 | 8.5449e-09 | 0.0015311 | 1.0001 | 2.8448e-06
3 | m | 0.003 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0066 | 0.00050468 | 0.0012912 | 0.00040547 | m | 0.001 | 0.0030066 | 5.0468e-07 | 1.2912e-06 | 4.0547e-07 | 0.0030066 | 0.00042944 | 500Hz3m | 51 | 0.0030066 | 4.7434e-08 | 1.1111e-06 | 1.5e-07 | 0.0030066 | 0.00036956 | 500Hz3m3m | -3.9688e-08 | 0.0010094 | -1.8007e-07 | 0.00081095 | 0.99999 | -5.9886e-05
3 | m | 0.003 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0198 | 0.001345 | -6.8942e-05 | 0.00013366 | m | 0.001 | 0.0030198 | 1.345e-06 | -6.8942e-08 | 1.3366e-07 | 0.0030198 | -2.283e-05 | 1kHz3m | 52 | 0.0030196 | 1.3416e-07 | 1.4353e-07 | 3e-07 | 0.0030196 | 4.7533e-05 | 1kHz3m3m | -1.4394e-07 | 0.0026901 | 2.1247e-07 | 0.00026733 | 0.99995 | 7.0363e-05
3 | m | 0.003 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0553 | 0.0030218 | 0.0016649 | 0.00033438 | m | 0.001 | 0.0030553 | 3.0218e-06 | 1.6649e-06 | 3.3438e-07 | 0.0030553 | 0.00054491 | 2kHz3m | 53 | 0.0030554 | 3.7947e-07 | 1.6117e-06 | 6e-07 | 0.0030554 | 0.00052749 | 2kHz3m3m | 1.0274e-07 | 0.0060437 | -5.3164e-08 | 0.00066877 | 1 | -1.7418e-05
3 | m | 0.003 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.2186 | 0.0040107 | 0.0040981 | 0.00019254 | m | 0.001 | 0.0032186 | 4.0107e-06 | 4.0981e-06 | 1.9254e-07 | 0.0032186 | 0.0012733 | 5kHz3m | 54 | 0.0032189 | 1.5e-06 | 4.1887e-06 | 1.5e-06 | 0.0032189 | 0.0013013 | 5kHz3m3m | 3.1454e-07 | 0.0080214 | 9.0519e-08 | 0.00038508 | 1.0001 | 2.7996e-05
10 | m | 0.01 | 0.01 | 10 | mHz | 3 | m | 0.003 | 2.9999 | 0.0014503 | -2.4146e-05 | 0.0036032 | m | 0.001 | 0.0029999 | 1.4503e-06 | -2.4146e-08 | 3.6032e-06 | 0.0029999 | -8.0489e-06 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 2.7585e-07 | 1.5e-07 | 0.0029999 | 9.1951e-05 | 10mHz10m3m | 1.8571e-08 | 0.0029007 | 3e-07 | 0.0072065 | 1 | 0.0001
10 | m | 0.01 | 0.02 | 20 | mHz | 3 | m | 0.003 | 3 | 0.004017 | 0.00044947 | 0.0031623 | m | 0.001 | 0.003 | 4.017e-06 | 4.4947e-07 | 3.1623e-06 | 0.003 | 0.00014982 | 20mHz3m | 38 | 0.003 | 3e-08 | 7.4948e-07 | 1.5e-07 | 0.003 | 0.00024982 | 20mHz10m3m | -3.5594e-10 | 0.0080341 | 3e-07 | 0.0063246 | 1 | 0.0001
10 | m | 0.01 | 0.05 | 50 | mHz | 3 | m | 0.003 | 3.0006 | 0.00082523 | 0.00033046 | 0.00046053 | m | 0.001 | 0.0030006 | 8.2523e-07 | 3.3046e-07 | 4.6053e-07 | 0.0030006 | 0.00011013 | 50mHz3m | 39 | 0.0030004 | 3e-08 | 6.3046e-07 | 1.5e-07 | 0.0030004 | 0.00021013 | 50mHz10m3m | -1.966e-07 | 0.0016505 | 3e-07 | 0.00092106 | 0.99993 | 9.9994e-05
10 | m | 0.01 | 0.1 | 100 | mHz | 3 | m | 0.003 | 2.9996 | 0.0021807 | -0.0011408 | 0.0033089 | m | 0.001 | 0.0029996 | 2.1807e-06 | -1.1408e-06 | 3.3089e-06 | 0.0029996 | -0.00038032 | 100mHz3m | 40 | 0.0029993 | 3e-08 | -8.4074e-07 | 1.5e-07 | 0.0029993 | -0.00028031 | 100mHz10m3m | -2.3891e-07 | 0.0043614 | 3.0005e-07 | 0.0066178 | 0.99992 | 0.00010001
10 | m | 0.01 | 0.2 | 200 | mHz | 3 | m | 0.003 | 2.9998 | 0.00053505 | -0.00012237 | 0.00027078 | m | 0.001 | 0.0029998 | 5.3505e-07 | -1.2237e-07 | 2.7078e-07 | 0.0029998 | -4.0792e-05 | 200mHz3m | 41 | 0.0029997 | 3e-08 | 1.7762e-07 | 1.5e-07 | 0.0029997 | 5.9212e-05 | 200mHz10m3m | -6.5943e-08 | 0.0010701 | 2.9999e-07 | 0.00054156 | 0.99998 | 0.0001
10 | m | 0.01 | 0.5 | 500 | mHz | 3 | m | 0.003 | 3.0002 | 0.0034779 | -0.00045153 | 0.0029062 | m | 0.001 | 0.0030002 | 3.4779e-06 | -4.5153e-07 | 2.9062e-06 | 0.0030002 | -0.0001505 | 500mHz3m | 42 | 0.0029999 | 3e-08 | -1.5149e-07 | 1.5e-07 | 0.0029999 | -5.0499e-05 | 500mHz10m3m | -2.333e-07 | 0.0069558 | 3.0003e-07 | 0.0058123 | 0.99992 | 0.0001
10 | m | 0.01 | 1 | 1 | Hz | 3 | m | 0.003 | 2.9996 | 0.0032693 | -0.00043222 | 0.0017132 | m | 0.001 | 0.0029996 | 3.2693e-06 | -4.3222e-07 | 1.7132e-06 | 0.0029996 | -0.00014409 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -1.3225e-07 | 1.5e-07 | 0.0029996 | -4.4088e-05 | 1Hz10m3m | -1.7506e-09 | 0.0065386 | 2.9997e-07 | 0.0034264 | 1 | 0.0001
10 | m | 0.01 | 2 | 2 | Hz | 3 | m | 0.003 | 2.9998 | 0.0040014 | -9.6738e-05 | 0.0046094 | m | 0.001 | 0.0029998 | 4.0014e-06 | -9.6738e-08 | 4.6094e-06 | 0.0029998 | -3.2248e-05 | 2Hz3m | 44 | 0.003 | 3e-08 | 2.0325e-07 | 1.5e-07 | 0.003 | 6.7752e-05 | 2Hz10m3m | 1.9667e-07 | 0.0080028 | 2.9999e-07 | 0.0092189 | 1.0001 | 0.0001
10 | m | 0.01 | 5 | 5 | Hz | 3 | m | 0.003 | 2.9999 | 0.0025238 | -0.00075247 | 0.0046712 | m | 0.001 | 0.0029999 | 2.5238e-06 | -7.5247e-07 | 4.6712e-06 | 0.0029999 | -0.00025083 | 5Hz3m | 45 | 0.0030001 | 3e-08 | -4.5252e-07 | 1.5e-07 | 0.0030001 | -0.00015083 | 5Hz10m3m | 2.3034e-07 | 0.0050475 | 2.9995e-07 | 0.0093424 | 1.0001 | 9.9999e-05
10 | m | 0.01 | 10 | 10 | Hz | 3 | m | 0.003 | 3.0004 | 0.00042945 | 0.00017374 | 0.0039483 | m | 0.001 | 0.0030004 | 4.2945e-07 | 1.7374e-07 | 3.9483e-06 | 0.0030004 | 5.7906e-05 | 10Hz3m | 46 | 0.0030002 | 3e-08 | 4.7373e-07 | 1.5e-07 | 0.0030002 | 0.0001579 | 10Hz10m3m | -2.2293e-07 | 0.00085889 | 2.9999e-07 | 0.0078965 | 0.99993 | 9.9996e-05
10 | m | 0.01 | 20 | 20 | Hz | 3 | m | 0.003 | 2.9999 | 0.001721 | -0.00010531 | 0.0019506 | m | 0.001 | 0.0029999 | 1.721e-06 | -1.0531e-07 | 1.9506e-06 | 0.0029999 | -3.5106e-05 | 20Hz3m | 47 | 0.0029999 | 3e-08 | 1.9468e-07 | 1.5e-07 | 0.0029999 | 6.4897e-05 | 20Hz10m3m | -3.5033e-08 | 0.003442 | 3e-07 | 0.0039013 | 0.99999 | 0.0001
10 | m | 0.01 | 50 | 50 | Hz | 3 | m | 0.003 | 3.0005 | 0.00079115 | -0.00036171 | 0.00045857 | m | 0.001 | 0.0030005 | 7.9115e-07 | -3.6171e-07 | 4.5857e-07 | 0.0030005 | -0.00012055 | 50Hz3m | 48 | 0.0030005 | 3e-08 | -6.1695e-08 | 1.5e-07 | 0.0030005 | -2.0562e-05 | 50Hz10m3m | 3.0977e-09 | 0.0015823 | 3.0001e-07 | 0.00091715 | 1 | 9.9987e-05
10 | m | 0.01 | 100 | 100 | Hz | 3 | m | 0.003 | 3.0003 | 0.0021633 | -5.0972e-05 | 0.0026567 | m | 0.001 | 0.0030003 | 2.1633e-06 | -5.0972e-08 | 2.6567e-06 | 0.0030003 | -1.6989e-05 | 100Hz3m | 49 | 0.0030003 | 3e-08 | 2.4904e-07 | 1.5e-07 | 0.0030003 | 8.3006e-05 | 100Hz10m3m | 2.2391e-09 | 0.0043266 | 3.0001e-07 | 0.0053134 | 1 | 9.9996e-05
10 | m | 0.01 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0016 | 0.0027177 | -0.00023012 | 0.001046 | m | 0.001 | 0.0030016 | 2.7177e-06 | -2.3012e-07 | 1.046e-06 | 0.0030016 | -7.6666e-05 | 200Hz3m | 50 | 0.0030016 | 3e-08 | 6.9938e-08 | 1.5e-07 | 0.0030016 | 2.3301e-05 | 200Hz10m3m | -7.8893e-08 | 0.0054354 | 3.0006e-07 | 0.0020919 | 0.99997 | 9.9966e-05
10 | m | 0.01 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0063 | 0.00070678 | 0.00081088 | 0.0013515 | m | 0.001 | 0.0030063 | 7.0678e-07 | 8.1088e-07 | 1.3515e-06 | 0.0030063 | 0.00026973 | 500Hz3m | 51 | 0.0030066 | 4.7434e-08 | 1.1111e-06 | 1.5e-07 | 0.0030066 | 0.00036956 | 500Hz10m3m | 2.7795e-07 | 0.0014136 | 3.0022e-07 | 0.002703 | 1.0001 | 9.9828e-05
10 | m | 0.01 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0199 | 0.0040958 | -0.00015702 | 0.0011401 | m | 0.001 | 0.0030199 | 4.0958e-06 | -1.5702e-07 | 1.1401e-06 | 0.0030199 | -5.1995e-05 | 1kHz3m | 52 | 0.0030196 | 1.3416e-07 | 1.4353e-07 | 3e-07 | 0.0030196 | 4.7533e-05 | 1kHz10m3m | -2.7712e-07 | 0.0081916 | 3.0055e-07 | 0.0022801 | 0.99991 | 9.9528e-05
10 | m | 0.01 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0556 | 0.002183 | 0.0013079 | 0.0019183 | m | 0.001 | 0.0030556 | 2.183e-06 | 1.3079e-06 | 1.9183e-06 | 0.0030556 | 0.00042804 | 2kHz3m | 53 | 0.0030554 | 3.7947e-07 | 1.6117e-06 | 6e-07 | 0.0030554 | 0.00052749 | 2kHz10m3m | -1.3617e-07 | 0.004366 | 3.0381e-07 | 0.0038366 | 0.99996 | 9.9453e-05
10 | m | 0.01 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.2189 | 0.0038888 | 0.0038854 | 0.00040611 | m | 0.001 | 0.0032189 | 3.8888e-06 | 3.8854e-06 | 4.0611e-07 | 0.0032189 | 0.0012071 | 5kHz3m | 54 | 0.0032189 | 1.5e-06 | 4.1887e-06 | 1.5e-06 | 0.0032189 | 0.0013013 | 5kHz10m3m | 3.8348e-08 | 0.0077777 | 3.0329e-07 | 0.00081221 | 1 | 9.4207e-05
10 | m | 0.01 | 0.01 | 10 | mHz | 10 | m | 0.01 | 9.9982 | 0.0049749 | -0.0022835 | 0.0034987 | m | 0.001 | 0.0099982 | 4.9749e-06 | -2.2835e-06 | 3.4987e-06 | 0.0099982 | -0.00022839 | 10mHz10m | 55 | 0.0099991 | 1e-07 | -1.7068e-06 | 5e-07 | 0.0099991 | -0.00017069 | 10mHz10m10m | 9.212e-07 | 0.0099497 | 5.767e-07 | 0.0069975 | 1.0001 | 5.7696e-05
10 | m | 0.01 | 0.02 | 20 | mHz | 10 | m | 0.01 | 9.9994 | 0.0015321 | -0.00085095 | 0.0040117 | m | 0.001 | 0.0099994 | 1.5321e-06 | -8.5095e-07 | 4.0117e-06 | 0.0099994 | -8.51e-05 | 20mHz10m | 56 | 0.0099998 | 1e-07 | 2.7632e-08 | 5e-07 | 0.0099998 | 2.7632e-06 | 20mHz10m10m | 3.5118e-07 | 0.0030642 | 8.7858e-07 | 0.0080233 | 1 | 8.7863e-05
10 | m | 0.01 | 0.05 | 50 | mHz | 10 | m | 0.01 | 10.001 | 0.0014952 | 0.0023917 | 0.0018398 | m | 0.001 | 0.010001 | 1.4952e-06 | 2.3917e-06 | 1.8398e-06 | 0.010001 | 0.00023914 | 50mHz10m | 57 | 0.010001 | 1e-07 | 2.1969e-06 | 5e-07 | 0.010001 | 0.00021967 | 50mHz10m10m | -2.6057e-07 | 0.0029905 | -1.9484e-07 | 0.0036795 | 0.99997 | -1.9476e-05
10 | m | 0.01 | 0.1 | 100 | mHz | 10 | m | 0.01 | 9.9996 | 0.00014925 | 0.00052405 | 0.0018637 | m | 0.001 | 0.0099996 | 1.4925e-07 | 5.2405e-07 | 1.8637e-06 | 0.0099996 | 5.2407e-05 | 100mHz10m | 58 | 0.01 | 1e-07 | 4.9047e-07 | 5e-07 | 0.01 | 4.9046e-05 | 100mHz10m10m | 6.6571e-07 | 0.00029851 | -3.3573e-08 | 0.0037274 | 1.0001 | -3.3607e-06
10 | m | 0.01 | 0.2 | 200 | mHz | 10 | m | 0.01 | 10.0015 | 0.0015098 | -0.0020127 | 0.0034631 | m | 0.001 | 0.010001 | 1.5098e-06 | -2.0127e-06 | 3.4631e-06 | 0.010001 | -0.00020124 | 200mHz10m | 59 | 0.010002 | 1e-07 | -1.8749e-06 | 5e-07 | 0.010002 | -0.00018745 | 200mHz10m10m | 2.3191e-07 | 0.0030196 | 1.3784e-07 | 0.0069261 | 1 | 1.3786e-05
10 | m | 0.01 | 0.5 | 500 | mHz | 10 | m | 0.01 | 10.002 | 0.004018 | -0.0013894 | 0.0038018 | m | 0.001 | 0.010002 | 4.018e-06 | -1.3894e-06 | 3.8018e-06 | 0.010002 | -0.00013892 | 500mHz10m | 60 | 0.010001 | 1e-07 | -1.5567e-06 | 5e-07 | 0.010001 | -0.00015565 | 500mHz10m10m | -6.9421e-07 | 0.008036 | -1.6727e-07 | 0.0076037 | 0.99993 | -1.6735e-05
10 | m | 0.01 | 1 | 1 | Hz | 10 | m | 0.01 | 9.9991 | 0.0019757 | -0.00021096 | 0.00044874 | m | 0.001 | 0.0099991 | 1.9757e-06 | -2.1096e-07 | 4.4874e-07 | 0.0099991 | -2.1098e-05 | 1Hz10m | 61 | 0.0099989 | 1e-07 | 3.6524e-07 | 5e-07 | 0.0099989 | 3.6528e-05 | 1Hz10m10m | -1.8985e-07 | 0.0039514 | 5.762e-07 | 0.00089748 | 0.99998 | 5.7626e-05
10 | m | 0.01 | 2 | 2 | Hz | 10 | m | 0.01 | 10.0002 | 0.0018506 | -0.00023831 | 0.0015725 | m | 0.001 | 0.01 | 1.8506e-06 | -2.3831e-07 | 1.5725e-06 | 0.01 | -2.383e-05 | 2Hz10m | 62 | 0.010001 | 1e-07 | 1.1614e-07 | 5e-07 | 0.010001 | 1.1613e-05 | 2Hz10m10m | 4.1052e-07 | 0.0037012 | 3.5445e-07 | 0.0031449 | 1 | 3.5443e-05
10 | m | 0.01 | 5 | 5 | Hz | 10 | m | 0.01 | 9.9992 | 0.0048731 | 0.0008821 | 0.0044325 | m | 0.001 | 0.0099992 | 4.8731e-06 | 8.821e-07 | 4.4325e-06 | 0.0099992 | 8.8217e-05 | 5Hz10m | 63 | 0.0099993 | 1e-07 | 1.4423e-06 | 5e-07 | 0.0099993 | 0.00014424 | 5Hz10m10m | 9.1093e-08 | 0.0097463 | 5.6021e-07 | 0.0088651 | 1 | 5.6024e-05
10 | m | 0.01 | 10 | 10 | Hz | 10 | m | 0.01 | 10.0013 | 0.0024643 | -0.00090992 | 0.00056462 | m | 0.001 | 0.010001 | 2.4643e-06 | -9.0992e-07 | 5.6462e-07 | 0.010001 | -9.0981e-05 | 10Hz10m | 64 | 0.010001 | 1e-07 | -1.481e-06 | 5e-07 | 0.010001 | -0.00014808 | 10Hz10m10m | 2.0131e-07 | 0.0049286 | -5.7107e-07 | 0.0011292 | 1 | -5.7097e-05
10 | m | 0.01 | 20 | 20 | Hz | 10 | m | 0.01 | 9.9989 | 0.00063667 | 0.0020583 | 0.00011386 | m | 0.001 | 0.0099989 | 6.3667e-07 | 2.0583e-06 | 1.1386e-07 | 0.0099989 | 0.00020585 | 20Hz10m | 65 | 0.0099995 | 1e-07 | 1.1958e-06 | 5e-07 | 0.0099995 | 0.00011958 | 20Hz10m10m | 6.3819e-07 | 0.0012733 | -8.6252e-07 | 0.00022773 | 1.0001 | -8.6269e-05
10 | m | 0.01 | 50 | 50 | Hz | 10 | m | 0.01 | 10.0025 | 0.003284 | 0.00039973 | 0.0045323 | m | 0.001 | 0.010003 | 3.284e-06 | 3.9973e-07 | 4.5323e-06 | 0.010003 | 3.9963e-05 | 50Hz10m | 66 | 0.010002 | 1e-07 | -3.658e-07 | 5e-07 | 0.010002 | -3.6572e-05 | 50Hz10m10m | -3.1031e-07 | 0.006568 | -7.6553e-07 | 0.0090646 | 0.99997 | -7.6535e-05
10 | m | 0.01 | 100 | 100 | Hz | 10 | m | 0.01 | 10.0018 | 0.00069426 | 0.00058825 | 0.0029669 | m | 0.001 | 0.010002 | 6.9426e-07 | 5.8825e-07 | 2.9669e-06 | 0.010002 | 5.8814e-05 | 100Hz10m | 67 | 0.010002 | 1e-07 | 4.8189e-07 | 5e-07 | 0.010002 | 4.8178e-05 | 100Hz10m10m | 6.3242e-07 | 0.0013885 | -1.0635e-07 | 0.0059338 | 1.0001 | -1.0637e-05
10 | m | 0.01 | 200 | 200 | Hz | 10 | m | 0.01 | 10.0048 | 0.0032281 | 0.00096769 | 0.0021591 | m | 0.001 | 0.010005 | 3.2281e-06 | 9.6769e-07 | 2.1591e-06 | 0.010005 | 9.6723e-05 | 200Hz10m | 68 | 0.010006 | 1e-07 | 4.5743e-07 | 5e-07 | 0.010006 | 4.5717e-05 | 200Hz10m10m | 9.884e-07 | 0.0064562 | -5.1026e-07 | 0.0043182 | 1.0001 | -5.1006e-05
10 | m | 0.01 | 500 | 500 | Hz | 10 | m | 0.01 | 10.021 | 0.0047096 | -0.00013829 | 0.00039345 | m | 0.001 | 0.010021 | 4.7096e-06 | -1.3829e-07 | 3.9345e-07 | 0.010021 | -1.38e-05 | 500Hz10m | 69 | 0.010022 | 1.5811e-07 | -9.748e-07 | 5e-07 | 0.010022 | -9.7269e-05 | 500Hz10m10m | 6.7384e-07 | 0.0094192 | -8.365e-07 | 0.00078689 | 1.0001 | -8.3468e-05
10 | m | 0.01 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.0656 | 0.0017331 | 0.0027824 | 0.00344 | m | 0.001 | 0.010066 | 1.7331e-06 | 2.7824e-06 | 3.44e-06 | 0.010066 | 0.00027643 | 1kHz10m | 70 | 0.010066 | 4.4721e-07 | 2.0772e-06 | 1e-06 | 0.010066 | 0.00020635 | 1kHz10m10m | 5.1261e-07 | 0.0034662 | -7.0524e-07 | 0.0068801 | 1.0001 | -7.0075e-05
10 | m | 0.01 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.1847 | 0.0038347 | 0.0054068 | 0.002815 | m | 0.001 | 0.010185 | 3.8347e-06 | 5.4068e-06 | 2.815e-06 | 0.010185 | 0.00053087 | 2kHz10m | 71 | 0.010185 | 1.2649e-06 | 5.7029e-06 | 2e-06 | 0.010185 | 0.00055994 | 2kHz10m10m | -1.5414e-08 | 0.0076693 | 2.9611e-07 | 0.00563 | 1 | 2.9074e-05
10 | m | 0.01 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.7299 | 0.002279 | 0.01455 | 0.0029989 | m | 0.001 | 0.01073 | 2.279e-06 | 1.455e-05 | 2.9989e-06 | 0.01073 | 0.0013561 | 5kHz10m | 72 | 0.01073 | 5e-06 | 1.3428e-05 | 5e-06 | 0.01073 | 0.0012515 | 5kHz10m10m | -2.9561e-07 | 0.004558 | -1.122e-06 | 0.0059978 | 0.99997 | -0.00010453
100 | m | 0.1 | 0.01 | 10 | mHz | 10 | m | 0.01 | 9.9982 | 0.0017036 | -0.0027065 | 0.0035288 | m | 0.001 | 0.0099982 | 1.7036e-06 | -2.7065e-06 | 3.5288e-06 | 0.0099982 | -0.00027069 | 10mHz10m | 55 | 0.0099991 | 1e-07 | -1.7068e-06 | 5e-07 | 0.0099991 | -0.00017069 | 10mHz100m10m | 8.6773e-07 | 0.0034071 | 9.9969e-07 | 0.0070576 | 1.0001 | 0.0001
100 | m | 0.1 | 0.02 | 20 | mHz | 10 | m | 0.01 | 9.9991 | 0.0038325 | -0.00097228 | 0.0020238 | m | 0.001 | 0.0099991 | 3.8325e-06 | -9.7228e-07 | 2.0238e-06 | 0.0099991 | -9.7237e-05 | 20mHz10m | 56 | 0.0099998 | 1e-07 | 2.7632e-08 | 5e-07 | 0.0099998 | 2.7632e-06 | 20mHz100m10m | 6.9034e-07 | 0.0076651 | 9.9991e-07 | 0.0040476 | 1.0001 | 0.0001
100 | m | 0.1 | 0.05 | 50 | mHz | 10 | m | 0.01 | 10.0016 | 0.00035193 | 0.0011969 | 0.0038617 | m | 0.001 | 0.010002 | 3.5193e-07 | 1.1969e-06 | 3.8617e-06 | 0.010002 | 0.00011967 | 50mHz10m | 57 | 0.010001 | 1e-07 | 2.1969e-06 | 5e-07 | 0.010001 | 0.00021967 | 50mHz100m10m | -8.812e-07 | 0.00070386 | 9.9995e-07 | 0.0077234 | 0.99991 | 9.9998e-05
100 | m | 0.1 | 0.1 | 100 | mHz | 10 | m | 0.01 | 9.9995 | 0.00011537 | -0.00050951 | 0.0013513 | m | 0.001 | 0.0099995 | 1.1537e-07 | -5.0951e-07 | 1.3513e-06 | 0.0099995 | -5.0953e-05 | 100mHz10m | 58 | 0.01 | 1e-07 | 4.9047e-07 | 5e-07 | 0.01 | 4.9046e-05 | 100mHz100m10m | 7.2738e-07 | 0.00023073 | 9.9998e-07 | 0.0027026 | 1.0001 | 0.0001
100 | m | 0.1 | 0.2 | 200 | mHz | 10 | m | 0.01 | 10.0011 | 0.0018417 | -0.0028748 | 0.0036152 | m | 0.001 | 0.010001 | 1.8417e-06 | -2.8748e-06 | 3.6152e-06 | 0.010001 | -0.00028745 | 200mHz10m | 59 | 0.010002 | 1e-07 | -1.8749e-06 | 5e-07 | 0.010002 | -0.00018745 | 200mHz100m10m | 6.4711e-07 | 0.0036835 | 9.9995e-07 | 0.0072304 | 1.0001 | 9.9997e-05
100 | m | 0.1 | 0.5 | 500 | mHz | 10 | m | 0.01 | 10.0021 | 8.8043e-05 | -0.002557 | 0.0002944 | m | 0.001 | 0.010002 | 8.8043e-08 | -2.557e-06 | 2.944e-07 | 0.010002 | -0.00025565 | 500mHz10m | 60 | 0.010001 | 1e-07 | -1.5567e-06 | 5e-07 | 0.010001 | -0.00015565 | 500mHz100m10m | -8.2605e-07 | 0.00017609 | 1.0003e-06 | 0.0005888 | 0.99992 | 9.9998e-05
100 | m | 0.1 | 1 | 1 | Hz | 10 | m | 0.01 | 9.9995 | 0.0047044 | -0.0006347 | 5.6579e-05 | m | 0.001 | 0.0099995 | 4.7044e-06 | -6.347e-07 | 5.6579e-08 | 0.0099995 | -6.3473e-05 | 1Hz10m | 61 | 0.0099989 | 1e-07 | 3.6524e-07 | 5e-07 | 0.0099989 | 3.6528e-05 | 1Hz100m10m | -5.7964e-07 | 0.0094088 | 9.9994e-07 | 0.00011316 | 0.99994 | 0.0001
100 | m | 0.1 | 2 | 2 | Hz | 10 | m | 0.01 | 10.0011 | 0.00069547 | -0.00088396 | 0.0031789 | m | 0.001 | 0.010001 | 6.9547e-07 | -8.8396e-07 | 3.1789e-06 | 0.010001 | -8.8387e-05 | 2Hz10m | 62 | 0.010001 | 1e-07 | 1.1614e-07 | 5e-07 | 0.010001 | 1.1613e-05 | 2Hz100m10m | -4.3996e-07 | 0.0013909 | 1.0001e-06 | 0.0063578 | 0.99996 | 0.0001
100 | m | 0.1 | 5 | 5 | Hz | 10 | m | 0.01 | 9.9986 | 0.0050754 | 0.00044234 | 0.0025159 | m | 0.001 | 0.0099986 | 5.0754e-06 | 4.4234e-07 | 2.5159e-06 | 0.0099986 | 4.424e-05 | 5Hz10m | 63 | 0.0099993 | 1e-07 | 1.4423e-06 | 5e-07 | 0.0099993 | 0.00014424 | 5Hz100m10m | 6.6278e-07 | 0.010151 | 9.9997e-07 | 0.0050317 | 1.0001 | 0.0001
100 | m | 0.1 | 10 | 10 | Hz | 10 | m | 0.01 | 10.0011 | 0.0041053 | -0.002481 | 0.0015122 | m | 0.001 | 0.010001 | 4.1053e-06 | -2.481e-06 | 1.5122e-06 | 0.010001 | -0.00024808 | 10Hz10m | 64 | 0.010001 | 1e-07 | -1.481e-06 | 5e-07 | 0.010001 | -0.00014808 | 10Hz100m10m | 3.7878e-07 | 0.0082107 | 1e-06 | 0.0030243 | 1 | 9.9998e-05
100 | m | 0.1 | 20 | 20 | Hz | 10 | m | 0.01 | 9.9998 | 0.0032701 | 0.00019582 | 0.0020599 | m | 0.001 | 0.0099998 | 3.2701e-06 | 1.9582e-07 | 2.0599e-06 | 0.0099998 | 1.9582e-05 | 20Hz10m | 65 | 0.0099995 | 1e-07 | 1.1958e-06 | 5e-07 | 0.0099995 | 0.00011958 | 20Hz100m10m | -3.5692e-07 | 0.0065402 | 9.9995e-07 | 0.0041199 | 0.99996 | 0.0001
100 | m | 0.1 | 50 | 50 | Hz | 10 | m | 0.01 | 10.0027 | 0.0029015 | -0.0013661 | 0.0035363 | m | 0.001 | 0.010003 | 2.9015e-06 | -1.3661e-06 | 3.5363e-06 | 0.010003 | -0.00013657 | 50Hz10m | 66 | 0.010002 | 1e-07 | -3.658e-07 | 5e-07 | 0.010002 | -3.6572e-05 | 50Hz100m10m | -4.4594e-07 | 0.005803 | 1.0003e-06 | 0.0070726 | 0.99996 | 9.9998e-05
100 | m | 0.1 | 100 | 100 | Hz | 10 | m | 0.01 | 10.0023 | 0.0032527 | -0.00051834 | 0.0044132 | m | 0.001 | 0.010002 | 3.2527e-06 | -5.1834e-07 | 4.4132e-06 | 0.010002 | -5.1822e-05 | 100Hz10m | 67 | 0.010002 | 1e-07 | 4.8189e-07 | 5e-07 | 0.010002 | 4.8178e-05 | 100Hz100m10m | 8.5742e-08 | 0.0065053 | 1.0002e-06 | 0.0088263 | 1 | 0.0001
100 | m | 0.1 | 200 | 200 | Hz | 10 | m | 0.01 | 10.0054 | 0.00018739 | -0.00054306 | 0.0044328 | m | 0.001 | 0.010005 | 1.8739e-07 | -5.4306e-07 | 4.4328e-06 | 0.010005 | -5.4277e-05 | 200Hz10m | 68 | 0.010006 | 1e-07 | 4.5743e-07 | 5e-07 | 0.010006 | 4.5717e-05 | 200Hz100m10m | 3.9575e-07 | 0.00037478 | 1.0005e-06 | 0.0088655 | 1 | 9.9994e-05
100 | m | 0.1 | 500 | 500 | Hz | 10 | m | 0.01 | 10.0225 | 0.00022205 | -0.0019771 | 0.0014786 | m | 0.001 | 0.010023 | 2.2205e-07 | -1.9771e-06 | 1.4786e-06 | 0.010023 | -0.00019727 | 500Hz10m | 69 | 0.010022 | 1.5811e-07 | -9.748e-07 | 5e-07 | 0.010022 | -9.7269e-05 | 500Hz100m10m | -8.5263e-07 | 0.0004441 | 1.0023e-06 | 0.0029572 | 0.99991 | 9.9999e-05
100 | m | 0.1 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.0653 | 0.004403 | 0.0010717 | 0.0012427 | m | 0.001 | 0.010065 | 4.403e-06 | 1.0717e-06 | 1.2427e-06 | 0.010065 | 0.00010648 | 1kHz10m | 70 | 0.010066 | 4.4721e-07 | 2.0772e-06 | 1e-06 | 0.010066 | 0.00020635 | 1kHz100m10m | 8.1145e-07 | 0.0088059 | 1.0054e-06 | 0.0024854 | 1.0001 | 9.9876e-05
100 | m | 0.1 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.1845 | 0.0024118 | 0.0046875 | 0.0025247 | m | 0.001 | 0.010185 | 2.4118e-06 | 4.6875e-06 | 2.5247e-06 | 0.010185 | 0.00046026 | 2kHz10m | 71 | 0.010185 | 1.2649e-06 | 5.7029e-06 | 2e-06 | 0.010185 | 0.00055994 | 2kHz100m10m | 1.8479e-07 | 0.0048235 | 1.0154e-06 | 0.0050493 | 1 | 9.9687e-05
100 | m | 0.1 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.7306 | 0.0030846 | 0.01236 | 0.0013746 | m | 0.001 | 0.010731 | 3.0846e-06 | 1.236e-05 | 1.3746e-06 | 0.010731 | 0.0011518 | 5kHz10m | 72 | 0.01073 | 5e-06 | 1.3428e-05 | 5e-06 | 0.01073 | 0.0012515 | 5kHz100m10m | -1.069e-06 | 0.0061692 | 1.0685e-06 | 0.0027491 | 0.9999 | 9.9703e-05
100 | m | 0.1 | 0.01 | 10 | mHz | 100 | m | 0.1 | 100.0045 | 0.0020323 | -0.0019606 | 0.0017697 | m | 0.001 | 0.1 | 2.0323e-06 | -1.9606e-06 | 1.7697e-06 | 0.1 | -1.9605e-05 | 10mHz100m | 73 | 0.099998 | 1e-06 | -3.6936e-06 | 5e-06 | 0.099998 | -3.6936e-05 | 10mHz100m100m | -6.7717e-06 | 0.0040646 | -1.733e-06 | 0.0035394 | 0.99993 | -1.7332e-05
100 | m | 0.1 | 0.02 | 20 | mHz | 100 | m | 0.1 | 100.0008 | 0.0055009 | 0.014614 | 0.0028175 | m | 0.001 | 0.1 | 5.5009e-06 | 1.4614e-05 | 2.8175e-06 | 0.1 | 0.00014613 | 20mHz100m | 74 | 0.099992 | 1e-06 | 7.0064e-06 | 5e-06 | 0.099992 | 7.0069e-05 | 20mHz100m100m | -8.554e-06 | 0.011002 | -7.6072e-06 | 0.005635 | 0.99991 | -7.6065e-05
100 | m | 0.1 | 0.05 | 50 | mHz | 100 | m | 0.1 | 100.0217 | 0.0034313 | 0.016804 | 0.0038737 | m | 0.001 | 0.10002 | 3.4313e-06 | 1.6804e-05 | 3.8737e-06 | 0.10002 | 0.000168 | 50mHz100m | 75 | 0.10002 | 1e-06 | 8.2678e-06 | 5e-06 | 0.10002 | 8.2664e-05 | 50mHz100m100m | -5.2374e-06 | 0.0068627 | -8.5358e-06 | 0.0077474 | 0.99995 | -8.5335e-05
100 | m | 0.1 | 0.1 | 100 | mHz | 100 | m | 0.1 | 99.997 | 0.00072563 | 0.01258 | 0.0036827 | m | 0.001 | 0.099997 | 7.2563e-07 | 1.258e-05 | 3.6827e-06 | 0.099997 | 0.00012581 | 100mHz100m | 76 | 0.099996 | 1e-06 | 1.2929e-05 | 5e-06 | 0.099996 | 0.00012929 | 100mHz100m100m | -1.0097e-06 | 0.0014513 | 3.4827e-07 | 0.0073655 | 0.99999 | 3.4841e-06
100 | m | 0.1 | 0.2 | 200 | mHz | 100 | m | 0.1 | 99.9956 | 0.0010857 | 0.004354 | 0.004103 | m | 0.001 | 0.099996 | 1.0857e-06 | 4.354e-06 | 4.103e-06 | 0.099996 | 4.3542e-05 | 200mHz100m | 77 | 0.1 | 1e-06 | -1.2329e-06 | 5e-06 | 0.1 | -1.2329e-05 | 200mHz100m100m | 4.5167e-06 | 0.0021714 | -5.5869e-06 | 0.0082059 | 1 | -5.5871e-05
100 | m | 0.1 | 0.5 | 500 | mHz | 100 | m | 0.1 | 100.0115 | 0.0043481 | 0.0090002 | 0.00042901 | m | 0.001 | 0.10001 | 4.3481e-06 | 9.0002e-06 | 4.2901e-07 | 0.10001 | 8.9992e-05 | 500mHz100m | 78 | 0.1 | 1e-06 | 2.8994e-06 | 5e-06 | 0.1 | 2.8993e-05 | 500mHz100m100m | -6.6614e-06 | 0.0086963 | -6.1008e-06 | 0.00085803 | 0.99993 | -6.0999e-05
100 | m | 0.1 | 1 | 1 | Hz | 100 | m | 0.1 | 99.9905 | 0.0039537 | -0.0012138 | 0.00014729 | m | 0.001 | 0.099991 | 3.9537e-06 | -1.2138e-06 | 1.4729e-07 | 0.099991 | -1.2139e-05 | 1Hz100m | 79 | 0.099983 | 1e-06 | -1.823e-06 | 5e-06 | 0.099983 | -1.8233e-05 | 1Hz100m100m | -7.7257e-06 | 0.0079075 | -6.0924e-07 | 0.00029463 | 0.99992 | -6.0944e-06
100 | m | 0.1 | 2 | 2 | Hz | 100 | m | 0.1 | 100.0136 | 0.0038725 | -0.0044384 | 0.0029632 | m | 0.001 | 0.10001 | 3.8725e-06 | -4.4384e-06 | 2.9632e-06 | 0.10001 | -4.4378e-05 | 2Hz100m | 80 | 0.10001 | 1e-06 | -1.1702e-06 | 5e-06 | 0.10001 | -1.17e-05 | 2Hz100m100m | -1.8898e-06 | 0.007745 | 3.2683e-06 | 0.0059264 | 0.99998 | 3.2678e-05
100 | m | 0.1 | 5 | 5 | Hz | 100 | m | 0.1 | 100.0092 | 0.0018042 | 0.006234 | 7.5286e-05 | m | 0.001 | 0.10001 | 1.8042e-06 | 6.234e-06 | 7.5286e-08 | 0.10001 | 6.2334e-05 | 5Hz100m | 81 | 0.10001 | 1e-06 | -2.9868e-06 | 5e-06 | 0.10001 | -2.9865e-05 | 5Hz100m100m | 1.8154e-06 | 0.0036084 | -9.2207e-06 | 0.00015066 | 1 | -9.2198e-05
100 | m | 0.1 | 10 | 10 | Hz | 100 | m | 0.1 | 99.9985 | 0.0050474 | -0.0010948 | 0.0030927 | m | 0.001 | 0.099999 | 5.0474e-06 | -1.0948e-06 | 3.0927e-06 | 0.099999 | -1.0948e-05 | 10Hz100m | 82 | 0.099997 | 1e-06 | -5.0857e-06 | 5e-06 | 0.099997 | -5.0858e-05 | 10Hz100m100m | -1.7212e-06 | 0.010095 | -3.9909e-06 | 0.0061855 | 0.99998 | -3.9911e-05
100 | m | 0.1 | 20 | 20 | Hz | 100 | m | 0.1 | 100.0044 | 0.004357 | 0.014598 | 0.00047022 | m | 0.001 | 0.1 | 4.357e-06 | 1.4598e-05 | 4.7022e-07 | 0.1 | 0.00014597 | 20Hz100m | 83 | 0.1 | 1e-06 | 7.8637e-06 | 5e-06 | 0.1 | 7.8634e-05 | 20Hz100m100m | -8.9075e-07 | 0.0087139 | -6.7344e-06 | 0.00094046 | 0.99999 | -6.734e-05
100 | m | 0.1 | 50 | 50 | Hz | 100 | m | 0.1 | 99.9901 | 0.0049644 | 0.0071892 | 0.0010985 | m | 0.001 | 0.09999 | 4.9644e-06 | 7.1892e-06 | 1.0985e-06 | 0.09999 | 7.1899e-05 | 50Hz100m | 84 | 0.099986 | 1e-06 | 7.0349e-06 | 5e-06 | 0.099986 | 7.0359e-05 | 50Hz100m100m | -4.2272e-06 | 0.0099288 | -1.5433e-07 | 0.002197 | 0.99996 | -1.5405e-06
100 | m | 0.1 | 100 | 100 | Hz | 100 | m | 0.1 | 100.0353 | 0.0026993 | -0.0087863 | 0.002258 | m | 0.001 | 0.10004 | 2.6993e-06 | -8.7863e-06 | 2.258e-06 | 0.10004 | -8.7832e-05 | 100Hz100m | 85 | 0.10003 | 1e-06 | 1.8682e-07 | 5e-06 | 0.10003 | 1.8675e-06 | 100Hz100m100m | -1.855e-06 | 0.0053987 | 8.9731e-06 | 0.0045161 | 0.99998 | 8.9699e-05
100 | m | 0.1 | 200 | 200 | Hz | 100 | m | 0.1 | 100.0669 | 0.0011625 | -0.0017512 | 0.0033275 | m | 0.001 | 0.10007 | 1.1625e-06 | -1.7512e-06 | 3.3275e-06 | 0.10007 | -1.7501e-05 | 200Hz100m | 86 | 0.10007 | 1e-06 | -6.8813e-07 | 5e-06 | 0.10007 | -6.8764e-06 | 200Hz100m100m | 4.1712e-06 | 0.0023249 | 1.0631e-06 | 0.006655 | 1 | 1.0624e-05
100 | m | 0.1 | 500 | 500 | Hz | 100 | m | 0.1 | 100.214 | 0.0044934 | -0.0082178 | 0.0026076 | m | 0.001 | 0.10021 | 4.4934e-06 | -8.2178e-06 | 2.6076e-06 | 0.10021 | -8.2003e-05 | 500Hz100m | 87 | 0.10022 | 1.5811e-06 | 1.3827e-06 | 5e-06 | 0.10022 | 1.3797e-05 | 500Hz100m100m | 6.5389e-06 | 0.0089867 | 9.6006e-06 | 0.0052153 | 1.0001 | 9.58e-05
100 | m | 0.1 | 1000 | 1 | kHz | 100 | m | 0.1 | 100.6516 | 0.0022905 | 0.031388 | 0.0034002 | m | 0.001 | 0.10065 | 2.2905e-06 | 3.1388e-05 | 3.4002e-06 | 0.10065 | 0.00031185 | 1kHz100m | 88 | 0.10066 | 4.4721e-06 | 2.4327e-05 | 1e-05 | 0.10066 | 0.00024167 | 1kHz100m100m | 9.6046e-06 | 0.0045809 | -7.0617e-06 | 0.0068005 | 1.0001 | -7.0183e-05
100 | m | 0.1 | 2000 | 2 | kHz | 100 | m | 0.1 | 101.8563 | 0.0051025 | 0.07242 | 0.0045086 | m | 0.001 | 0.10186 | 5.1025e-06 | 7.242e-05 | 4.5086e-06 | 0.10186 | 0.000711 | 2kHz100m | 89 | 0.10185 | 1.2649e-05 | 6.6749e-05 | 2e-05 | 0.10185 | 0.00065534 | 2kHz100m100m | -2.5564e-06 | 0.010205 | -5.671e-06 | 0.0090173 | 0.99997 | -5.566e-05
100 | m | 0.1 | 5000 | 5 | kHz | 100 | m | 0.1 | 107.299 | 0.0032562 | 0.13665 | 0.0010664 | m | 0.001 | 0.1073 | 3.2562e-06 | 0.00013665 | 1.0664e-06 | 0.1073 | 0.0012736 | 5kHz100m | 90 | 0.10729 | 5e-05 | 0.00014313 | 5e-05 | 0.10729 | 0.001334 | 5kHz100m100m | -1.0304e-05 | 0.0065127 | 6.4732e-06 | 0.0021333 | 0.9999 | 6.0456e-05

## Ref Z list
Ref. Z list
¦Z¦ |  | ¦Z¦ | Z tag
Ohm |  | Ohm | -
0 | m | 0 | 0m
1 | m | 0.001 | 1m
3 | m | 0.003 | 3m
10 | m | 0.01 | 10m
100 | m | 0.1 | 100m

## Ref Z
This is sheet containing calibration data of reference standards to be used for calibration of RLC meter.
f | f |  | Nominal Z |  | Nom Z | Rs | U(Rs) | Xs | U(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | tag
Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | -
0.01 | 10 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10mHz0m
0.02 | 20 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20mHz0m
0.05 | 50 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50mHz0m
0.1 | 100 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100mHz0m
0.2 | 200 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200mHz0m
0.5 | 500 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500mHz0m
1 | 1 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1Hz0m
2 | 2 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2Hz0m
5 | 5 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5Hz0m
10 | 10 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10Hz0m
20 | 20 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20Hz0m
50 | 50 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50Hz0m
100 | 100 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100Hz0m
200 | 200 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200Hz0m
500 | 500 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500Hz0m
1000 | 1 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1kHz0m
2000 | 2 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2kHz0m
5000 | 5 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5kHz0m
0.01 | 10 | mHz | 1 | m | 0.001 | 0.99995 | 1e-05 | 6.7794e-05 | 5e-05 | m | 0.001 | 0.00099995 | 1e-08 | 6.7794e-08 | 5e-08 | 10mHz1m
0.02 | 20 | mHz | 1 | m | 0.001 | 0.99995 | 1e-05 | -2.6701e-05 | 5e-05 | m | 0.001 | 0.00099995 | 1e-08 | -2.6701e-08 | 5e-08 | 20mHz1m
0.05 | 50 | mHz | 1 | m | 0.001 | 1 | 1e-05 | 0.00016256 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | 1.6256e-07 | 5e-08 | 50mHz1m
0.1 | 100 | mHz | 1 | m | 0.001 | 1.0001 | 1e-05 | 6.7205e-05 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 6.7205e-08 | 5e-08 | 100mHz1m
0.2 | 200 | mHz | 1 | m | 0.001 | 1 | 1e-05 | -3.1063e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | -3.1063e-08 | 5e-08 | 200mHz1m
0.5 | 500 | mHz | 1 | m | 0.001 | 0.99999 | 1e-05 | -3.3372e-05 | 5e-05 | m | 0.001 | 0.00099999 | 1e-08 | -3.3372e-08 | 5e-08 | 500mHz1m
1 | 1 | Hz | 1 | m | 0.001 | 1 | 1e-05 | -4.9729e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | -4.9729e-08 | 5e-08 | 1Hz1m
2 | 2 | Hz | 1 | m | 0.001 | 0.99998 | 1e-05 | -3.0857e-05 | 5e-05 | m | 0.001 | 0.00099998 | 1e-08 | -3.0857e-08 | 5e-08 | 2Hz1m
5 | 5 | Hz | 1 | m | 0.001 | 1 | 1e-05 | -7.5652e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | -7.5652e-08 | 5e-08 | 5Hz1m
10 | 10 | Hz | 1 | m | 0.001 | 0.99993 | 1e-05 | -0.00012146 | 5e-05 | m | 0.001 | 0.00099993 | 1e-08 | -1.2146e-07 | 5e-08 | 10Hz1m
20 | 20 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | -5.5145e-05 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | -5.5145e-08 | 5e-08 | 20Hz1m
50 | 50 | Hz | 1 | m | 0.001 | 1 | 1e-05 | 4.9221e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | 4.9221e-08 | 5e-08 | 50Hz1m
100 | 100 | Hz | 1 | m | 0.001 | 1.0004 | 1e-05 | -8.2709e-06 | 5e-05 | m | 0.001 | 0.0010004 | 1e-08 | -8.2709e-09 | 5e-08 | 100Hz1m
200 | 200 | Hz | 1 | m | 0.001 | 1.0005 | 1e-05 | -3.8301e-05 | 5e-05 | m | 0.001 | 0.0010005 | 1e-08 | -3.8301e-08 | 5e-08 | 200Hz1m
500 | 500 | Hz | 1 | m | 0.001 | 1.0023 | 1.5811e-05 | -2.2623e-05 | 5e-05 | m | 0.001 | 0.0010023 | 1.5811e-08 | -2.2623e-08 | 5e-08 | 500Hz1m
1000 | 1 | kHz | 1 | m | 0.001 | 1.0064 | 4.4721e-05 | 0.00031807 | 0.0001 | m | 0.001 | 0.0010064 | 4.4721e-08 | 3.1807e-07 | 1e-07 | 1kHz1m
2000 | 2 | kHz | 1 | m | 0.001 | 1.0183 | 0.00012649 | 0.00049418 | 0.0002 | m | 0.001 | 0.0010183 | 1.2649e-07 | 4.9418e-07 | 2e-07 | 2kHz1m
5000 | 5 | kHz | 1 | m | 0.001 | 1.0729 | 0.0005 | 0.0013752 | 0.0005 | m | 0.001 | 0.0010729 | 5e-07 | 1.3752e-06 | 5e-07 | 5kHz1m
0.01 | 10 | mHz | 3 | m | 0.003 | 2.9999 | 3e-05 | 0.00027585 | 0.00015 | m | 0.001 | 0.0029999 | 3e-08 | 2.7585e-07 | 1.5e-07 | 10mHz3m
0.02 | 20 | mHz | 3 | m | 0.003 | 3 | 3e-05 | 0.00074948 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | 7.4948e-07 | 1.5e-07 | 20mHz3m
0.05 | 50 | mHz | 3 | m | 0.003 | 3.0004 | 3e-05 | 0.00063046 | 0.00015 | m | 0.001 | 0.0030004 | 3e-08 | 6.3046e-07 | 1.5e-07 | 50mHz3m
0.1 | 100 | mHz | 3 | m | 0.003 | 2.9993 | 3e-05 | -0.00084074 | 0.00015 | m | 0.001 | 0.0029993 | 3e-08 | -8.4074e-07 | 1.5e-07 | 100mHz3m
0.2 | 200 | mHz | 3 | m | 0.003 | 2.9997 | 3e-05 | 0.00017762 | 0.00015 | m | 0.001 | 0.0029997 | 3e-08 | 1.7762e-07 | 1.5e-07 | 200mHz3m
0.5 | 500 | mHz | 3 | m | 0.003 | 2.9999 | 3e-05 | -0.00015149 | 0.00015 | m | 0.001 | 0.0029999 | 3e-08 | -1.5149e-07 | 1.5e-07 | 500mHz3m
1 | 1 | Hz | 3 | m | 0.003 | 2.9996 | 3e-05 | -0.00013225 | 0.00015 | m | 0.001 | 0.0029996 | 3e-08 | -1.3225e-07 | 1.5e-07 | 1Hz3m
2 | 2 | Hz | 3 | m | 0.003 | 3 | 3e-05 | 0.00020325 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | 2.0325e-07 | 1.5e-07 | 2Hz3m
5 | 5 | Hz | 3 | m | 0.003 | 3.0001 | 3e-05 | -0.00045252 | 0.00015 | m | 0.001 | 0.0030001 | 3e-08 | -4.5252e-07 | 1.5e-07 | 5Hz3m
10 | 10 | Hz | 3 | m | 0.003 | 3.0002 | 3e-05 | 0.00047373 | 0.00015 | m | 0.001 | 0.0030002 | 3e-08 | 4.7373e-07 | 1.5e-07 | 10Hz3m
20 | 20 | Hz | 3 | m | 0.003 | 2.9999 | 3e-05 | 0.00019468 | 0.00015 | m | 0.001 | 0.0029999 | 3e-08 | 1.9468e-07 | 1.5e-07 | 20Hz3m
50 | 50 | Hz | 3 | m | 0.003 | 3.0005 | 3e-05 | -6.1695e-05 | 0.00015 | m | 0.001 | 0.0030005 | 3e-08 | -6.1695e-08 | 1.5e-07 | 50Hz3m
100 | 100 | Hz | 3 | m | 0.003 | 3.0003 | 3e-05 | 0.00024904 | 0.00015 | m | 0.001 | 0.0030003 | 3e-08 | 2.4904e-07 | 1.5e-07 | 100Hz3m
200 | 200 | Hz | 3 | m | 0.003 | 3.0016 | 3e-05 | 6.9938e-05 | 0.00015 | m | 0.001 | 0.0030016 | 3e-08 | 6.9938e-08 | 1.5e-07 | 200Hz3m
500 | 500 | Hz | 3 | m | 0.003 | 3.0066 | 4.7434e-05 | 0.0011111 | 0.00015 | m | 0.001 | 0.0030066 | 4.7434e-08 | 1.1111e-06 | 1.5e-07 | 500Hz3m
1000 | 1 | kHz | 3 | m | 0.003 | 3.0196 | 0.00013416 | 0.00014353 | 0.0003 | m | 0.001 | 0.0030196 | 1.3416e-07 | 1.4353e-07 | 3e-07 | 1kHz3m
2000 | 2 | kHz | 3 | m | 0.003 | 3.0554 | 0.00037947 | 0.0016117 | 0.0006 | m | 0.001 | 0.0030554 | 3.7947e-07 | 1.6117e-06 | 6e-07 | 2kHz3m
5000 | 5 | kHz | 3 | m | 0.003 | 3.2189 | 0.0015 | 0.0041887 | 0.0015 | m | 0.001 | 0.0032189 | 1.5e-06 | 4.1887e-06 | 1.5e-06 | 5kHz3m
0.01 | 10 | mHz | 10 | m | 0.01 | 9.9991 | 0.0001 | -0.0017068 | 0.0005 | m | 0.001 | 0.0099991 | 1e-07 | -1.7068e-06 | 5e-07 | 10mHz10m
0.02 | 20 | mHz | 10 | m | 0.01 | 9.9998 | 0.0001 | 2.7632e-05 | 0.0005 | m | 0.001 | 0.0099998 | 1e-07 | 2.7632e-08 | 5e-07 | 20mHz10m
0.05 | 50 | mHz | 10 | m | 0.01 | 10.0007 | 0.0001 | 0.0021969 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 2.1969e-06 | 5e-07 | 50mHz10m
0.1 | 100 | mHz | 10 | m | 0.01 | 10.0002 | 0.0001 | 0.00049047 | 0.0005 | m | 0.001 | 0.01 | 1e-07 | 4.9047e-07 | 5e-07 | 100mHz10m
0.2 | 200 | mHz | 10 | m | 0.01 | 10.0017 | 0.0001 | -0.0018749 | 0.0005 | m | 0.001 | 0.010002 | 1e-07 | -1.8749e-06 | 5e-07 | 200mHz10m
0.5 | 500 | mHz | 10 | m | 0.01 | 10.0013 | 0.0001 | -0.0015567 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | -1.5567e-06 | 5e-07 | 500mHz10m
1 | 1 | Hz | 10 | m | 0.01 | 9.9989 | 0.0001 | 0.00036524 | 0.0005 | m | 0.001 | 0.0099989 | 1e-07 | 3.6524e-07 | 5e-07 | 1Hz10m
2 | 2 | Hz | 10 | m | 0.01 | 10.0006 | 0.0001 | 0.00011614 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 1.1614e-07 | 5e-07 | 2Hz10m
5 | 5 | Hz | 10 | m | 0.01 | 9.9993 | 0.0001 | 0.0014423 | 0.0005 | m | 0.001 | 0.0099993 | 1e-07 | 1.4423e-06 | 5e-07 | 5Hz10m
10 | 10 | Hz | 10 | m | 0.01 | 10.0015 | 0.0001 | -0.001481 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | -1.481e-06 | 5e-07 | 10Hz10m
20 | 20 | Hz | 10 | m | 0.01 | 9.9995 | 0.0001 | 0.0011958 | 0.0005 | m | 0.001 | 0.0099995 | 1e-07 | 1.1958e-06 | 5e-07 | 20Hz10m
50 | 50 | Hz | 10 | m | 0.01 | 10.0022 | 0.0001 | -0.0003658 | 0.0005 | m | 0.001 | 0.010002 | 1e-07 | -3.658e-07 | 5e-07 | 50Hz10m
100 | 100 | Hz | 10 | m | 0.01 | 10.0024 | 0.0001 | 0.00048189 | 0.0005 | m | 0.001 | 0.010002 | 1e-07 | 4.8189e-07 | 5e-07 | 100Hz10m
200 | 200 | Hz | 10 | m | 0.01 | 10.0058 | 0.0001 | 0.00045743 | 0.0005 | m | 0.001 | 0.010006 | 1e-07 | 4.5743e-07 | 5e-07 | 200Hz10m
500 | 500 | Hz | 10 | m | 0.01 | 10.0217 | 0.00015811 | -0.0009748 | 0.0005 | m | 0.001 | 0.010022 | 1.5811e-07 | -9.748e-07 | 5e-07 | 500Hz10m
1000 | 1 | kHz | 10 | m | 0.01 | 10.0661 | 0.00044721 | 0.0020772 | 0.001 | m | 0.001 | 0.010066 | 4.4721e-07 | 2.0772e-06 | 1e-06 | 1kHz10m
2000 | 2 | kHz | 10 | m | 0.01 | 10.1847 | 0.0012649 | 0.0057029 | 0.002 | m | 0.001 | 0.010185 | 1.2649e-06 | 5.7029e-06 | 2e-06 | 2kHz10m
5000 | 5 | kHz | 10 | m | 0.01 | 10.7296 | 0.005 | 0.013428 | 0.005 | m | 0.001 | 0.01073 | 5e-06 | 1.3428e-05 | 5e-06 | 5kHz10m
0.01 | 10 | mHz | 100 | m | 0.1 | 99.9977 | 0.001 | -0.0036936 | 0.005 | m | 0.001 | 0.099998 | 1e-06 | -3.6936e-06 | 5e-06 | 10mHz100m
0.02 | 20 | mHz | 100 | m | 0.1 | 99.9923 | 0.001 | 0.0070064 | 0.005 | m | 0.001 | 0.099992 | 1e-06 | 7.0064e-06 | 5e-06 | 20mHz100m
0.05 | 50 | mHz | 100 | m | 0.1 | 100.0165 | 0.001 | 0.0082678 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 8.2678e-06 | 5e-06 | 50mHz100m
0.1 | 100 | mHz | 100 | m | 0.1 | 99.996 | 0.001 | 0.012929 | 0.005 | m | 0.001 | 0.099996 | 1e-06 | 1.2929e-05 | 5e-06 | 100mHz100m
0.2 | 200 | mHz | 100 | m | 0.1 | 100.0001 | 0.001 | -0.0012329 | 0.005 | m | 0.001 | 0.1 | 1e-06 | -1.2329e-06 | 5e-06 | 200mHz100m
0.5 | 500 | mHz | 100 | m | 0.1 | 100.0049 | 0.001 | 0.0028994 | 0.005 | m | 0.001 | 0.1 | 1e-06 | 2.8994e-06 | 5e-06 | 500mHz100m
1 | 1 | Hz | 100 | m | 0.1 | 99.9828 | 0.001 | -0.001823 | 0.005 | m | 0.001 | 0.099983 | 1e-06 | -1.823e-06 | 5e-06 | 1Hz100m
2 | 2 | Hz | 100 | m | 0.1 | 100.0117 | 0.001 | -0.0011702 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -1.1702e-06 | 5e-06 | 2Hz100m
5 | 5 | Hz | 100 | m | 0.1 | 100.0111 | 0.001 | -0.0029868 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -2.9868e-06 | 5e-06 | 5Hz100m
10 | 10 | Hz | 100 | m | 0.1 | 99.9968 | 0.001 | -0.0050857 | 0.005 | m | 0.001 | 0.099997 | 1e-06 | -5.0857e-06 | 5e-06 | 10Hz100m
20 | 20 | Hz | 100 | m | 0.1 | 100.0035 | 0.001 | 0.0078637 | 0.005 | m | 0.001 | 0.1 | 1e-06 | 7.8637e-06 | 5e-06 | 20Hz100m
50 | 50 | Hz | 100 | m | 0.1 | 99.9859 | 0.001 | 0.0070349 | 0.005 | m | 0.001 | 0.099986 | 1e-06 | 7.0349e-06 | 5e-06 | 50Hz100m
100 | 100 | Hz | 100 | m | 0.1 | 100.0334 | 0.001 | 0.00018682 | 0.005 | m | 0.001 | 0.10003 | 1e-06 | 1.8682e-07 | 5e-06 | 100Hz100m
200 | 200 | Hz | 100 | m | 0.1 | 100.0711 | 0.001 | -0.00068813 | 0.005 | m | 0.001 | 0.10007 | 1e-06 | -6.8813e-07 | 5e-06 | 200Hz100m
500 | 500 | Hz | 100 | m | 0.1 | 100.2205 | 0.0015811 | 0.0013827 | 0.005 | m | 0.001 | 0.10022 | 1.5811e-06 | 1.3827e-06 | 5e-06 | 500Hz100m
1000 | 1 | kHz | 100 | m | 0.1 | 100.6612 | 0.0044721 | 0.024327 | 0.01 | m | 0.001 | 0.10066 | 4.4721e-06 | 2.4327e-05 | 1e-05 | 1kHz100m
2000 | 2 | kHz | 100 | m | 0.1 | 101.8538 | 0.012649 | 0.066749 | 0.02 | m | 0.001 | 0.10185 | 1.2649e-05 | 6.6749e-05 | 2e-05 | 2kHz100m
5000 | 5 | kHz | 100 | m | 0.1 | 107.2887 | 0.05 | 0.14313 | 0.05 | m | 0.001 | 0.10729 | 5e-05 | 0.00014313 | 5e-05 | 5kHz100m