## Measurement
This sheet uses correction data to fix bridge errors.
[5 empty] | Measured Z |  |  |  |  | Corrected Z (results) | [5 empty] | Validation data (debuging) |  |  | Deviation of XLS calculated from Octave calculated: |  |  | Measured Z (basic units) | [24 empty] | Interpolated correction |  |  |  | corrected Z
Range |  | f | f |  | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | Fixed Rs | Fixed U(Rs) | Fixed Xs | Fixed U(Xs) |  | Status | Valid Rs | Valid Xs |  | Dev Rs | Dev Xs |  | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | ¦Z¦ | lower Z ID | lower Z tag | higher Z tag | lower Z | higher Z | lower tag | higher tag | lower ID | higher ID | low check | low cor(Rs) | low U(cor(Rs)) | low cor(Xs) | low U(cor(Xs)) | high check | high cor(Rs) | high U(cor(Rs)) | high cor(Xs) | high U(cor(Xs)) | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | Rs | U(Rs) | Xs | U(Xs)
Ohm |  | Hz | - |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm |  | - | Ohm | Ohm |  | Ohm | Ohm |  | - | Ohm | Ohm | Ohm | Ohm | Ohm | - | - | - | Ohm | Ohm | - | - | - | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm
100 | m | 0.01 | 10 | mHz | -25.2001 | 0 | 22.8053 | 0 | m | -25.2 | 5.4964 | 22.8 | 7.5118 | m | OK | -25.2 | 22.8 | m | 0 | 0 | m | 0.001 | -0.0252 | 0 | 0.022805 | 0 | 0.033987 | 4 | 10m | 100m | 0.01 | 0.1 | 10mHz100m10m | 10mHz100m100m | 109 | 127 | 10mHz100m10m | 3.5728e-07 | 0.0037267 | -7.0448e-06 | 0.0093271 | 10mHz100m100m | -7.2913e-07 | 0.0054964 | -5.4688e-07 | 0.0025161 | 6.7728e-08 | 0.0054964 | -5.313e-06 | 0.0075118 | -0.0252 | 0.0054964 | 0.0228 | 0.0075118
10 | m | 0.2 | 200 | mHz | -4.8377 | 0 | -4.0113 | 0 | m | -4.84 | 0.39046 | -4.01 | 5.0042 | m | OK | -4.84 | -4.01 | m | 3.3622e-11 | 5.195e-12 | m | 0.001 | -0.0048377 | 0 | -0.0040113 | 0 | 0.0062844 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | 4.5087e-06 | 0.0058445 | 2.3197e-06 | 0.0066638 | 200mHz10m10m | -1.01e-05 | 0.00039046 | 6.262e-08 | 0.0031267 | -2.3458e-06 | 0.00039046 | 1.2607e-06 | 0.0050042 | -0.00484 | 0.00039046 | -0.00401 | 0.0050042
3 | m | 2000 | 2 | kHz | 1.8167 | 0 | -0.2484 | 0 | m | 1.82 | 2.7128 | -0.253 | 8.3693 | m | OK | 1.82 | -0.253 | m | -1.551e-11 | 5.8731e-13 | m | 0.001 | 0.0018167 | 0 | -0.0002484 | 0 | 0.0018336 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 1.9461e-06 | 0.0068581 | -4.5464e-06 | 0.0091974 | 2kHz3m3m | 5.1913e-06 | 0.0027128 | -4.6693e-06 | 0.0072106 | 3.2987e-06 | 0.0027128 | -4.5976e-06 | 0.0083693 | 0.00182 | 0.0027128 | -0.000253 | 0.0083693
10 | m | 200 | 200 | Hz | -8.0425 | 0 | -0.68228 | 0 | m | -8.04 | 3.3919 | -0.682 | 7.9987 | m | OK | -8.04 | -0.682 | m | 5.9508e-13 | 1.6098e-13 | m | 0.001 | -0.0080425 | 0 | -0.00068228 | 0 | 0.0080714 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -1.5114e-06 | 0.0067426 | -7.9441e-07 | 0.0040442 | 200Hz10m10m | 3.9822e-06 | 0.0033919 | 6.9205e-07 | 0.0095026 | 2.4686e-06 | 0.0033919 | 2.825e-07 | 0.0079987 | -0.00804 | 0.0033919 | -0.000682 | 0.0079987
3 | m | 0.1 | 100 | mHz | 0.31144 | 0 | 2.943 | 0 | m | 0.308 | 5.4921 | 2.94 | 4.7185 | m | OK | 0.308 | 2.94 | m | 8.7174e-12 | 5.925e-12 | m | 0.001 | 0.00031144 | 0 | 0.002943 | 0 | 0.0029594 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 2.9129e-06 | 0.0039352 | 1.3377e-06 | 9.2827e-05 | 100mHz3m3m | -3.5713e-06 | 0.0054921 | -3.0694e-06 | 0.0048143 | -3.4397e-06 | 0.0054921 | -2.98e-06 | 0.0047185 | 0.000308 | 0.0054921 | 0.00294 | 0.0047185
10 | m | 0.2 | 200 | mHz | -3.5829 | 0 | -1.1921 | 0 | m | -3.58 | 0.39046 | -1.19 | 6.2717 | m | OK | -3.58 | -1.19 | m | -2.5554e-11 | -3.948e-12 | m | 0.001 | -0.0035829 | 0 | -0.0011921 | 0 | 0.003776 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | 4.5087e-06 | 0.0058445 | 2.3197e-06 | 0.0066638 | 200mHz10m10m | -1.01e-05 | 0.00039046 | 6.262e-08 | 0.0031267 | 2.8892e-06 | 0.00039046 | 2.0695e-06 | 0.0062717 | -0.00358 | 0.00039046 | -0.00119 | 0.0062717
10 | m | 2 | 2 | Hz | 0.1028 | 0 | -8.4735 | 0 | m | 0.102 | 4.0675 | -8.48 | 5.4958 | m | OK | 0.102 | -8.48 | m | 9.4882e-14 | 2.6894e-12 | m | 0.001 | 0.0001028 | 0 | -0.0084735 | 0 | 0.0084741 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | -6.0987e-07 | 0.0079362 | -1.0353e-06 | 0.0039997 | 2Hz10m10m | -8.5536e-07 | 0.0040675 | -7.9933e-06 | 0.0059128 | -8.0185e-07 | 0.0040675 | -6.4766e-06 | 0.0054958 | 0.000102 | 0.0040675 | -0.00848 | 0.0054958
100 | m | 0.05 | 50 | mHz | -2.7198 | 0 | -60.9008 | 0 | m | -2.73 | 3.0998 | -60.9 | 4.4863 | m | OK | -2.73 | -60.9 | m | 0 | 0 | m | 0.001 | -0.0027198 | 0 | -0.060901 | 0 | 0.060962 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -2.9309e-06 | 0.0071846 | -6.5988e-07 | 0.0073072 | 50mHz100m100m | -1.5712e-05 | 0.0030998 | 1.9391e-06 | 0.0023254 | -1.0168e-05 | 0.0030998 | 8.1174e-07 | 0.0044863 | -0.00273 | 0.0030998 | -0.0609 | 0.0044863
1 | m | 0.2 | 200 | mHz | -0.58592 | 0 | -0.38735 | 0 | m | -0.586 | 2.7825 | -0.388 | 4.9537 | m | OK | -0.586 | -0.388 | m | -3.1186e-13 | 4.8472e-13 | m | 0.001 | -0.00058592 | 0 | -0.00038735 | 0 | 0.00070239 | 1 | 0m | 1m | 0 | 0.001 | 200mHz1m0m | 200mHz1m1m | 5 | 23 | 200mHz1m0m | -7.3927e-06 | 0.0017256 | 1.0729e-05 | 0.0084651 | 200mHz1m1m | 3.0245e-06 | 0.0027825 | -5.4646e-06 | 0.0034658 | -7.5749e-08 | 0.0027825 | -6.453e-07 | 0.0049537 | -0.000586 | 0.0027825 | -0.000388 | 0.0049537
100 | m | 50 | 50 | Hz | 58.6944 | 0 | -50.8971 | 0 | m | 58.7 | 10.0431 | -50.9 | 4.6878 | m | OK | 58.7 | -50.9 | m | 1.1369e-13 | 0 | m | 0.001 | 0.058694 | 0 | -0.050897 | 0 | 0.077689 | 4 | 10m | 100m | 0.01 | 0.1 | 50Hz100m10m | 50Hz100m100m | 120 | 138 | 50Hz100m10m | -6.2662e-06 | 0.0080157 | -1.0542e-07 | 0.0085401 | 50Hz100m100m | 9.4487e-06 | 0.010043 | -3.8231e-06 | 0.003418 | 5.5529e-06 | 0.010043 | -2.9015e-06 | 0.0046878 | 0.0587 | 0.010043 | -0.0509 | 0.0046878
1 | m | 50 | 50 | Hz | -0.55402 | 0 | -0.68347 | 0 | m | -0.554 | 0.31621 | -0.686 | 2.3194 | m | OK | -0.554 | -0.686 | m | -1.2401e-10 | 2.5977e-10 | m | 0.001 | -0.00055402 | 0 | -0.00068347 | 0 | 0.00087981 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | -3.4708e-06 | 0.00012295 | 4.7889e-06 | 0.00028219 | 50Hz1m1m | 5.01e-07 | 0.00031621 | -3.5309e-06 | 0.0025977 | 2.3643e-08 | 0.00031621 | -2.531e-06 | 0.0023194 | -0.000554 | 0.00031621 | -0.000686 | 0.0023194
3 | m | 100 | 100 | Hz | -1.01 | 0 | -1.2125 | 0 | m | -1.01 | 9.1079 | -1.21 | 1.9522 | m | OK | -1.01 | -1.21 | m | -8.8375e-11 | 7.7833e-12 | m | 0.001 | -0.00101 | 0 | -0.0012125 | 0 | 0.001578 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | 1.8965e-06 | 0.0049624 | 2.3088e-06 | 0.0015494 | 100Hz3m3m | -4.6761e-06 | 0.0091079 | 2.8876e-06 | 0.0029432 | -3.0814e-09 | 0.0091079 | 2.4761e-06 | 0.0019522 | -0.00101 | 0.0091079 | -0.00121 | 0.0019522
1 | m | 1000 | 1 | kHz | -0.78906 | 0 | 0.33888 | 0 | m | -0.786 | 9.6391 | 0.34 | 3.9058 | m | OK | -0.786 | 0.34 | m | 1.4494e-10 | -8.9664e-11 | m | 0.001 | -0.00078906 | 0 | 0.00033888 | 0 | 0.00085875 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 9.8775e-08 | 0.0082149 | 2.9458e-06 | 0.0038491 | 1kHz1m1m | 3.5429e-06 | 0.0096391 | 8.1518e-07 | 0.0039152 | 3.0564e-06 | 0.0096391 | 1.1161e-06 | 0.0039058 | -0.000786 | 0.0096391 | 0.00034 | 0.0039058
10 | m | 0.5 | 500 | mHz | -1.9222 | 0 | 5.2319 | 0 | m | -1.92 | 4.3624 | 5.23 | 2.5841 | m | OK | -1.92 | 5.23 | m | 3.5949e-13 | 9.5888e-12 | m | 0.001 | -0.0019222 | 0 | 0.0052319 | 0 | 0.0055738 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | 2.0303e-06 | 0.004134 | -5.0781e-06 | 0.00049429 | 500mHz10m10m | 2.3588e-06 | 0.0043624 | 3.6816e-06 | 0.006178 | 2.151e-06 | 0.0043624 | -1.8573e-06 | 0.0025841 | -0.00192 | 0.0043624 | 0.00523 | 0.0025841
10 | m | 5 | 5 | Hz | -6.4564 | 0 | 6.8282 | 0 | m | -6.46 | 8.5804 | 6.82 | 1.9607 | m | OK | -6.46 | 6.82 | m | -1.5099e-14 | -2.2204e-14 | m | 0.001 | -0.0064564 | 0 | 0.0068282 | 0 | 0.0093973 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -1.8062e-06 | 0.0017251 | -5.3584e-06 | 0.0015636 | 5Hz10m10m | -3.7766e-06 | 0.0085804 | -8.5028e-06 | 0.0019981 | -3.6069e-06 | 0.0085804 | -8.232e-06 | 0.0019607 | -0.00646 | 0.0085804 | 0.00682 | 0.0019607
10 | m | 200 | 200 | Hz | 4.0702 | 0 | -2.3996 | 0 | m | 4.07 | 3.3919 | -2.4 | 5.3892 | m | OK | 4.07 | -2.4 | m | 6.1817e-13 | 1.6698e-13 | m | 0.001 | 0.0040702 | 0 | -0.0023996 | 0 | 0.0047248 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | -1.5114e-06 | 0.0067426 | -7.9441e-07 | 0.0040442 | 200Hz10m10m | 3.9822e-06 | 0.0033919 | 6.9205e-07 | 0.0095026 | -1.5777e-07 | 0.0033919 | -4.2813e-07 | 0.0053892 | 0.00407 | 0.0033919 | -0.0024 | 0.0053892
1 | m | 5 | 5 | Hz | -0.10714 | 0 | -0.081382 | 0 | m | -0.108 | 6.7277 | -0.0872 | 3.3066 | m | OK | -0.108 | -0.0872 | m | 2.2658e-09 | 3.4164e-10 | m | 0.001 | -0.00010714 | 0 | -8.1382e-05 | 0 | 0.00013454 | 1 | 0m | 1m | 0 | 0.001 | 5Hz1m0m | 5Hz1m1m | 9 | 27 | 5Hz1m0m | 3.4079e-07 | 0.0079365 | -5.6364e-06 | 0.0035676 | 5Hz1m1m | -8.5912e-06 | 0.0067277 | -6.9831e-06 | 0.0016279 | -8.6094e-07 | 0.0067277 | -5.8176e-06 | 0.0033066 | -0.000108 | 0.0067277 | -8.72e-05 | 0.0033066
10 | m | 0.05 | 50 | mHz | 6.6464 | 0 | 3.6537 | 0 | m | 6.65 | 5.1427 | 3.65 | 4.6254 | m | OK | 6.65 | 3.65 | m | -9.1482e-14 | 4.1078e-13 | m | 0.001 | 0.0066464 | 0 | 0.0036537 | 0 | 0.0075845 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 2.1976e-06 | 0.0066352 | 2.4712e-06 | 0.0083113 | 50mHz10m10m | 4.3218e-06 | 0.0051427 | -7.0219e-06 | 0.0026834 | 3.5888e-06 | 0.0051427 | -3.7461e-06 | 0.0046254 | 0.00665 | 0.0051427 | 0.00365 | 0.0046254
3 | m | 2000 | 2 | kHz | 0.88063 | 0 | -1.6554 | 0 | m | 0.884 | 2.7128 | -1.66 | 8.3281 | m | OK | 0.884 | -1.66 | m | -2.4347e-12 | 9.2371e-14 | m | 0.001 | 0.00088063 | 0 | -0.0016554 | 0 | 0.0018751 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 1.9461e-06 | 0.0068581 | -4.5464e-06 | 0.0091974 | 2kHz3m3m | 5.1913e-06 | 0.0027128 | -4.6693e-06 | 0.0072106 | 3.366e-06 | 0.0027128 | -4.6001e-06 | 0.0083281 | 0.000884 | 0.0027128 | -0.00166 | 0.0083281
100 | m | 2000 | 2 | kHz | 33.9968 | 0 | 17.4933 | 0 | m | 34 | 8.0617 | 17.5 | 1.2337 | m | OK | 34 | 17.5 | m | 0 | 0 | m | 0.001 | 0.033997 | 0 | 0.017493 | 0 | 0.038233 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | 7.1951e-07 | 0.0044607 | 9.3729e-06 | 0.00046839 | 2kHz100m100m | 8.7809e-06 | 0.0080617 | 1.0075e-06 | 0.002908 | 3.2484e-06 | 0.0080617 | 6.7486e-06 | 0.0012337 | 0.034 | 0.0080617 | 0.0175 | 0.0012337
3 | m | 0.02 | 20 | mHz | -0.53103 | 0 | 1.5958 | 0 | m | -0.527 | 0.88313 | 1.59 | 4.9905 | m | OK | -0.527 | 1.59 | m | -4.6582e-10 | 2.1876e-10 | m | 0.001 | -0.00053103 | 0 | 0.0015958 | 0 | 0.0016819 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | 7.2004e-06 | 0.0048633 | -7.3096e-06 | 0.0045106 | 20mHz3m3m | -2.091e-06 | 0.00088313 | -2.9462e-06 | 0.0059184 | 4.0327e-06 | 0.00088313 | -5.822e-06 | 0.0049905 | -0.000527 | 0.00088313 | 0.00159 | 0.0049905
10 | m | 0.05 | 50 | mHz | -3.7424 | 0 | 0.82166 | 0 | m | -3.74 | 5.1427 | 0.823 | 7.6427 | m | OK | -3.74 | 0.823 | m | -2.1227e-13 | 9.4702e-13 | m | 0.001 | -0.0037424 | 0 | 0.00082166 | 0 | 0.0038316 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 2.1976e-06 | 0.0066352 | 2.4712e-06 | 0.0083113 | 50mHz10m10m | 4.3218e-06 | 0.0051427 | -7.0219e-06 | 0.0026834 | 2.4499e-06 | 0.0051427 | 1.3435e-06 | 0.0076427 | -0.00374 | 0.0051427 | 0.000823 | 0.0076427
3 | m | 0.05 | 50 | mHz | 2.3468 | 0 | -1.5254 | 0 | m | 2.35 | 5.9409 | -1.52 | 6.7809 | m | OK | 2.35 | -1.52 | m | -1.7262e-12 | -1.3582e-11 | m | 0.001 | 0.0023468 | 0 | -0.0015254 | 0 | 0.002799 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362 | 3.1764e-06 | 0.0059409 | 5.4057e-06 | 0.0067809 | 0.00235 | 0.0059409 | -0.00152 | 0.0067809
3 | m | 2000 | 2 | kHz | -1.465 | 0 | 2.5047 | 0 | m | -1.46 | 2.7128 | 2.5 | 7.3083 | m | OK | -1.46 | 2.5 | m | 9.6676e-12 | -3.6593e-13 | m | 0.001 | -0.001465 | 0 | 0.0025047 | 0 | 0.0029017 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 1.9461e-06 | 0.0068581 | -4.5464e-06 | 0.0091974 | 2kHz3m3m | 5.1913e-06 | 0.0027128 | -4.6693e-06 | 0.0072106 | 5.0317e-06 | 0.0027128 | -4.6632e-06 | 0.0073083 | -0.00146 | 0.0027128 | 0.0025 | 0.0073083
3 | m | 5000 | 5 | kHz | -1.6278 | 0 | 1.2123 | 0 | m | -1.62 | 3.6563 | 1.21 | 5.6564 | m | OK | -1.62 | 1.21 | m | -1.1769e-09 | 8.0892e-10 | m | 0.001 | -0.0016278 | 0 | 0.0012123 | 0 | 0.0020296 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | 1.1929e-05 | 0.0054327 | -5.075e-06 | 0.0063123 | 5kHz3m3m | 3.9732e-06 | 0.0036563 | 3.9338e-07 | 0.0050382 | 7.8333e-06 | 0.0036563 | -2.2598e-06 | 0.0056564 | -0.00162 | 0.0036563 | 0.00121 | 0.0056564
10 | m | 1 | 1 | Hz | 5.9623 | 0 | 0.09551 | 0 | m | 5.96 | 8.3131 | 0.0909 | 5.2815 | m | OK | 5.96 | 0.0909 | m | 0 | 3.0628e-14 | m | 0.001 | 0.0059623 | 0 | 9.551e-05 | 0 | 0.0059631 | 3 | 3m | 10m | 0.003 | 0.01 | 1Hz10m3m | 1Hz10m10m | 79 | 97 | 1Hz10m3m | -2.4705e-06 | 0.0093143 | -7.1774e-06 | 0.0087054 | 1Hz10m10m | -2.0327e-06 | 0.0083131 | -1.1117e-06 | 0.00061653 | -2.2852e-06 | 0.0083131 | -4.6098e-06 | 0.0052815 | 0.00596 | 0.0083131 | 9.09e-05 | 0.0052815
3 | m | 500 | 500 | Hz | -2.3681 | 0 | -0.80495 | 0 | m | -2.37 | 6.441 | -0.806 | 5.1783 | m | OK | -2.37 | -0.806 | m | 1.532e-10 | -1.4621e-10 | m | 0.001 | -0.0023681 | 0 | -0.00080495 | 0 | 0.0025012 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | 4.9361e-06 | 0.00012674 | -7.5628e-06 | 0.00051809 | 500Hz3m3m | -4.1615e-06 | 0.006441 | 1.1194e-06 | 0.0067268 | -1.8925e-06 | 0.006441 | -1.046e-06 | 0.0051783 | -0.00237 | 0.006441 | -0.000806 | 0.0051783
3 | m | 0.1 | 100 | mHz | 0.56527 | 0 | -0.89321 | 0 | m | 0.568 | 5.4921 | -0.892 | 0.22751 | m | OK | 0.568 | -0.892 | m | 2.387e-14 | 1.632e-14 | m | 0.001 | 0.00056527 | 0 | -0.00089321 | 0 | 0.0010571 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 2.9129e-06 | 0.0039352 | 1.3377e-06 | 9.2827e-05 | 100mHz3m3m | -3.5713e-06 | 0.0054921 | -3.0694e-06 | 0.0048143 | 2.7279e-06 | 0.0054921 | 1.212e-06 | 0.00022751 | 0.000568 | 0.0054921 | -0.000892 | 0.00022751
3 | m | 0.05 | 50 | mHz | 0.66643 | 0 | 2.1294 | 0 | m | 0.669 | 5.9409 | 2.13 | 7.1896 | m | OK | 0.669 | 2.13 | m | -4.109e-10 | -3.2334e-09 | m | 0.001 | 0.00066643 | 0 | 0.0021294 | 0 | 0.0022312 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362 | 2.5686e-06 | 0.0059409 | 6.2344e-07 | 0.0071896 | 0.000669 | 0.0059409 | 0.00213 | 0.0071896
3 | m | 0.1 | 100 | mHz | 0.94914 | 0 | 1.3301 | 0 | m | 0.95 | 5.4921 | 1.33 | 1.5895 | m | OK | 0.95 | 1.33 | m | 1.6717e-11 | 1.1362e-11 | m | 0.001 | 0.00094914 | 0 | 0.0013301 | 0 | 0.001634 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 2.9129e-06 | 0.0039352 | 1.3377e-06 | 9.2827e-05 | 100mHz3m3m | -3.5713e-06 | 0.0054921 | -3.0694e-06 | 0.0048143 | 8.5742e-07 | 0.0054921 | -5.9319e-08 | 0.0015895 | 0.00095 | 0.0054921 | 0.00133 | 0.0015895
1 | m | 50 | 50 | Hz | -0.61296 | 0 | -0.017687 | 0 | m | -0.614 | 0.31621 | -0.018 | 1.7021 | m | OK | -0.614 | -0.018 | m | -8.4423e-11 | 1.7685e-10 | m | 0.001 | -0.00061296 | 0 | -1.7687e-05 | 0 | 0.00061322 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | -3.4708e-06 | 0.00012295 | 4.7889e-06 | 0.00028219 | 50Hz1m1m | 5.01e-07 | 0.00031621 | -3.5309e-06 | 0.0025977 | -1.0352e-06 | 0.00031621 | -3.1297e-07 | 0.0017021 | -0.000614 | 0.00031621 | -1.8e-05 | 0.0017021
1 | m | 2000 | 2 | kHz | 0.88778 | 0 | 0.21789 | 0 | m | 0.885 | 4.9096 | 0.216 | 8.4043 | m | OK | 0.885 | 0.216 | m | 3.1014e-11 | 1.3171e-11 | m | 0.001 | 0.00088778 | 0 | 0.00021789 | 0 | 0.00091413 | 1 | 0m | 1m | 0 | 0.001 | 2kHz1m0m | 2kHz1m1m | 17 | 35 | 2kHz1m0m | -4.7259e-06 | 0.0088746 | -2.7211e-06 | 0.0078804 | 2kHz1m1m | -2.5952e-06 | 0.0049096 | -1.8163e-06 | 0.0084535 | -2.7781e-06 | 0.0049096 | -1.894e-06 | 0.0084043 | 0.000885 | 0.0049096 | 0.000216 | 0.0084043
3 | m | 2 | 2 | Hz | 2.1904 | 0 | 0.68164 | 0 | m | 2.19 | 7.3121 | 0.687 | 5.7684 | m | OK | 2.19 | 0.687 | m | -7.5886e-12 | 1.4075e-11 | m | 0.001 | 0.0021904 | 0 | 0.00068164 | 0 | 0.002294 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | -3.0351e-06 | 0.0044469 | 1.032e-05 | 0.0019628 | 2Hz3m3m | 1.1014e-06 | 0.0073121 | 2.6483e-06 | 0.0078449 | -3.5887e-07 | 0.0073121 | 5.3566e-06 | 0.0057684 | 0.00219 | 0.0073121 | 0.000687 | 0.0057684
1 | m | 50 | 50 | Hz | -0.29523 | 0 | 0.31078 | 0 | m | -0.297 | 0.31621 | 0.312 | 1.2747 | m | OK | -0.297 | 0.312 | m | -5.5132e-10 | 1.1549e-09 | m | 0.001 | -0.00029523 | 0 | 0.00031078 | 0 | 0.00042865 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | -3.4708e-06 | 0.00012295 | 4.7889e-06 | 0.00028219 | 50Hz1m1m | 5.01e-07 | 0.00031621 | -3.5309e-06 | 0.0025977 | -1.7683e-06 | 0.00031621 | 1.2226e-06 | 0.0012747 | -0.000297 | 0.00031621 | 0.000312 | 0.0012747
10 | m | 5000 | 5 | kHz | -4.7071 | 0 | -1.1951 | 0 | m | -4.71 | 6.4858 | -1.2 | 6.5206 | m | OK | -4.71 | -1.2 | m | 4.1744e-13 | 1.6067e-12 | m | 0.001 | -0.0047071 | 0 | -0.0011951 | 0 | 0.0048565 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | -2.3167e-06 | 0.0060027 | -2.8645e-06 | 0.0055319 | 5kHz10m10m | -4.3437e-06 | 0.0064858 | -1.0677e-05 | 0.0092597 | -2.8543e-06 | 0.0064858 | -4.9365e-06 | 0.0065206 | -0.00471 | 0.0064858 | -0.0012 | 0.0065206
100 | m | 2 | 2 | Hz | -20.6133 | 0 | 54.5988 | 0 | m | -20.6 | 10.2754 | 54.6 | 3.9741 | m | OK | -20.6 | 54.6 | m | 0 | 5.6843e-14 | m | 0.001 | -0.020613 | 0 | 0.054599 | 0 | 0.05836 | 4 | 10m | 100m | 0.01 | 0.1 | 2Hz100m10m | 2Hz100m100m | 116 | 134 | 2Hz100m10m | 1.0865e-05 | 0.0083693 | -4.6751e-06 | 0.0076711 | 2Hz100m100m | 1.5456e-05 | 0.010275 | 6.332e-06 | 0.00079084 | 1.3332e-05 | 0.010275 | 1.2395e-06 | 0.0039741 | -0.0206 | 0.010275 | 0.0546 | 0.0039741
1 | m | 50 | 50 | Hz | -0.64047 | 0 | 0.37137 | 0 | m | -0.641 | 0.31621 | 0.37 | 1.9964 | m | OK | -0.641 | 0.37 | m | -5.3473e-11 | 1.1201e-10 | m | 0.001 | -0.00064047 | 0 | 0.00037137 | 0 | 0.00074035 | 1 | 0m | 1m | 0 | 0.001 | 50Hz1m0m | 50Hz1m1m | 12 | 30 | 50Hz1m0m | -3.4708e-06 | 0.00012295 | 4.7889e-06 | 0.00028219 | 50Hz1m1m | 5.01e-07 | 0.00031621 | -3.5309e-06 | 0.0025977 | -5.3028e-07 | 0.00031621 | -1.3707e-06 | 0.0019964 | -0.000641 | 0.00031621 | 0.00037 | 0.0019964
3 | m | 2 | 2 | Hz | -1.4891 | 0 | -1.3864 | 0 | m | -1.49 | 7.3121 | -1.38 | 5.0055 | m | OK | -1.49 | -1.38 | m | 8.6087e-12 | -1.5966e-11 | m | 0.001 | -0.0014891 | 0 | -0.0013864 | 0 | 0.0020346 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | -3.0351e-06 | 0.0044469 | 1.032e-05 | 0.0019628 | 2Hz3m3m | 1.1014e-06 | 0.0073121 | 2.6483e-06 | 0.0078449 | -8.9541e-07 | 0.0073121 | 6.3517e-06 | 0.0050055 | -0.00149 | 0.0073121 | -0.00138 | 0.0050055
1 | m | 0.02 | 20 | mHz | 0.47194 | 0 | -0.049101 | 0 | m | 0.472 | 7.7127 | -0.0523 | 3.9313 | m | OK | 0.472 | -0.0523 | m | -8.9365e-11 | 1.7084e-11 | m | 0.001 | 0.00047194 | 0 | -4.9101e-05 | 0 | 0.00047449 | 1 | 0m | 1m | 0 | 0.001 | 20mHz1m0m | 20mHz1m1m | 2 | 20 | 20mHz1m0m | -2.5644e-06 | 0.0049929 | -2.6984e-06 | 0.0056488 | 20mHz1m1m | 2.9575e-06 | 0.0077127 | -3.754e-06 | 0.0020291 | 5.5723e-08 | 0.0077127 | -3.1993e-06 | 0.0039313 | 0.000472 | 0.0077127 | -5.23e-05 | 0.0039313
3 | m | 1 | 1 | Hz | 1.5663 | 0 | 0.29849 | 0 | m | 1.57 | 7.8306 | 0.297 | 6.1655 | m | OK | 1.57 | 0.297 | m | 3.9037e-11 | -2.0886e-11 | m | 0.001 | 0.0015663 | 0 | 0.00029849 | 0 | 0.0015944 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | 7.9524e-06 | 0.0018873 | -3.7396e-06 | 0.0066752 | 1Hz3m3m | -6.1901e-06 | 0.0078306 | 3.8269e-06 | 0.0049602 | 3.749e-06 | 0.0078306 | -1.4907e-06 | 0.0061655 | 0.00157 | 0.0078306 | 0.000297 | 0.0061655
3 | m | 500 | 500 | Hz | -2.1785 | 0 | 1.0314 | 0 | m | -2.18 | 6.441 | 1.03 | 4.8964 | m | OK | -2.18 | 1.03 | m | 8.7833e-10 | -8.3822e-10 | m | 0.001 | -0.0021785 | 0 | 0.0010314 | 0 | 0.0024104 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | 4.9361e-06 | 0.00012674 | -7.5628e-06 | 0.00051809 | 500Hz3m3m | -4.1615e-06 | 0.006441 | 1.1194e-06 | 0.0067268 | -1.4794e-06 | 0.006441 | -1.4403e-06 | 0.0048964 | -0.00218 | 0.006441 | 0.00103 | 0.0048964
3 | m | 5000 | 5 | kHz | -0.012159 | 0 | 1.0948 | 0 | m | -0.000607 | 3.6563 | 1.09 | 6.2519 | m | OK | -0.000607 | 1.09 | m | -2.7691e-10 | 1.9033e-10 | m | 0.001 | -1.2159e-05 | 0 | 0.0010948 | 0 | 0.0010949 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | 1.1929e-05 | 0.0054327 | -5.075e-06 | 0.0063123 | 5kHz3m3m | 3.9732e-06 | 0.0036563 | 3.9338e-07 | 0.0050382 | 1.1552e-05 | 0.0036563 | -4.8156e-06 | 0.0062519 | -6.07e-07 | 0.0036563 | 0.00109 | 0.0062519
10 | m | 0.05 | 50 | mHz | 2.1875 | 0 | 3.2488 | 0 | m | 2.19 | 5.1427 | 3.25 | 7.5743 | m | OK | 2.19 | 3.25 | m | -2.6734e-13 | 1.195e-12 | m | 0.001 | 0.0021875 | 0 | 0.0032488 | 0 | 0.0039166 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 2.1976e-06 | 0.0066352 | 2.4712e-06 | 0.0083113 | 50mHz10m10m | 4.3218e-06 | 0.0051427 | -7.0219e-06 | 0.0026834 | 2.4757e-06 | 0.0051427 | 1.2282e-06 | 0.0075743 | 0.00219 | 0.0051427 | 0.00325 | 0.0075743
10 | m | 0.01 | 10 | mHz | 1.2134 | 0 | -3.8045 | 0 | m | 1.22 | 6.2631 | -3.81 | 6.645 | m | OK | 1.22 | -3.81 | m | 2.7311e-14 | 1.199e-14 | m | 0.001 | 0.0012134 | 0 | -0.0038045 | 0 | 0.0039934 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | 7.7028e-06 | 0.0098114 | -4.9888e-06 | 0.0070397 | 10mHz10m10m | -2.8151e-07 | 0.0062631 | -8.333e-06 | 0.0042586 | 6.5698e-06 | 0.0062631 | -5.4633e-06 | 0.006645 | 0.00122 | 0.0062631 | -0.00381 | 0.006645
10 | m | 10 | 10 | Hz | 4.1843 | 0 | 3.2065 | 0 | m | 4.18 | 3.8182 | 3.21 | 5.3714 | m | OK | 4.18 | 3.21 | m | 0 | 0 | m | 0.001 | 0.0041843 | 0 | 0.0032065 | 0 | 0.0052716 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -4.1407e-06 | 0.0068863 | 3.2695e-06 | 0.003358 | 10Hz10m10m | -4.7765e-06 | 0.0038182 | 4.0862e-06 | 0.0095625 | -4.3471e-06 | 0.0038182 | 3.5346e-06 | 0.0053714 | 0.00418 | 0.0038182 | 0.00321 | 0.0053714
100 | m | 100 | 100 | Hz | -24.5031 | 0 | -30.9967 | 0 | m | -24.5 | 10.9725 | -31 | 6.1226 | m | OK | -24.5 | -31 | m | 0 | 0 | m | 0.001 | -0.024503 | 0 | -0.030997 | 0 | 0.039512 | 4 | 10m | 100m | 0.01 | 0.1 | 100Hz100m10m | 100Hz100m100m | 121 | 139 | 100Hz100m10m | 2.5938e-06 | 0.00063357 | -3.9209e-06 | 0.0056125 | 100Hz100m100m | 4.0011e-06 | 0.010973 | -2.0418e-06 | 0.007168 | 3.0553e-06 | 0.010973 | -3.3047e-06 | 0.0061226 | -0.0245 | 0.010973 | -0.031 | 0.0061226
10 | m | 500 | 500 | Hz | -5.3238 | 0 | -2.7368 | 0 | m | -5.32 | 5.1253 | -2.73 | 3.4228 | m | OK | -5.32 | -2.73 | m | -1.4238e-12 | -7.4341e-13 | m | 0.001 | -0.0053238 | 0 | -0.0027368 | 0 | 0.0059861 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 5.1494e-06 | 0.0078664 | 7.5307e-06 | 0.00054201 | 500Hz10m10m | 2.0292e-06 | 0.0051253 | 5.9026e-06 | 0.0072951 | 3.8184e-06 | 0.0051253 | 6.8362e-06 | 0.0034228 | -0.00532 | 0.0051253 | -0.00273 | 0.0034228
10 | m | 0.1 | 100 | mHz | -4.6477 | 0 | 7.5407 | 0 | m | -4.65 | 5.9461 | 7.54 | 8.1212 | m | OK | -4.65 | 7.54 | m | -2.5668e-13 | 2.351e-12 | m | 0.001 | -0.0046477 | 0 | 0.0075407 | 0 | 0.0088579 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | -2.9039e-06 | 0.0077852 | 4.8302e-06 | 0.0025706 | 100mHz10m10m | -2.1867e-06 | 0.0059461 | -1.7372e-06 | 0.0092034 | -2.3037e-06 | 0.0059461 | -6.6568e-07 | 0.0081212 | -0.00465 | 0.0059461 | 0.00754 | 0.0081212
100 | m | 0.05 | 50 | mHz | 37.6155 | 0 | 91.2981 | 0 | m | 37.6 | 3.0998 | 91.3 | 2.3949 | m | OK | 37.6 | 91.3 | m | 0 | 0 | m | 0.001 | 0.037616 | 0 | 0.091298 | 0 | 0.098743 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -2.9309e-06 | 0.0071846 | -6.5988e-07 | 0.0073072 | 50mHz100m100m | -1.5712e-05 | 0.0030998 | 1.9391e-06 | 0.0023254 | -1.5533e-05 | 0.0030998 | 1.9028e-06 | 0.0023949 | 0.0376 | 0.0030998 | 0.0913 | 0.0023949
3 | m | 0.05 | 50 | mHz | 1.3174 | 0 | 1.8093 | 0 | m | 1.32 | 5.9409 | 1.81 | 7.1846 | m | OK | 1.32 | 1.81 | m | -3.8018e-10 | -2.9917e-09 | m | 0.001 | 0.0013174 | 0 | 0.0018093 | 0 | 0.0022381 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362 | 2.576e-06 | 0.0059409 | 6.816e-07 | 0.0071846 | 0.00132 | 0.0059409 | 0.00181 | 0.0071846
1 | m | 20 | 20 | Hz | -0.17709 | 0 | 0.3608 | 0 | m | -0.173 | 8.0799 | 0.357 | 4.8794 | m | OK | -0.173 | 0.357 | m | 1.1744e-09 | 3.4074e-09 | m | 0.001 | -0.00017709 | 0 | 0.0003608 | 0 | 0.00040191 | 1 | 0m | 1m | 0 | 0.001 | 20Hz1m0m | 20Hz1m1m | 11 | 29 | 20Hz1m0m | 2.6642e-06 | 0.0087438 | -7.9212e-06 | 0.0040584 | 20Hz1m1m | 6.2024e-06 | 0.0080799 | 2.3448e-06 | 0.0061013 | 4.0863e-06 | 0.0080799 | -3.7952e-06 | 0.0048794 | -0.000173 | 0.0080799 | 0.000357 | 0.0048794
100 | m | 500 | 500 | Hz | -1.8767 | 0 | -55.4977 | 0 | m | -1.88 | 8.9432 | -55.5 | 7.7151 | m | OK | -1.88 | -55.5 | m | 6.4393e-15 | 0 | m | 0.001 | -0.0018767 | 0 | -0.055498 | 0 | 0.055529 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | -1.8758e-06 | 0.0098276 | 2.6486e-07 | 0.0081162 | 500Hz100m100m | -4.7634e-06 | 0.0089432 | -4.7152e-06 | 0.0073233 | -3.3366e-06 | 0.0089432 | -2.2545e-06 | 0.0077151 | -0.00188 | 0.0089432 | -0.0555 | 0.0077151
100 | m | 1 | 1 | Hz | -16.2021 | 0 | 70.1956 | 0 | m | -16.2 | 3.3847 | 70.2 | 5.4191 | m | OK | -16.2 | 70.2 | m | 0 | 0 | m | 0.001 | -0.016202 | 0 | 0.070196 | 0 | 0.072041 | 4 | 10m | 100m | 0.01 | 0.1 | 1Hz100m10m | 1Hz100m100m | 115 | 133 | 1Hz100m10m | 2.3046e-06 | 0.00054254 | 5.8148e-07 | 0.0029675 | 1Hz100m100m | 2.0044e-06 | 0.0033847 | 6.1626e-06 | 0.0065239 | 2.0977e-06 | 0.0033847 | 4.4288e-06 | 0.0054191 | -0.0162 | 0.0033847 | 0.0702 | 0.0054191
3 | m | 50 | 50 | Hz | -0.64833 | 0 | -1.5805 | 0 | m | -0.646 | 9.9373 | -1.58 | 6.5546 | m | OK | -0.646 | -1.58 | m | 3.2317e-12 | -7.0832e-14 | m | 0.001 | -0.00064833 | 0 | -0.0015805 | 0 | 0.0017083 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | 4.1395e-06 | 0.0044916 | 4.5172e-07 | 0.0056302 | 50Hz3m3m | -9.748e-07 | 0.0099373 | 5.6373e-07 | 0.0082406 | 2.3283e-06 | 0.0099373 | 4.9139e-07 | 0.0065546 | -0.000646 | 0.0099373 | -0.00158 | 0.0065546
100 | m | 1 | 1 | Hz | -83.2021 | 0 | 4.0049 | 0 | m | -83.2 | 3.3847 | 4.01 | 5.8639 | m | OK | -83.2 | 4.01 | m | 0 | 0 | m | 0.001 | -0.083202 | 0 | 0.0040049 | 0 | 0.083298 | 4 | 10m | 100m | 0.01 | 0.1 | 1Hz100m10m | 1Hz100m100m | 115 | 133 | 1Hz100m10m | 2.3046e-06 | 0.00054254 | 5.8148e-07 | 0.0029675 | 1Hz100m100m | 2.0044e-06 | 0.0033847 | 6.1626e-06 | 0.0065239 | 2.0601e-06 | 0.0033847 | 5.1269e-06 | 0.0058639 | -0.0832 | 0.0033847 | 0.00401 | 0.0058639
1 | m | 0.05 | 50 | mHz | -0.37242 | 0 | 0.5506 | 0 | m | -0.371 | 9.5046 | 0.557 | 6.0052 | m | OK | -0.371 | 0.557 | m | -1.1258e-09 | -4.0034e-10 | m | 0.001 | -0.00037242 | 0 | 0.0005506 | 0 | 0.00066473 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | -9.0717e-06 | 0.006266 | 2.6659e-06 | 0.0017195 | 50mHz1m1m | 6.7167e-06 | 0.0095046 | 8.2803e-06 | 0.0081667 | 1.4233e-06 | 0.0095046 | 6.398e-06 | 0.0060052 | -0.000371 | 0.0095046 | 0.000557 | 0.0060052
100 | m | 2000 | 2 | kHz | 33.3968 | 0 | -18.3068 | 0 | m | 33.4 | 8.0617 | -18.3 | 1.2297 | m | OK | 33.4 | -18.3 | m | 0 | 2.8422e-14 | m | 0.001 | 0.033397 | 0 | -0.018307 | 0 | 0.038085 | 4 | 10m | 100m | 0.01 | 0.1 | 2kHz100m10m | 2kHz100m100m | 125 | 143 | 2kHz100m10m | 7.1951e-07 | 0.0044607 | 9.3729e-06 | 0.00046839 | 2kHz100m100m | 8.7809e-06 | 0.0080617 | 1.0075e-06 | 0.002908 | 3.2351e-06 | 0.0080617 | 6.7624e-06 | 0.0012297 | 0.0334 | 0.0080617 | -0.0183 | 0.0012297
10 | m | 0.05 | 50 | mHz | -1.4128 | 0 | -4.8397 | 0 | m | -1.41 | 5.1427 | -4.84 | 6.6698 | m | OK | -1.41 | -4.84 | m | -1.4626e-12 | 6.5361e-12 | m | 0.001 | -0.0014128 | 0 | -0.0048397 | 0 | 0.0050417 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | 2.1976e-06 | 0.0066352 | 2.4712e-06 | 0.0083113 | 50mHz10m10m | 4.3218e-06 | 0.0051427 | -7.0219e-06 | 0.0026834 | 2.8171e-06 | 0.0051427 | -2.9766e-07 | 0.0066698 | -0.00141 | 0.0051427 | -0.00484 | 0.0066698
10 | m | 500 | 500 | Hz | -0.29882 | 0 | 8.2237 | 0 | m | -0.296 | 5.1253 | 8.23 | 5.5866 | m | OK | -0.296 | 8.23 | m | 4.6241e-14 | 2.4869e-14 | m | 0.001 | -0.00029882 | 0 | 0.0082237 | 0 | 0.0082291 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 5.1494e-06 | 0.0078664 | 7.5307e-06 | 0.00054201 | 500Hz10m10m | 2.0292e-06 | 0.0051253 | 5.9026e-06 | 0.0072951 | 2.8186e-06 | 0.0051253 | 6.3145e-06 | 0.0055866 | -0.000296 | 0.0051253 | 0.00823 | 0.0055866
10 | m | 500 | 500 | Hz | -9.8021 | 0 | 0.53205 | 0 | m | -9.8 | 5.1253 | 0.538 | 7.1181 | m | OK | -9.8 | 0.538 | m | -1.35e-13 | -7.0166e-14 | m | 0.001 | -0.0098021 | 0 | 0.00053205 | 0 | 0.0098165 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 5.1494e-06 | 0.0078664 | 7.5307e-06 | 0.00054201 | 500Hz10m10m | 2.0292e-06 | 0.0051253 | 5.9026e-06 | 0.0072951 | 2.111e-06 | 0.0051253 | 5.9453e-06 | 0.0071181 | -0.0098 | 0.0051253 | 0.000538 | 0.0071181
3 | m | 1000 | 1 | kHz | 1.0773 | 0 | -0.93372 | 0 | m | 1.07 | 7.0375 | -0.924 | 3.0361 | m | OK | 1.07 | -0.924 | m | 1.0443e-09 | -1.032e-09 | m | 0.001 | 0.0010773 | 0 | -0.00093372 | 0 | 0.0014256 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | -8.9779e-06 | 0.00073379 | 1.1372e-05 | 0.0023675 | 1kHz3m3m | -1.1055e-06 | 0.0070375 | 3.593e-06 | 0.0055089 | -7.3025e-06 | 0.0070375 | 9.7167e-06 | 0.0030361 | 0.00107 | 0.0070375 | -0.000924 | 0.0030361
100 | m | 0.05 | 50 | mHz | 4.9849 | 0 | -23.5997 | 0 | m | 4.98 | 3.0998 | -23.6 | 6.5256 | m | OK | 4.98 | -23.6 | m | 0 | 0 | m | 0.001 | 0.0049849 | 0 | -0.0236 | 0 | 0.02412 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -2.9309e-06 | 0.0071846 | -6.5988e-07 | 0.0073072 | 50mHz100m100m | -1.5712e-05 | 0.0030998 | 1.9391e-06 | 0.0023254 | -4.9361e-06 | 0.0030998 | -2.5212e-07 | 0.0065256 | 0.00498 | 0.0030998 | -0.0236 | 0.0065256
10 | m | 5000 | 5 | kHz | 0.17543 | 0 | -3.3967 | 0 | m | 0.173 | 6.4858 | -3.4 | 5.7456 | m | OK | 0.173 | -3.4 | m | 4.81e-14 | 1.8519e-13 | m | 0.001 | 0.00017543 | 0 | -0.0033967 | 0 | 0.0034012 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | -2.3167e-06 | 0.0060027 | -2.8645e-06 | 0.0055319 | 5kHz10m10m | -4.3437e-06 | 0.0064858 | -1.0677e-05 | 0.0092597 | -2.4329e-06 | 0.0064858 | -3.3123e-06 | 0.0057456 | 0.000173 | 0.0064858 | -0.0034 | 0.0057456
3 | m | 1000 | 1 | kHz | -1.0817 | 0 | -0.44571 | 0 | m | -1.09 | 7.0375 | -0.435 | 2.6344 | m | OK | -1.09 | -0.435 | m | -7.7439e-11 | 7.6522e-11 | m | 0.001 | -0.0010817 | 0 | -0.00044571 | 0 | 0.0011699 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | -8.9779e-06 | 0.00073379 | 1.1372e-05 | 0.0023675 | 1kHz3m3m | -1.1055e-06 | 0.0070375 | 3.593e-06 | 0.0055089 | -8.309e-06 | 0.0070375 | 1.0711e-05 | 0.0026344 | -0.00109 | 0.0070375 | -0.000435 | 0.0026344
100 | m | 5 | 5 | Hz | -1.456 | 0 | 10.5973 | 0 | m | -1.46 | 7.0843 | 10.6 | 8.5052 | m | OK | -1.46 | 10.6 | m | 0 | 0 | m | 0.001 | -0.001456 | 0 | 0.010597 | 0 | 0.010697 | 4 | 10m | 100m | 0.01 | 0.1 | 5Hz100m10m | 5Hz100m100m | 117 | 135 | 5Hz100m10m | -4.0641e-06 | 0.009757 | 2.7582e-06 | 0.0085653 | 5Hz100m100m | 6.3364e-07 | 0.0070843 | 7.604e-07 | 0.00079673 | -4.0277e-06 | 0.0070843 | 2.7427e-06 | 0.0085052 | -0.00146 | 0.0070843 | 0.0106 | 0.0085052
10 | m | 2000 | 2 | kHz | -0.50887 | 0 | -4.0868 | 0 | m | -0.515 | 6.4033 | -4.09 | 2.6835 | m | OK | -0.515 | -4.09 | m | -2.2704e-13 | -1.2701e-13 | m | 0.001 | -0.00050887 | 0 | -0.0040868 | 0 | 0.0041183 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -6.5665e-06 | 0.0016504 | -3.4596e-06 | 0.0028854 | 2kHz10m10m | -3.8097e-06 | 0.0064033 | -1.9173e-06 | 0.0016218 | -6.1261e-06 | 0.0064033 | -3.2132e-06 | 0.0026835 | -0.000515 | 0.0064033 | -0.00409 | 0.0026835
3 | m | 50 | 50 | Hz | -2.6396 | 0 | 0.88245 | 0 | m | -2.64 | 9.9373 | 0.883 | 7.9576 | m | OK | -2.64 | 0.883 | m | 5.4138e-11 | -1.1857e-12 | m | 0.001 | -0.0026396 | 0 | 0.00088245 | 0 | 0.0027832 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | 4.1395e-06 | 0.0044916 | 4.5172e-07 | 0.0056302 | 50Hz3m3m | -9.748e-07 | 0.0099373 | 5.6373e-07 | 0.0082406 | -4.2036e-07 | 0.0099373 | 5.5159e-07 | 0.0079576 | -0.00264 | 0.0099373 | 0.000883 | 0.0079576
10 | m | 0.5 | 500 | mHz | -1.1721 | 0 | 4.1934 | 0 | m | -1.17 | 4.3624 | 4.19 | 1.5938 | m | OK | -1.17 | 4.19 | m | 2.458e-13 | 6.5565e-12 | m | 0.001 | -0.0011721 | 0 | 0.0041934 | 0 | 0.0043541 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | 2.0303e-06 | 0.004134 | -5.0781e-06 | 0.00049429 | 500mHz10m10m | 2.3588e-06 | 0.0043624 | 3.6816e-06 | 0.006178 | 2.0938e-06 | 0.0043624 | -3.3836e-06 | 0.0015938 | -0.00117 | 0.0043624 | 0.00419 | 0.0015938
10 | m | 2000 | 2 | kHz | 6.6151 | 0 | 1.4726 | 0 | m | 6.61 | 6.4033 | 1.47 | 2.2036 | m | OK | 6.61 | 1.47 | m | 6.0929e-13 | 3.4039e-13 | m | 0.001 | 0.0066151 | 0 | 0.0014726 | 0 | 0.006777 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | -6.5665e-06 | 0.0016504 | -3.4596e-06 | 0.0028854 | 2kHz10m10m | -3.8097e-06 | 0.0064033 | -1.9173e-06 | 0.0016218 | -5.079e-06 | 0.0064033 | -2.6274e-06 | 0.0022036 | 0.00661 | 0.0064033 | 0.00147 | 0.0022036
10 | m | 100 | 100 | Hz | 8.5058 | 0 | 2.0111 | 0 | m | 8.51 | 8.1937 | 2.01 | 3.1229 | m | OK | 8.51 | 2.01 | m | 0 | 0 | m | 0.001 | 0.0085058 | 0 | 0.0020111 | 0 | 0.0087404 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 1.2373e-06 | 0.0093016 | 9.7094e-06 | 0.0033799 | 100Hz10m10m | 4.8106e-06 | 0.0081937 | -3.4914e-06 | 0.0030666 | 4.1675e-06 | 0.0081937 | -1.116e-06 | 0.0031229 | 0.00851 | 0.0081937 | 0.00201 | 0.0031229
1 | m | 0.05 | 50 | mHz | 0.36014 | 0 | -0.35049 | 0 | m | 0.359 | 9.5046 | -0.345 | 4.9595 | m | OK | 0.359 | -0.345 | m | 3.932e-09 | 1.3982e-09 | m | 0.001 | 0.00036014 | 0 | -0.00035049 | 0 | 0.00050253 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | -9.0717e-06 | 0.006266 | 2.6659e-06 | 0.0017195 | 50mHz1m1m | 6.7167e-06 | 0.0095046 | 8.2803e-06 | 0.0081667 | -1.1375e-06 | 0.0095046 | 5.4873e-06 | 0.0049595 | 0.000359 | 0.0095046 | -0.000345 | 0.0049595
3 | m | 0.05 | 50 | mHz | -1.1414 | 0 | 0.0048259 | 0 | m | -1.14 | 5.9409 | -0.00373 | 7.9742 | m | OK | -1.14 | -0.00373 | m | -2.1654e-12 | -1.704e-11 | m | 0.001 | -0.0011414 | 0 | 4.8259e-06 | 0 | 0.0011414 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362 | 1.4021e-06 | 0.0059409 | -8.5559e-06 | 0.0079742 | -0.00114 | 0.0059409 | -3.73e-06 | 0.0079742
10 | m | 1000 | 1 | kHz | -4.5091 | 0 | -3.6705 | 0 | m | -4.51 | 8.8571 | -3.67 | 6.9948 | m | OK | -4.51 | -3.67 | m | 4.8956e-12 | 1.136e-12 | m | 0.001 | -0.0045091 | 0 | -0.0036705 | 0 | 0.0058141 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | 1.9599e-06 | 0.0028237 | 1.1734e-06 | 0.0083582 | 1kHz10m10m | -5.2726e-06 | 0.0088571 | -5.0497e-07 | 0.0049668 | -9.4771e-07 | 0.0088571 | 4.9868e-07 | 0.0069948 | -0.00451 | 0.0088571 | -0.00367 | 0.0069948
1 | m | 10 | 10 | Hz | -0.14624 | 0 | 0.50974 | 0 | m | -0.149 | 3.6966 | 0.512 | 4.0579 | m | OK | -0.149 | 0.512 | m | -8.4677e-10 | 7.9745e-10 | m | 0.001 | -0.00014624 | 0 | 0.00050974 | 0 | 0.00053031 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | -5.9086e-06 | 0.0073162 | 5.218e-06 | 0.00010656 | 10Hz1m1m | 2.3121e-08 | 0.0036966 | -3.6821e-07 | 0.0075576 | -2.763e-06 | 0.0036966 | 2.2556e-06 | 0.0040579 | -0.000149 | 0.0036966 | 0.000512 | 0.0040579
10 | m | 2 | 2 | Hz | 0.52475 | 0 | -6.8551 | 0 | m | 0.524 | 4.0675 | -6.86 | 5.0588 | m | OK | 0.524 | -6.86 | m | 2.4913e-13 | 7.0619e-12 | m | 0.001 | 0.00052475 | 0 | -0.0068551 | 0 | 0.0068752 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | -6.0987e-07 | 0.0079362 | -1.0353e-06 | 0.0039997 | 2Hz10m10m | -8.5536e-07 | 0.0040675 | -7.9933e-06 | 0.0059128 | -7.4577e-07 | 0.0040675 | -4.8872e-06 | 0.0050588 | 0.000524 | 0.0040675 | -0.00686 | 0.0050588
3 | m | 0.05 | 50 | mHz | -2.3828 | 0 | -0.39517 | 0 | m | -2.38 | 5.9409 | -0.393 | 7.0571 | m | OK | -2.38 | -0.393 | m | 6.6462e-12 | 5.23e-11 | m | 0.001 | -0.0023828 | 0 | -0.00039517 | 0 | 0.0024153 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362 | 2.7657e-06 | 0.0059409 | 2.174e-06 | 0.0070571 | -0.00238 | 0.0059409 | -0.000393 | 0.0070571
10 | m | 50 | 50 | Hz | -3.6083 | 0 | 1.819 | 0 | m | -3.61 | 2.8821 | 1.81 | 4.8669 | m | OK | -3.61 | 1.81 | m | -2.2915e-13 | -1.013e-12 | m | 0.001 | -0.0036083 | 0 | 0.001819 | 0 | 0.0040408 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | -2.328e-06 | 0.0018291 | -1.1645e-05 | 0.0055758 | 50Hz10m10m | 1.7163e-06 | 0.0028821 | 6.2636e-06 | 0.00080787 | -1.7266e-06 | 0.0028821 | -8.9825e-06 | 0.0048669 | -0.00361 | 0.0028821 | 0.00181 | 0.0048669
100 | m | 0.2 | 200 | mHz | -72.3048 | 0 | 53.998 | 0 | m | -72.3 | 3.3814 | 54 | 7.8565 | m | OK | -72.3 | 54 | m | 0 | 0 | m | 0.001 | -0.072305 | 0 | 0.053998 | 0 | 0.090243 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | -5.4865e-06 | 0.0023032 | -5.4027e-07 | 0.0086516 | 200mHz100m100m | 6.0537e-06 | 0.0033814 | 2.319e-06 | 0.0077598 | 4.8026e-06 | 0.0033814 | 2.009e-06 | 0.0078565 | -0.0723 | 0.0033814 | 0.054 | 0.0078565
1 | m | 1000 | 1 | kHz | 0.31343 | 0 | 0.28996 | 0 | m | 0.315 | 9.6391 | 0.292 | 3.8773 | m | OK | 0.315 | 0.292 | m | -1.0565e-11 | 6.5355e-12 | m | 0.001 | 0.00031343 | 0 | 0.00028996 | 0 | 0.00042699 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 9.8775e-08 | 0.0082149 | 2.9458e-06 | 0.0038491 | 1kHz1m1m | 3.5429e-06 | 0.0096391 | 8.1518e-07 | 0.0039152 | 1.5694e-06 | 0.0096391 | 2.036e-06 | 0.0038773 | 0.000315 | 0.0096391 | 0.000292 | 0.0038773
3 | m | 1000 | 1 | kHz | -1.3323 | 0 | 0.08243 | 0 | m | -1.34 | 7.0375 | 0.0925 | 2.8935 | m | OK | -1.34 | 0.0925 | m | -2.7201e-10 | 2.6879e-10 | m | 0.001 | -0.0013323 | 0 | 8.243e-05 | 0 | 0.0013349 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | -8.9779e-06 | 0.00073379 | 1.1372e-05 | 0.0023675 | 1kHz3m3m | -1.1055e-06 | 0.0070375 | 3.593e-06 | 0.0055089 | -7.6597e-06 | 0.0070375 | 1.007e-05 | 0.0028935 | -0.00134 | 0.0070375 | 9.25e-05 | 0.0028935
3 | m | 0.2 | 200 | mHz | -0.026059 | 0 | 1.2551 | 0 | m | -0.0255 | 2.9446 | 1.25 | 5.3254 | m | OK | -0.0255 | 1.25 | m | 1.9682e-11 | 2.2784e-11 | m | 0.001 | -2.6059e-05 | 0 | 0.0012551 | 0 | 0.0012554 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | -1.0459e-07 | 0.0035905 | -5.8486e-06 | 0.0050804 | 200mHz3m3m | 5.0916e-06 | 0.0029446 | 1.6667e-07 | 0.0069992 | 5.5884e-07 | 0.0029446 | -5.0806e-06 | 0.0053254 | -2.55e-05 | 0.0029446 | 0.00125 | 0.0053254
10 | m | 100 | 100 | Hz | 6.0559 | 0 | 6.1309 | 0 | m | 6.06 | 8.1937 | 6.13 | 3.1284 | m | OK | 6.06 | 6.13 | m | -4.5741e-13 | 1.6893e-12 | m | 0.001 | 0.0060559 | 0 | 0.0061309 | 0 | 0.0086175 | 3 | 3m | 10m | 0.003 | 0.01 | 100Hz10m3m | 100Hz10m10m | 85 | 103 | 100Hz10m3m | 1.2373e-06 | 0.0093016 | 9.7094e-06 | 0.0033799 | 100Hz10m10m | 4.8106e-06 | 0.0081937 | -3.4914e-06 | 0.0030666 | 4.1048e-06 | 0.0081937 | -8.843e-07 | 0.0031284 | 0.00606 | 0.0081937 | 0.00613 | 0.0031284
10 | m | 50 | 50 | Hz | 5.7591 | 0 | 6.4073 | 0 | m | 5.76 | 2.8821 | 6.41 | 1.7512 | m | OK | 5.76 | 6.41 | m | -2.1733e-11 | -9.6232e-11 | m | 0.001 | 0.0057591 | 0 | 0.0064073 | 0 | 0.0086151 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | -2.328e-06 | 0.0018291 | -1.1645e-05 | 0.0055758 | 50Hz10m10m | 1.7163e-06 | 0.0028821 | 6.2636e-06 | 0.00080787 | 9.1622e-07 | 0.0028821 | 2.7205e-06 | 0.0017512 | 0.00576 | 0.0028821 | 0.00641 | 0.0017512
100 | m | 50 | 50 | Hz | 41.3004 | 0 | -14.1985 | 0 | m | 41.3 | 10.0431 | -14.2 | 6.6237 | m | OK | 41.3 | -14.2 | m | -2.2737e-13 | 5.5067e-14 | m | 0.001 | 0.0413 | 0 | -0.014199 | 0 | 0.043673 | 4 | 10m | 100m | 0.01 | 0.1 | 50Hz100m10m | 50Hz100m100m | 120 | 138 | 50Hz100m10m | -6.2662e-06 | 0.0080157 | -1.0542e-07 | 0.0085401 | 50Hz100m100m | 9.4487e-06 | 0.010043 | -3.8231e-06 | 0.003418 | -3.8658e-07 | 0.010043 | -1.4964e-06 | 0.0066237 | 0.0413 | 0.010043 | -0.0142 | 0.0066237
10 | m | 50 | 50 | Hz | 0.19789 | 0 | 5.4853 | 0 | m | 0.197 | 2.8821 | 5.48 | 3.8806 | m | OK | 0.197 | 5.48 | m | 3.3922e-11 | 1.5021e-10 | m | 0.001 | 0.00019789 | 0 | 0.0054853 | 0 | 0.0054888 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | -2.328e-06 | 0.0018291 | -1.1645e-05 | 0.0055758 | 50Hz10m10m | 1.7163e-06 | 0.0028821 | 6.2636e-06 | 0.00080787 | -8.9001e-07 | 0.0028821 | -5.2778e-06 | 0.0038806 | 0.000197 | 0.0028821 | 0.00548 | 0.0038806
1 | m | 1000 | 1 | kHz | 0.28188 | 0 | -0.51569 | 0 | m | 0.284 | 9.6391 | -0.514 | 3.8879 | m | OK | 0.284 | -0.514 | m | 5.9232e-12 | -3.6642e-12 | m | 0.001 | 0.00028188 | 0 | -0.00051569 | 0 | 0.0005877 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 9.8775e-08 | 0.0082149 | 2.9458e-06 | 0.0038491 | 1kHz1m1m | 3.5429e-06 | 0.0096391 | 8.1518e-07 | 0.0039152 | 2.1229e-06 | 0.0096391 | 1.6936e-06 | 0.0038879 | 0.000284 | 0.0096391 | -0.000514 | 0.0038879
100 | m | 500 | 500 | Hz | 39.903 | 0 | 18.3016 | 0 | m | 39.9 | 8.9432 | 18.3 | 7.8176 | m | OK | 39.9 | 18.3 | m | 0 | 0 | m | 0.001 | 0.039903 | 0 | 0.018302 | 0 | 0.0439 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | -1.8758e-06 | 0.0098276 | 2.6486e-07 | 0.0081162 | 500Hz100m100m | -4.7634e-06 | 0.0089432 | -4.7152e-06 | 0.0073233 | -2.9634e-06 | 0.0089432 | -1.611e-06 | 0.0078176 | 0.0399 | 0.0089432 | 0.0183 | 0.0078176
1 | m | 5 | 5 | Hz | 0.28323 | 0 | 0.79878 | 0 | m | 0.276 | 6.7277 | 0.792 | 1.9237 | m | OK | 0.276 | 0.792 | m | -1.4718e-09 | -2.2192e-10 | m | 0.001 | 0.00028323 | 0 | 0.00079878 | 0 | 0.0008475 | 1 | 0m | 1m | 0 | 0.001 | 5Hz1m0m | 5Hz1m1m | 9 | 27 | 5Hz1m0m | 3.4079e-07 | 0.0079365 | -5.6364e-06 | 0.0035676 | 5Hz1m1m | -8.5912e-06 | 0.0067277 | -6.9831e-06 | 0.0016279 | -7.2291e-06 | 0.0067277 | -6.7778e-06 | 0.0019237 | 0.000276 | 0.0067277 | 0.000792 | 0.0019237
100 | m | 1000 | 1 | kHz | 52.9043 | 0 | 50.2023 | 0 | m | 52.9 | 11.5183 | 50.2 | 2.7954 | m | OK | 52.9 | 50.2 | m | 0 | 0 | m | 0.001 | 0.052904 | 0 | 0.050202 | 0 | 0.072932 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | -2.5073e-06 | 0.0056074 | 6.5281e-06 | 0.0091804 | 1kHz100m100m | -5.0663e-06 | 0.011518 | -6.0504e-06 | 4.916e-05 | -4.2967e-06 | 0.011518 | -2.2674e-06 | 0.0027954 | 0.0529 | 0.011518 | 0.0502 | 0.0027954
3 | m | 0.1 | 100 | mHz | 0.079747 | 0 | -1.2009 | 0 | m | 0.082 | 5.4921 | -1.2 | 0.57332 | m | OK | 0.082 | -1.2 | m | 9.1546e-12 | 6.2221e-12 | m | 0.001 | 7.9747e-05 | 0 | -0.0012009 | 0 | 0.0012035 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | 2.9129e-06 | 0.0039352 | 1.3377e-06 | 9.2827e-05 | 100mHz3m3m | -3.5713e-06 | 0.0054921 | -3.0694e-06 | 0.0048143 | 2.253e-06 | 0.0054921 | 8.8923e-07 | 0.00057332 | 8.2e-05 | 0.0054921 | -0.0012 | 0.00057332
100 | m | 200 | 200 | Hz | -28.7991 | 0 | 73.5032 | 0 | m | -28.8 | 5.8761 | 73.5 | 1.7539 | m | OK | -28.8 | 73.5 | m | 0 | 0 | m | 0.001 | -0.028799 | 0 | 0.073503 | 0 | 0.078944 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | 3.5932e-06 | 0.0042566 | -1.7526e-06 | 0.0025144 | 200Hz100m100m | -2.277e-06 | 0.0058761 | -3.5953e-06 | 0.0015216 | -9.036e-07 | 0.0058761 | -3.1642e-06 | 0.0017539 | -0.0288 | 0.0058761 | 0.0735 | 0.0017539
10 | m | 5 | 5 | Hz | -2.8378 | 0 | -3.484 | 0 | m | -2.84 | 8.5804 | -3.49 | 1.6563 | m | OK | -2.84 | -3.49 | m | 9.77e-14 | 1.5588e-13 | m | 0.001 | -0.0028378 | 0 | -0.003484 | 0 | 0.0044934 | 3 | 3m | 10m | 0.003 | 0.01 | 5Hz10m3m | 5Hz10m10m | 81 | 99 | 5Hz10m3m | -1.8062e-06 | 0.0017251 | -5.3584e-06 | 0.0015636 | 5Hz10m10m | -3.7766e-06 | 0.0085804 | -8.5028e-06 | 0.0019981 | -2.2266e-06 | 0.0085804 | -6.0292e-06 | 0.0016563 | -0.00284 | 0.0085804 | -0.00349 | 0.0016563
10 | m | 5000 | 5 | kHz | -7.3162 | 0 | 3.6887 | 0 | m | -7.32 | 6.4858 | 3.68 | 8.2976 | m | OK | -7.32 | 3.68 | m | -3.7303e-14 | -1.4255e-13 | m | 0.001 | -0.0073162 | 0 | 0.0036887 | 0 | 0.0081935 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | -2.3167e-06 | 0.0060027 | -2.8645e-06 | 0.0055319 | 5kHz10m10m | -4.3437e-06 | 0.0064858 | -1.0677e-05 | 0.0092597 | -3.8206e-06 | 0.0064858 | -8.6608e-06 | 0.0082976 | -0.00732 | 0.0064858 | 0.00368 | 0.0082976
100 | m | 5000 | 5 | kHz | -5.8635 | 0 | -12.4991 | 0 | m | -5.86 | 11.3309 | -12.5 | 1.6241 | m | OK | -5.86 | -12.5 | m | 0 | 0 | m | 0.001 | -0.0058635 | 0 | -0.012499 | 0 | 0.013806 | 4 | 10m | 100m | 0.01 | 0.1 | 5kHz100m10m | 5kHz100m100m | 126 | 144 | 5kHz100m10m | 3.1651e-06 | 0.0054299 | -8.8348e-07 | 0.0016282 | 5kHz100m100m | 1.0411e-05 | 0.011331 | -3.3759e-07 | 0.0015322 | 3.4715e-06 | 0.011331 | -8.6039e-07 | 0.0016241 | -0.00586 | 0.011331 | -0.0125 | 0.0016241
3 | m | 0.2 | 200 | mHz | -0.82183 | 0 | -2.7799 | 0 | m | -0.817 | 2.9446 | -2.78 | 6.9021 | m | OK | -0.817 | -2.78 | m | 8.0097e-11 | 9.2722e-11 | m | 0.001 | -0.00082183 | 0 | -0.0027799 | 0 | 0.0028988 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | -1.0459e-07 | 0.0035905 | -5.8486e-06 | 0.0050804 | 200mHz3m3m | 5.0916e-06 | 0.0029446 | 1.6667e-07 | 0.0069992 | 4.8287e-06 | 0.0029446 | -1.377e-07 | 0.0069021 | -0.000817 | 0.0029446 | -0.00278 | 0.0069021

## Ranges
Range list
Range |  | Range Z | Min ¦Z¦ |  | Min Z | Max ¦Z¦ |  | Max Z
Ohm |  | Ohm | Ohm |  | Ohm | Ohm |  | Ohm
1 | m | 0.001 | 0 | m | 0 | 1 | m | 0.001
3 | m | 0.003 | 1 | m | 0.001 | 3 | m | 0.003
10 | m | 0.01 | 3 | m | 0.003 | 10 | m | 0.01
100 | m | 0.1 | 10 | m | 0.01 | 100 | m | 0.1

## Cal Data
Calibration data: measured values of reference standards.
[9 empty] | Measured calibration Z | [10 empty] | Reference Z (auto search from Ref Z sheet) | [5 empty] | Absolute correction
Range |  | Range Z | f | f |  | Nominal Z |  | Nom Z | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | raf tag | ref ID | Rs | U(Rs) | Xs | U(Xs) | tag | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs))
Ohm |  | Ohm | Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm
1 | m | 0.001 | 0.01 | 10 | mHz | 0 | m | 0 | 0.0048876 | 0.0039755 | -0.0045148 | 0.0042932 | m | 0.001 | 4.8876e-06 | 3.9755e-06 | -4.5148e-06 | 4.2932e-06 | 6.6537e-06 | 10mHz0m | 1 | 0 | 0 | 0 | 0 | 10mHz1m0m | -4.8876e-06 | 0.0079511 | 4.5148e-06 | 0.0085864
1 | m | 0.001 | 0.02 | 20 | mHz | 0 | m | 0 | 0.0025644 | 0.0024965 | 0.0026984 | 0.0028244 | m | 0.001 | 2.5644e-06 | 2.4965e-06 | 2.6984e-06 | 2.8244e-06 | 3.7225e-06 | 20mHz0m | 2 | 0 | 0 | 0 | 0 | 20mHz1m0m | -2.5644e-06 | 0.0049929 | -2.6984e-06 | 0.0056488
1 | m | 0.001 | 0.05 | 50 | mHz | 0 | m | 0 | 0.0090717 | 0.003133 | -0.0026659 | 0.00085977 | m | 0.001 | 9.0717e-06 | 3.133e-06 | -2.6659e-06 | 8.5977e-07 | 9.4553e-06 | 50mHz0m | 3 | 0 | 0 | 0 | 0 | 50mHz1m0m | -9.0717e-06 | 0.006266 | 2.6659e-06 | 0.0017195
1 | m | 0.001 | 0.1 | 100 | mHz | 0 | m | 0 | 0.0058269 | 0.0023732 | -0.0073356 | 0.0041098 | m | 0.001 | 5.8269e-06 | 2.3732e-06 | -7.3356e-06 | 4.1098e-06 | 9.3682e-06 | 100mHz0m | 4 | 0 | 0 | 0 | 0 | 100mHz1m0m | -5.8269e-06 | 0.0047465 | 7.3356e-06 | 0.0082195
1 | m | 0.001 | 0.2 | 200 | mHz | 0 | m | 0 | 0.0073927 | 0.00086282 | -0.010729 | 0.0042326 | m | 0.001 | 7.3927e-06 | 8.6282e-07 | -1.0729e-05 | 4.2326e-06 | 1.3029e-05 | 200mHz0m | 5 | 0 | 0 | 0 | 0 | 200mHz1m0m | -7.3927e-06 | 0.0017256 | 1.0729e-05 | 0.0084651
1 | m | 0.001 | 0.5 | 500 | mHz | 0 | m | 0 | 0.0011465 | 0.0036619 | 0.0055894 | 0.0030952 | m | 0.001 | 1.1465e-06 | 3.6619e-06 | 5.5894e-06 | 3.0952e-06 | 5.7057e-06 | 500mHz0m | 6 | 0 | 0 | 0 | 0 | 500mHz1m0m | -1.1465e-06 | 0.0073238 | -5.5894e-06 | 0.0061905
1 | m | 0.001 | 1 | 1 | Hz | 0 | m | 0 | -0.0034144 | 0.0015414 | -0.006249 | 0.0047731 | m | 0.001 | -3.4144e-06 | 1.5414e-06 | -6.249e-06 | 4.7731e-06 | 7.121e-06 | 1Hz0m | 7 | 0 | 0 | 0 | 0 | 1Hz1m0m | 3.4144e-06 | 0.0030828 | 6.249e-06 | 0.0095462
1 | m | 0.001 | 2 | 2 | Hz | 0 | m | 0 | -0.0027014 | 0.0024326 | 0.0025448 | 0.001932 | m | 0.001 | -2.7014e-06 | 2.4326e-06 | 2.5448e-06 | 1.932e-06 | 3.7112e-06 | 2Hz0m | 8 | 0 | 0 | 0 | 0 | 2Hz1m0m | 2.7014e-06 | 0.0048652 | -2.5448e-06 | 0.0038639
1 | m | 0.001 | 5 | 5 | Hz | 0 | m | 0 | -0.00034079 | 0.0039683 | 0.0056364 | 0.0017838 | m | 0.001 | -3.4079e-07 | 3.9683e-06 | 5.6364e-06 | 1.7838e-06 | 5.6467e-06 | 5Hz0m | 9 | 0 | 0 | 0 | 0 | 5Hz1m0m | 3.4079e-07 | 0.0079365 | -5.6364e-06 | 0.0035676
1 | m | 0.001 | 10 | 10 | Hz | 0 | m | 0 | 0.0059086 | 0.0036581 | -0.005218 | 5.3282e-05 | m | 0.001 | 5.9086e-06 | 3.6581e-06 | -5.218e-06 | 5.3282e-08 | 7.8828e-06 | 10Hz0m | 10 | 0 | 0 | 0 | 0 | 10Hz1m0m | -5.9086e-06 | 0.0073162 | 5.218e-06 | 0.00010656
1 | m | 0.001 | 20 | 20 | Hz | 0 | m | 0 | -0.0026642 | 0.0043719 | 0.0079212 | 0.0020292 | m | 0.001 | -2.6642e-06 | 4.3719e-06 | 7.9212e-06 | 2.0292e-06 | 8.3573e-06 | 20Hz0m | 11 | 0 | 0 | 0 | 0 | 20Hz1m0m | 2.6642e-06 | 0.0087438 | -7.9212e-06 | 0.0040584
1 | m | 0.001 | 50 | 50 | Hz | 0 | m | 0 | 0.0034708 | 6.1477e-05 | -0.0047889 | 0.00014109 | m | 0.001 | 3.4708e-06 | 6.1477e-08 | -4.7889e-06 | 1.4109e-07 | 5.9144e-06 | 50Hz0m | 12 | 0 | 0 | 0 | 0 | 50Hz1m0m | -3.4708e-06 | 0.00012295 | 4.7889e-06 | 0.00028219
1 | m | 0.001 | 100 | 100 | Hz | 0 | m | 0 | 0.00037614 | 0.0003423 | 0.00039282 | 0.0042604 | m | 0.001 | 3.7614e-07 | 3.423e-07 | 3.9282e-07 | 4.2604e-06 | 5.4387e-07 | 100Hz0m | 13 | 0 | 0 | 0 | 0 | 100Hz1m0m | -3.7614e-07 | 0.00068461 | -3.9282e-07 | 0.0085207
1 | m | 0.001 | 200 | 200 | Hz | 0 | m | 0 | -0.0063507 | 0.0036533 | 0.0041062 | 0.0012186 | m | 0.001 | -6.3507e-06 | 3.6533e-06 | 4.1062e-06 | 1.2186e-06 | 7.5625e-06 | 200Hz0m | 14 | 0 | 0 | 0 | 0 | 200Hz1m0m | 6.3507e-06 | 0.0073066 | -4.1062e-06 | 0.0024372
1 | m | 0.001 | 500 | 500 | Hz | 0 | m | 0 | 0.00033528 | 0.0010699 | 0.0029916 | 0.0013795 | m | 0.001 | 3.3528e-07 | 1.0699e-06 | 2.9916e-06 | 1.3795e-06 | 3.0103e-06 | 500Hz0m | 15 | 0 | 0 | 0 | 0 | 500Hz1m0m | -3.3528e-07 | 0.0021398 | -2.9916e-06 | 0.0027591
1 | m | 0.001 | 1000 | 1 | kHz | 0 | m | 0 | -9.8775e-05 | 0.0041074 | -0.0029458 | 0.0019245 | m | 0.001 | -9.8775e-08 | 4.1074e-06 | -2.9458e-06 | 1.9245e-06 | 2.9474e-06 | 1kHz0m | 16 | 0 | 0 | 0 | 0 | 1kHz1m0m | 9.8775e-08 | 0.0082149 | 2.9458e-06 | 0.0038491
1 | m | 0.001 | 2000 | 2 | kHz | 0 | m | 0 | 0.0047259 | 0.0044373 | 0.0027211 | 0.0039402 | m | 0.001 | 4.7259e-06 | 4.4373e-06 | 2.7211e-06 | 3.9402e-06 | 5.4533e-06 | 2kHz0m | 17 | 0 | 0 | 0 | 0 | 2kHz1m0m | -4.7259e-06 | 0.0088746 | -2.7211e-06 | 0.0078804
1 | m | 0.001 | 5000 | 5 | kHz | 0 | m | 0 | 0.0061967 | 0.0047044 | 0.0044432 | 0.0033856 | m | 0.001 | 6.1967e-06 | 4.7044e-06 | 4.4432e-06 | 3.3856e-06 | 7.625e-06 | 5kHz0m | 18 | 0 | 0 | 0 | 0 | 5kHz1m0m | -6.1967e-06 | 0.0094088 | -4.4432e-06 | 0.0067713
1 | m | 0.001 | 0.01 | 10 | mHz | 1 | m | 0.001 | 1.0015 | 0.0049338 | 0.005569 | 0.00076478 | m | 0.001 | 0.0010015 | 4.9338e-06 | 5.569e-06 | 7.6478e-07 | 0.0010015 | 10mHz1m | 19 | 0.00099993 | 1e-08 | -1.1463e-07 | 5e-08 | 10mHz1m1m | -1.5379e-06 | 0.0098676 | -5.6836e-06 | 0.0015296
1 | m | 0.001 | 0.02 | 20 | mHz | 1 | m | 0.001 | 0.99691 | 0.0038564 | 0.0038403 | 0.0010146 | m | 0.001 | 0.00099691 | 3.8564e-06 | 3.8403e-06 | 1.0146e-06 | 0.00099692 | 20mHz1m | 20 | 0.00099987 | 1e-08 | 8.6308e-08 | 5e-08 | 20mHz1m1m | 2.9575e-06 | 0.0077127 | -3.754e-06 | 0.0020291
1 | m | 0.001 | 0.05 | 50 | mHz | 1 | m | 0.001 | 0.99307 | 0.0047523 | -0.0081906 | 0.0040834 | m | 0.001 | 0.00099307 | 4.7523e-06 | -8.1906e-06 | 4.0834e-06 | 0.0009931 | 50mHz1m | 21 | 0.00099979 | 1e-08 | 8.9699e-08 | 5e-08 | 50mHz1m1m | 6.7167e-06 | 0.0095046 | 8.2803e-06 | 0.0081667
1 | m | 0.001 | 0.1 | 100 | mHz | 1 | m | 0.001 | 1.0019 | 0.0042957 | -0.0068079 | 0.0016579 | m | 0.001 | 0.0010019 | 4.2957e-06 | -6.8079e-06 | 1.6579e-06 | 0.0010019 | 100mHz1m | 22 | 0.00099986 | 1e-08 | -1.692e-08 | 5e-08 | 100mHz1m1m | -2.0685e-06 | 0.0085913 | 6.7909e-06 | 0.0033158
1 | m | 0.001 | 0.2 | 200 | mHz | 1 | m | 0.001 | 0.99683 | 0.0013912 | 0.0054 | 0.0017329 | m | 0.001 | 0.00099683 | 1.3912e-06 | 5.4e-06 | 1.7329e-06 | 0.00099685 | 200mHz1m | 23 | 0.00099986 | 1e-08 | -6.4602e-08 | 5e-08 | 200mHz1m1m | 3.0245e-06 | 0.0027825 | -5.4646e-06 | 0.0034658
1 | m | 0.001 | 0.5 | 500 | mHz | 1 | m | 0.001 | 0.99652 | 0.0029778 | 0.0013118 | 0.0047161 | m | 0.001 | 0.00099652 | 2.9778e-06 | 1.3118e-06 | 4.7161e-06 | 0.00099652 | 500mHz1m | 24 | 0.00099999 | 1e-08 | -1.903e-07 | 5e-08 | 500mHz1m1m | 3.4718e-06 | 0.0059555 | -1.5021e-06 | 0.0094322
1 | m | 0.001 | 1 | 1 | Hz | 1 | m | 0.001 | 0.99281 | 0.0036531 | 0.0045258 | 0.0012655 | m | 0.001 | 0.00099281 | 3.6531e-06 | 4.5258e-06 | 1.2655e-06 | 0.00099283 | 1Hz1m | 25 | 0.001 | 1e-08 | 3.7586e-08 | 5e-08 | 1Hz1m1m | 7.2287e-06 | 0.0073063 | -4.4882e-06 | 0.002531
1 | m | 0.001 | 2 | 2 | Hz | 1 | m | 0.001 | 1.005 | 0.00075154 | 0.007409 | 0.00061613 | m | 0.001 | 0.001005 | 7.5154e-07 | 7.409e-06 | 6.1613e-07 | 0.001005 | 2Hz1m | 26 | 0.0010001 | 1e-08 | 1.4693e-08 | 5e-08 | 2Hz1m1m | -4.8561e-06 | 0.0015031 | -7.3943e-06 | 0.0012323
1 | m | 0.001 | 5 | 5 | Hz | 1 | m | 0.001 | 1.0087 | 0.0033639 | 0.0067735 | 0.00081397 | m | 0.001 | 0.0010087 | 3.3639e-06 | 6.7735e-06 | 8.1397e-07 | 0.0010087 | 5Hz1m | 27 | 0.0010001 | 1e-08 | -2.0967e-07 | 5e-08 | 5Hz1m1m | -8.5912e-06 | 0.0067277 | -6.9831e-06 | 0.0016279
1 | m | 0.001 | 10 | 10 | Hz | 1 | m | 0.001 | 1.0001 | 0.0018483 | 0.00036201 | 0.0037788 | m | 0.001 | 0.0010001 | 1.8483e-06 | 3.6201e-07 | 3.7788e-06 | 0.0010001 | 10Hz1m | 28 | 0.0010001 | 1e-08 | -6.2055e-09 | 5e-08 | 10Hz1m1m | 2.3121e-08 | 0.0036966 | -3.6821e-07 | 0.0075576
1 | m | 0.001 | 20 | 20 | Hz | 1 | m | 0.001 | 0.99385 | 0.0040399 | -0.0024252 | 0.0030506 | m | 0.001 | 0.00099385 | 4.0399e-06 | -2.4252e-06 | 3.0506e-06 | 0.00099385 | 20Hz1m | 29 | 0.0010001 | 1e-08 | -8.0442e-08 | 5e-08 | 20Hz1m1m | 6.2024e-06 | 0.0080799 | 2.3448e-06 | 0.0061013
1 | m | 0.001 | 50 | 50 | Hz | 1 | m | 0.001 | 0.99968 | 0.00015811 | 0.0034818 | 0.0012988 | m | 0.001 | 0.00099968 | 1.5811e-07 | 3.4818e-06 | 1.2988e-06 | 0.00099968 | 50Hz1m | 30 | 0.0010002 | 1e-08 | -4.9127e-08 | 5e-08 | 50Hz1m1m | 5.01e-07 | 0.00031621 | -3.5309e-06 | 0.0025977
1 | m | 0.001 | 100 | 100 | Hz | 1 | m | 0.001 | 1.007 | 0.0040613 | 0.0013269 | 0.00066986 | m | 0.001 | 0.001007 | 4.0613e-06 | 1.3269e-06 | 6.6986e-07 | 0.001007 | 100Hz1m | 31 | 0.0010001 | 1e-08 | 2.3187e-07 | 5e-08 | 100Hz1m1m | -6.9131e-06 | 0.0081227 | -1.0951e-06 | 0.0013397
1 | m | 0.001 | 200 | 200 | Hz | 1 | m | 0.001 | 0.99728 | 0.00063399 | 0.0012619 | 0.0010677 | m | 0.001 | 0.00099728 | 6.3399e-07 | 1.2619e-06 | 1.0677e-06 | 0.00099728 | 200Hz1m | 32 | 0.0010008 | 1e-08 | 4.2778e-08 | 5e-08 | 200Hz1m1m | 3.4761e-06 | 0.001268 | -1.2191e-06 | 0.0021353
1 | m | 0.001 | 500 | 500 | Hz | 1 | m | 0.001 | 0.99541 | 0.0046075 | -0.0027512 | 0.0028313 | m | 0.001 | 0.00099541 | 4.6075e-06 | -2.7512e-06 | 2.8313e-06 | 0.00099541 | 500Hz1m | 33 | 0.0010026 | 1.5811e-08 | 8.1299e-07 | 5e-08 | 500Hz1m1m | 7.2434e-06 | 0.009215 | 3.5641e-06 | 0.0056626
1 | m | 0.001 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0038 | 0.0048195 | 0.00039948 | 0.0019576 | m | 0.001 | 0.0010038 | 4.8195e-06 | 3.9948e-07 | 1.9576e-06 | 0.0010038 | 1kHz1m | 34 | 0.0010073 | 4.4721e-08 | 1.2147e-06 | 1e-07 | 1kHz1m1m | 3.5429e-06 | 0.0096391 | 8.1518e-07 | 0.0039152
1 | m | 0.001 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0234 | 0.0024548 | 0.0044125 | 0.0042267 | m | 0.001 | 0.0010234 | 2.4548e-06 | 4.4125e-06 | 4.2267e-06 | 0.0010234 | 2kHz1m | 35 | 0.0010208 | 1.2649e-07 | 2.5962e-06 | 2e-07 | 2kHz1m1m | -2.5952e-06 | 0.0049096 | -1.8163e-06 | 0.0084535
1 | m | 0.001 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.0881 | 0.00078947 | 0.0036946 | 0.004321 | m | 0.001 | 0.0010881 | 7.8947e-07 | 3.6946e-06 | 4.321e-06 | 0.0010881 | 5kHz1m | 36 | 0.0010827 | 5e-07 | 6.277e-06 | 5e-07 | 5kHz1m1m | -5.4046e-06 | 0.0015789 | 2.5824e-06 | 0.0086419
3 | m | 0.003 | 0.01 | 10 | mHz | 1 | m | 0.001 | 1.0061 | 0.0034544 | 0.0022999 | 0.004051 | m | 0.001 | 0.0010061 | 3.4544e-06 | 2.2999e-06 | 4.051e-06 | 0.0010061 | 10mHz1m | 19 | 0.00099993 | 1e-08 | -1.1463e-07 | 5e-08 | 10mHz3m1m | -6.1621e-06 | 0.0069087 | -2.4146e-06 | 0.008102
3 | m | 0.003 | 0.02 | 20 | mHz | 1 | m | 0.001 | 0.99267 | 0.0024317 | 0.0073959 | 0.0022553 | m | 0.001 | 0.00099267 | 2.4317e-06 | 7.3959e-06 | 2.2553e-06 | 0.0009927 | 20mHz1m | 20 | 0.00099987 | 1e-08 | 8.6308e-08 | 5e-08 | 20mHz3m1m | 7.2004e-06 | 0.0048633 | -7.3096e-06 | 0.0045106
3 | m | 0.003 | 0.05 | 50 | mHz | 1 | m | 0.001 | 0.99853 | 0.0033246 | 0.0098367 | 0.004038 | m | 0.001 | 0.00099853 | 3.3246e-06 | 9.8367e-06 | 4.038e-06 | 0.00099858 | 50mHz1m | 21 | 0.00099979 | 1e-08 | 8.9699e-08 | 5e-08 | 50mHz3m1m | 1.2508e-06 | 0.0066493 | -9.747e-06 | 0.008076
3 | m | 0.003 | 0.1 | 100 | mHz | 1 | m | 0.001 | 0.99694 | 0.0019676 | -0.0013546 | 4.6413e-05 | m | 0.001 | 0.00099694 | 1.9676e-06 | -1.3546e-06 | 4.6413e-08 | 0.00099694 | 100mHz1m | 22 | 0.00099986 | 1e-08 | -1.692e-08 | 5e-08 | 100mHz3m1m | 2.9129e-06 | 0.0039352 | 1.3377e-06 | 9.2827e-05
3 | m | 0.003 | 0.2 | 200 | mHz | 1 | m | 0.001 | 0.99996 | 0.0017952 | 0.005784 | 0.0025402 | m | 0.001 | 0.00099996 | 1.7952e-06 | 5.784e-06 | 2.5402e-06 | 0.00099998 | 200mHz1m | 23 | 0.00099986 | 1e-08 | -6.4602e-08 | 5e-08 | 200mHz3m1m | -1.0459e-07 | 0.0035905 | -5.8486e-06 | 0.0050804
3 | m | 0.003 | 0.5 | 500 | mHz | 1 | m | 0.001 | 1.0001 | 0.0024122 | 0.0016233 | 0.0011363 | m | 0.001 | 0.0010001 | 2.4122e-06 | 1.6233e-06 | 1.1363e-06 | 0.0010002 | 500mHz1m | 24 | 0.00099999 | 1e-08 | -1.903e-07 | 5e-08 | 500mHz3m1m | -1.612e-07 | 0.0048245 | -1.8136e-06 | 0.0022726
3 | m | 0.003 | 1 | 1 | Hz | 1 | m | 0.001 | 0.99209 | 0.00094365 | 0.0037772 | 0.0033376 | m | 0.001 | 0.00099209 | 9.4365e-07 | 3.7772e-06 | 3.3376e-06 | 0.0009921 | 1Hz1m | 25 | 0.001 | 1e-08 | 3.7586e-08 | 5e-08 | 1Hz3m1m | 7.9524e-06 | 0.0018873 | -3.7396e-06 | 0.0066752
3 | m | 0.003 | 2 | 2 | Hz | 1 | m | 0.001 | 1.0032 | 0.0022234 | -0.010305 | 0.00098142 | m | 0.001 | 0.0010032 | 2.2234e-06 | -1.0305e-05 | 9.8142e-07 | 0.0010032 | 2Hz1m | 26 | 0.0010001 | 1e-08 | 1.4693e-08 | 5e-08 | 2Hz3m1m | -3.0351e-06 | 0.0044469 | 1.032e-05 | 0.0019628
3 | m | 0.003 | 5 | 5 | Hz | 1 | m | 0.001 | 0.99858 | 0.0038942 | 0.0053534 | 4.1345e-05 | m | 0.001 | 0.00099858 | 3.8942e-06 | 5.3534e-06 | 4.1345e-08 | 0.0009986 | 5Hz1m | 27 | 0.0010001 | 1e-08 | -2.0967e-07 | 5e-08 | 5Hz3m1m | 1.494e-06 | 0.0077883 | -5.5631e-06 | 8.269e-05
3 | m | 0.003 | 10 | 10 | Hz | 1 | m | 0.001 | 0.989 | 0.0025187 | 0.0031034 | 0.004503 | m | 0.001 | 0.000989 | 2.5187e-06 | 3.1034e-06 | 4.503e-06 | 0.00098901 | 10Hz1m | 28 | 0.0010001 | 1e-08 | -6.2055e-09 | 5e-08 | 10Hz3m1m | 1.1083e-05 | 0.0050374 | -3.1096e-06 | 0.009006
3 | m | 0.003 | 20 | 20 | Hz | 1 | m | 0.001 | 1.0015 | 0.0013293 | -0.0015936 | 0.0023252 | m | 0.001 | 0.0010015 | 1.3293e-06 | -1.5936e-06 | 2.3252e-06 | 0.0010015 | 20Hz1m | 29 | 0.0010001 | 1e-08 | -8.0442e-08 | 5e-08 | 20Hz3m1m | -1.4817e-06 | 0.0026586 | 1.5132e-06 | 0.0046503
3 | m | 0.003 | 50 | 50 | Hz | 1 | m | 0.001 | 0.99604 | 0.0022458 | -0.00050085 | 0.0028151 | m | 0.001 | 0.00099604 | 2.2458e-06 | -5.0085e-07 | 2.8151e-06 | 0.00099604 | 50Hz1m | 30 | 0.0010002 | 1e-08 | -4.9127e-08 | 5e-08 | 50Hz3m1m | 4.1395e-06 | 0.0044916 | 4.5172e-07 | 0.0056302
3 | m | 0.003 | 100 | 100 | Hz | 1 | m | 0.001 | 0.99823 | 0.0024812 | -0.0020769 | 0.00077468 | m | 0.001 | 0.00099823 | 2.4812e-06 | -2.0769e-06 | 7.7468e-07 | 0.00099823 | 100Hz1m | 31 | 0.0010001 | 1e-08 | 2.3187e-07 | 5e-08 | 100Hz3m1m | 1.8965e-06 | 0.0049624 | 2.3088e-06 | 0.0015494
3 | m | 0.003 | 200 | 200 | Hz | 1 | m | 0.001 | 1.0094 | 0.0012896 | -0.0048473 | 0.0032838 | m | 0.001 | 0.0010094 | 1.2896e-06 | -4.8473e-06 | 3.2838e-06 | 0.0010094 | 200Hz1m | 32 | 0.0010008 | 1e-08 | 4.2778e-08 | 5e-08 | 200Hz3m1m | -8.6395e-06 | 0.0025792 | 4.8901e-06 | 0.0065675
3 | m | 0.003 | 500 | 500 | Hz | 1 | m | 0.001 | 0.99771 | 6.337e-05 | 0.0083757 | 0.00025904 | m | 0.001 | 0.00099771 | 6.337e-08 | 8.3757e-06 | 2.5904e-07 | 0.00099775 | 500Hz1m | 33 | 0.0010026 | 1.5811e-08 | 8.1299e-07 | 5e-08 | 500Hz3m1m | 4.9361e-06 | 0.00012674 | -7.5628e-06 | 0.00051809
3 | m | 0.003 | 1000 | 1 | kHz | 1 | m | 0.001 | 1.0163 | 0.0003669 | -0.010158 | 0.0011838 | m | 0.001 | 0.0010163 | 3.669e-07 | -1.0158e-05 | 1.1838e-06 | 0.0010164 | 1kHz1m | 34 | 0.0010073 | 4.4721e-08 | 1.2147e-06 | 1e-07 | 1kHz3m1m | -8.9779e-06 | 0.00073379 | 1.1372e-05 | 0.0023675
3 | m | 0.003 | 2000 | 2 | kHz | 1 | m | 0.001 | 1.0189 | 0.0034291 | 0.0071426 | 0.0045987 | m | 0.001 | 0.0010189 | 3.4291e-06 | 7.1426e-06 | 4.5987e-06 | 0.0010189 | 2kHz1m | 35 | 0.0010208 | 1.2649e-07 | 2.5962e-06 | 2e-07 | 2kHz3m1m | 1.9461e-06 | 0.0068581 | -4.5464e-06 | 0.0091974
3 | m | 0.003 | 5000 | 5 | kHz | 1 | m | 0.001 | 1.0707 | 0.0027163 | 0.011352 | 0.0031562 | m | 0.001 | 0.0010707 | 2.7163e-06 | 1.1352e-05 | 3.1562e-06 | 0.0010708 | 5kHz1m | 36 | 0.0010827 | 5e-07 | 6.277e-06 | 5e-07 | 5kHz3m1m | 1.1929e-05 | 0.0054327 | -5.075e-06 | 0.0063123
3 | m | 0.003 | 0.01 | 10 | mHz | 3 | m | 0.003 | 3.0011 | 0.0019051 | -0.0016302 | 0.0022889 | m | 0.001 | 0.0030011 | 1.9051e-06 | -1.6302e-06 | 2.2889e-06 | 0.0030011 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 1.8535e-07 | 1.5e-07 | 10mHz3m3m | -1.1457e-06 | 0.0038102 | 1.8155e-06 | 0.0045779
3 | m | 0.003 | 0.02 | 20 | mHz | 3 | m | 0.003 | 3.0021 | 0.00044157 | 0.0027113 | 0.0029592 | m | 0.001 | 0.0030021 | 4.4157e-07 | 2.7113e-06 | 2.9592e-06 | 0.0030021 | 20mHz3m | 38 | 0.003 | 3e-08 | -2.3489e-07 | 1.5e-07 | 20mHz3m3m | -2.091e-06 | 0.00088313 | -2.9462e-06 | 0.0059184
3 | m | 0.003 | 0.05 | 50 | mHz | 3 | m | 0.003 | 2.9963 | 0.0029704 | -0.007214 | 0.0033181 | m | 0.001 | 0.0029963 | 2.9704e-06 | -7.214e-06 | 3.3181e-06 | 0.0029963 | 50mHz3m | 39 | 0.0029997 | 3e-08 | -1.1531e-07 | 1.5e-07 | 50mHz3m3m | 3.3915e-06 | 0.0059409 | 7.0987e-06 | 0.0066362
3 | m | 0.003 | 0.1 | 100 | mHz | 3 | m | 0.003 | 3.0039 | 0.0027461 | 0.0031885 | 0.0024072 | m | 0.001 | 0.0030039 | 2.7461e-06 | 3.1885e-06 | 2.4072e-06 | 0.0030039 | 100mHz3m | 40 | 0.0030003 | 3e-08 | 1.1905e-07 | 1.5e-07 | 100mHz3m3m | -3.5713e-06 | 0.0054921 | -3.0694e-06 | 0.0048143
3 | m | 0.003 | 0.2 | 200 | mHz | 3 | m | 0.003 | 2.9945 | 0.0014723 | -0.00039135 | 0.0034996 | m | 0.001 | 0.0029945 | 1.4723e-06 | -3.9135e-07 | 3.4996e-06 | 0.0029945 | 200mHz3m | 41 | 0.0029996 | 3e-08 | -2.2468e-07 | 1.5e-07 | 200mHz3m3m | 5.0916e-06 | 0.0029446 | 1.6667e-07 | 0.0069992
3 | m | 0.003 | 0.5 | 500 | mHz | 3 | m | 0.003 | 3.0073 | 0.0012295 | 0.0091379 | 0.0026387 | m | 0.001 | 0.0030073 | 1.2295e-06 | 9.1379e-06 | 2.6387e-06 | 0.0030074 | 500mHz3m | 42 | 0.003 | 3e-08 | -1.4963e-08 | 1.5e-07 | 500mHz3m3m | -7.3022e-06 | 0.002459 | -9.1528e-06 | 0.0052774
3 | m | 0.003 | 1 | 1 | Hz | 3 | m | 0.003 | 3.0058 | 0.0039153 | -0.0043565 | 0.0024801 | m | 0.001 | 0.0030058 | 3.9153e-06 | -4.3565e-06 | 2.4801e-06 | 0.0030058 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -5.2952e-07 | 1.5e-07 | 1Hz3m3m | -6.1901e-06 | 0.0078306 | 3.8269e-06 | 0.0049602
3 | m | 0.003 | 2 | 2 | Hz | 3 | m | 0.003 | 2.9988 | 0.003656 | -0.0022979 | 0.0039224 | m | 0.001 | 0.0029988 | 3.656e-06 | -2.2979e-06 | 3.9224e-06 | 0.0029988 | 2Hz3m | 44 | 0.0029999 | 3e-08 | 3.5039e-07 | 1.5e-07 | 2Hz3m3m | 1.1014e-06 | 0.0073121 | 2.6483e-06 | 0.0078449
3 | m | 0.003 | 5 | 5 | Hz | 3 | m | 0.003 | 3.0026 | 0.0048619 | -0.00062195 | 0.0046038 | m | 0.001 | 0.0030026 | 4.8619e-06 | -6.2195e-07 | 4.6038e-06 | 0.0030026 | 5Hz3m | 45 | 0.0029998 | 3e-08 | -1.8605e-07 | 1.5e-07 | 5Hz3m3m | -2.732e-06 | 0.0097238 | 4.359e-07 | 0.0092077
3 | m | 0.003 | 10 | 10 | Hz | 3 | m | 0.003 | 2.994 | 0.0020312 | -0.0068213 | 0.0035176 | m | 0.001 | 0.002994 | 2.0312e-06 | -6.8213e-06 | 3.5176e-06 | 0.0029941 | 10Hz3m | 46 | 0.003 | 3e-08 | -4.4661e-07 | 1.5e-07 | 10Hz3m3m | 5.9601e-06 | 0.0040625 | 6.3747e-06 | 0.0070352
3 | m | 0.003 | 20 | 20 | Hz | 3 | m | 0.003 | 2.9982 | 0.0046637 | 0.0048636 | 0.00033042 | m | 0.001 | 0.0029982 | 4.6637e-06 | 4.8636e-06 | 3.3042e-07 | 0.0029982 | 20Hz3m | 47 | 0.0030001 | 3e-08 | 1.7986e-07 | 1.5e-07 | 20Hz3m3m | 1.9055e-06 | 0.0093274 | -4.6838e-06 | 0.00066084
3 | m | 0.003 | 50 | 50 | Hz | 3 | m | 0.003 | 3.0008 | 0.0049686 | -0.00036121 | 0.0041203 | m | 0.001 | 0.0030008 | 4.9686e-06 | -3.6121e-07 | 4.1203e-06 | 0.0030008 | 50Hz3m | 48 | 0.0029998 | 3e-08 | 2.0252e-07 | 1.5e-07 | 50Hz3m3m | -9.748e-07 | 0.0099373 | 5.6373e-07 | 0.0082406
3 | m | 0.003 | 100 | 100 | Hz | 3 | m | 0.003 | 3.0059 | 0.004554 | -0.0023842 | 0.0014716 | m | 0.001 | 0.0030059 | 4.554e-06 | -2.3842e-06 | 1.4716e-06 | 0.0030059 | 100Hz3m | 49 | 0.0030012 | 3e-08 | 5.0344e-07 | 1.5e-07 | 100Hz3m3m | -4.6761e-06 | 0.0091079 | 2.8876e-06 | 0.0029432
3 | m | 0.003 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0048 | 0.0038227 | -0.002133 | 0.00077182 | m | 0.001 | 0.0030048 | 3.8227e-06 | -2.133e-06 | 7.7182e-07 | 0.0030048 | 200Hz3m | 50 | 0.0030018 | 3e-08 | 1.0905e-06 | 1.5e-07 | 200Hz3m3m | -2.9874e-06 | 0.0076454 | 3.2235e-06 | 0.0015436
3 | m | 0.003 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0121 | 0.0032205 | 0.000572 | 0.0033634 | m | 0.001 | 0.0030121 | 3.2205e-06 | 5.72e-07 | 3.3634e-06 | 0.0030121 | 500Hz3m | 51 | 0.0030079 | 4.7434e-08 | 1.6914e-06 | 1.5e-07 | 500Hz3m3m | -4.1615e-06 | 0.006441 | 1.1194e-06 | 0.0067268
3 | m | 0.003 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0225 | 0.0035187 | -0.00020224 | 0.0027545 | m | 0.001 | 0.0030225 | 3.5187e-06 | -2.0224e-07 | 2.7545e-06 | 0.0030225 | 1kHz3m | 52 | 0.0030214 | 1.3416e-07 | 3.3908e-06 | 3e-07 | 1kHz3m3m | -1.1055e-06 | 0.0070375 | 3.593e-06 | 0.0055089
3 | m | 0.003 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0575 | 0.0013564 | 0.012004 | 0.0036053 | m | 0.001 | 0.0030575 | 1.3564e-06 | 1.2004e-05 | 3.6053e-06 | 0.0030575 | 2kHz3m | 53 | 0.0030627 | 3.7947e-07 | 7.3349e-06 | 6e-07 | 2kHz3m3m | 5.1913e-06 | 0.0027128 | -4.6693e-06 | 0.0072106
3 | m | 0.003 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.2438 | 0.0018282 | 0.018644 | 0.0025191 | m | 0.001 | 0.0032438 | 1.8282e-06 | 1.8644e-05 | 2.5191e-06 | 0.0032439 | 5kHz3m | 54 | 0.0032478 | 1.5e-06 | 1.9038e-05 | 1.5e-06 | 5kHz3m3m | 3.9732e-06 | 0.0036563 | 3.9338e-07 | 0.0050382
10 | m | 0.01 | 0.01 | 10 | mHz | 3 | m | 0.003 | 2.9922 | 0.0049057 | 0.0051741 | 0.0035198 | m | 0.001 | 0.0029922 | 4.9057e-06 | 5.1741e-06 | 3.5198e-06 | 0.0029922 | 10mHz3m | 37 | 0.0029999 | 3e-08 | 1.8535e-07 | 1.5e-07 | 10mHz10m3m | 7.7028e-06 | 0.0098114 | -4.9888e-06 | 0.0070397
10 | m | 0.01 | 0.02 | 20 | mHz | 3 | m | 0.003 | 2.9979 | 0.0016305 | 0.00059738 | 0.00019594 | m | 0.001 | 0.0029979 | 1.6305e-06 | 5.9738e-07 | 1.9594e-07 | 0.0029979 | 20mHz3m | 38 | 0.003 | 3e-08 | -2.3489e-07 | 1.5e-07 | 20mHz10m3m | 2.1312e-06 | 0.003261 | -8.3227e-07 | 0.00039189
10 | m | 0.01 | 0.05 | 50 | mHz | 3 | m | 0.003 | 2.9975 | 0.0033176 | -0.0025865 | 0.0041556 | m | 0.001 | 0.0029975 | 3.3176e-06 | -2.5865e-06 | 4.1556e-06 | 0.0029975 | 50mHz3m | 39 | 0.0029997 | 3e-08 | -1.1531e-07 | 1.5e-07 | 50mHz10m3m | 2.1976e-06 | 0.0066352 | 2.4712e-06 | 0.0083113
10 | m | 0.01 | 0.1 | 100 | mHz | 3 | m | 0.003 | 3.0032 | 0.0038926 | -0.0047112 | 0.0012853 | m | 0.001 | 0.0030032 | 3.8926e-06 | -4.7112e-06 | 1.2853e-06 | 0.0030033 | 100mHz3m | 40 | 0.0030003 | 3e-08 | 1.1905e-07 | 1.5e-07 | 100mHz10m3m | -2.9039e-06 | 0.0077852 | 4.8302e-06 | 0.0025706
10 | m | 0.01 | 0.2 | 200 | mHz | 3 | m | 0.003 | 2.9951 | 0.0029222 | -0.0025444 | 0.0033319 | m | 0.001 | 0.0029951 | 2.9222e-06 | -2.5444e-06 | 3.3319e-06 | 0.0029951 | 200mHz3m | 41 | 0.0029996 | 3e-08 | -2.2468e-07 | 1.5e-07 | 200mHz10m3m | 4.5087e-06 | 0.0058445 | 2.3197e-06 | 0.0066638
10 | m | 0.01 | 0.5 | 500 | mHz | 3 | m | 0.003 | 2.998 | 0.002067 | 0.0050631 | 0.00024715 | m | 0.001 | 0.002998 | 2.067e-06 | 5.0631e-06 | 2.4715e-07 | 0.002998 | 500mHz3m | 42 | 0.003 | 3e-08 | -1.4963e-08 | 1.5e-07 | 500mHz10m3m | 2.0303e-06 | 0.004134 | -5.0781e-06 | 0.00049429
10 | m | 0.01 | 1 | 1 | Hz | 3 | m | 0.003 | 3.0021 | 0.0046572 | 0.0066478 | 0.0043527 | m | 0.001 | 0.0030021 | 4.6572e-06 | 6.6478e-06 | 4.3527e-06 | 0.0030021 | 1Hz3m | 43 | 0.0029996 | 3e-08 | -5.2952e-07 | 1.5e-07 | 1Hz10m3m | -2.4705e-06 | 0.0093143 | -7.1774e-06 | 0.0087054
10 | m | 0.01 | 2 | 2 | Hz | 3 | m | 0.003 | 3.0005 | 0.0039681 | 0.0013857 | 0.0019999 | m | 0.001 | 0.0030005 | 3.9681e-06 | 1.3857e-06 | 1.9999e-06 | 0.0030005 | 2Hz3m | 44 | 0.0029999 | 3e-08 | 3.5039e-07 | 1.5e-07 | 2Hz10m3m | -6.0987e-07 | 0.0079362 | -1.0353e-06 | 0.0039997
10 | m | 0.01 | 5 | 5 | Hz | 3 | m | 0.003 | 3.0017 | 0.00086257 | 0.0051723 | 0.00078178 | m | 0.001 | 0.0030017 | 8.6257e-07 | 5.1723e-06 | 7.8178e-07 | 0.0030017 | 5Hz3m | 45 | 0.0029998 | 3e-08 | -1.8605e-07 | 1.5e-07 | 5Hz10m3m | -1.8062e-06 | 0.0017251 | -5.3584e-06 | 0.0015636
10 | m | 0.01 | 10 | 10 | Hz | 3 | m | 0.003 | 3.0041 | 0.0034432 | -0.0037161 | 0.001679 | m | 0.001 | 0.0030041 | 3.4432e-06 | -3.7161e-06 | 1.679e-06 | 0.0030041 | 10Hz3m | 46 | 0.003 | 3e-08 | -4.4661e-07 | 1.5e-07 | 10Hz10m3m | -4.1407e-06 | 0.0068863 | 3.2695e-06 | 0.003358
10 | m | 0.01 | 20 | 20 | Hz | 3 | m | 0.003 | 2.9964 | 0.0048111 | -0.0043065 | 0.00043948 | m | 0.001 | 0.0029964 | 4.8111e-06 | -4.3065e-06 | 4.3948e-07 | 0.0029964 | 20Hz3m | 47 | 0.0030001 | 3e-08 | 1.7986e-07 | 1.5e-07 | 20Hz10m3m | 3.724e-06 | 0.0096223 | 4.4863e-06 | 0.00087897
10 | m | 0.01 | 50 | 50 | Hz | 3 | m | 0.003 | 3.0022 | 0.00091457 | 0.011848 | 0.0027879 | m | 0.001 | 0.0030022 | 9.1457e-07 | 1.1848e-05 | 2.7879e-06 | 0.0030022 | 50Hz3m | 48 | 0.0029998 | 3e-08 | 2.0252e-07 | 1.5e-07 | 50Hz10m3m | -2.328e-06 | 0.0018291 | -1.1645e-05 | 0.0055758
10 | m | 0.01 | 100 | 100 | Hz | 3 | m | 0.003 | 3 | 0.0046508 | -0.009206 | 0.00169 | m | 0.001 | 0.003 | 4.6508e-06 | -9.206e-06 | 1.69e-06 | 0.003 | 100Hz3m | 49 | 0.0030012 | 3e-08 | 5.0344e-07 | 1.5e-07 | 100Hz10m3m | 1.2373e-06 | 0.0093016 | 9.7094e-06 | 0.0033799
10 | m | 0.01 | 200 | 200 | Hz | 3 | m | 0.003 | 3.0033 | 0.0033713 | 0.0018849 | 0.0020221 | m | 0.001 | 0.0030033 | 3.3713e-06 | 1.8849e-06 | 2.0221e-06 | 0.0030033 | 200Hz3m | 50 | 0.0030018 | 3e-08 | 1.0905e-06 | 1.5e-07 | 200Hz10m3m | -1.5114e-06 | 0.0067426 | -7.9441e-07 | 0.0040442
10 | m | 0.01 | 500 | 500 | Hz | 3 | m | 0.003 | 3.0027 | 0.0039332 | -0.0058392 | 0.000271 | m | 0.001 | 0.0030027 | 3.9332e-06 | -5.8392e-06 | 2.71e-07 | 0.0030028 | 500Hz3m | 51 | 0.0030079 | 4.7434e-08 | 1.6914e-06 | 1.5e-07 | 500Hz10m3m | 5.1494e-06 | 0.0078664 | 7.5307e-06 | 0.00054201
10 | m | 0.01 | 1000 | 1 | kHz | 3 | m | 0.003 | 3.0194 | 0.0014119 | 0.0022174 | 0.0041791 | m | 0.001 | 0.0030194 | 1.4119e-06 | 2.2174e-06 | 4.1791e-06 | 0.0030194 | 1kHz3m | 52 | 0.0030214 | 1.3416e-07 | 3.3908e-06 | 3e-07 | 1kHz10m3m | 1.9599e-06 | 0.0028237 | 1.1734e-06 | 0.0083582
10 | m | 0.01 | 2000 | 2 | kHz | 3 | m | 0.003 | 3.0693 | 0.00082518 | 0.010795 | 0.0014427 | m | 0.001 | 0.0030693 | 8.2518e-07 | 1.0795e-05 | 1.4427e-06 | 0.0030693 | 2kHz3m | 53 | 0.0030627 | 3.7947e-07 | 7.3349e-06 | 6e-07 | 2kHz10m3m | -6.5665e-06 | 0.0016504 | -3.4596e-06 | 0.0028854
10 | m | 0.01 | 5000 | 5 | kHz | 3 | m | 0.003 | 3.2501 | 0.0030014 | 0.021902 | 0.002766 | m | 0.001 | 0.0032501 | 3.0014e-06 | 2.1902e-05 | 2.766e-06 | 0.0032502 | 5kHz3m | 54 | 0.0032478 | 1.5e-06 | 1.9038e-05 | 1.5e-06 | 5kHz10m3m | -2.3167e-06 | 0.0060027 | -2.8645e-06 | 0.0055319
10 | m | 0.01 | 0.01 | 10 | mHz | 10 | m | 0.01 | 10.0012 | 0.0031316 | 0.0087339 | 0.0021293 | m | 0.001 | 0.010001 | 3.1316e-06 | 8.7339e-06 | 2.1293e-06 | 0.010001 | 10mHz10m | 55 | 0.010001 | 1e-07 | 4.0086e-07 | 5e-07 | 10mHz10m10m | -2.8151e-07 | 0.0062631 | -8.333e-06 | 0.0042586
10 | m | 0.01 | 0.02 | 20 | mHz | 10 | m | 0.01 | 9.997 | 0.0013249 | 0.0078008 | 7.8571e-05 | m | 0.001 | 0.009997 | 1.3249e-06 | 7.8008e-06 | 7.8571e-08 | 0.009997 | 20mHz10m | 56 | 0.0099999 | 1e-07 | 6.4265e-07 | 5e-07 | 20mHz10m10m | 2.9426e-06 | 0.0026498 | -7.1581e-06 | 0.00015714
10 | m | 0.01 | 0.05 | 50 | mHz | 10 | m | 0.01 | 9.9967 | 0.0025713 | 0.0067015 | 0.0013417 | m | 0.001 | 0.0099967 | 2.5713e-06 | 6.7015e-06 | 1.3417e-06 | 0.0099967 | 50mHz10m | 57 | 0.010001 | 1e-07 | -3.2046e-07 | 5e-07 | 50mHz10m10m | 4.3218e-06 | 0.0051427 | -7.0219e-06 | 0.0026834
10 | m | 0.01 | 0.1 | 100 | mHz | 10 | m | 0.01 | 10.001 | 0.0029731 | 0.002022 | 0.0046017 | m | 0.001 | 0.010001 | 2.9731e-06 | 2.022e-06 | 4.6017e-06 | 0.010001 | 100mHz10m | 58 | 0.0099989 | 1e-07 | 2.8477e-07 | 5e-07 | 100mHz10m10m | -2.1867e-06 | 0.0059461 | -1.7372e-06 | 0.0092034
10 | m | 0.01 | 0.2 | 200 | mHz | 10 | m | 0.01 | 10.0114 | 0.00019523 | -0.0027063 | 0.0015634 | m | 0.001 | 0.010011 | 1.9523e-07 | -2.7063e-06 | 1.5634e-06 | 0.010011 | 200mHz10m | 59 | 0.010001 | 1e-07 | -2.6436e-06 | 5e-07 | 200mHz10m10m | -1.01e-05 | 0.00039046 | 6.262e-08 | 0.0031267
10 | m | 0.01 | 0.5 | 500 | mHz | 10 | m | 0.01 | 9.999 | 0.0021812 | -0.0034303 | 0.003089 | m | 0.001 | 0.009999 | 2.1812e-06 | -3.4303e-06 | 3.089e-06 | 0.009999 | 500mHz10m | 60 | 0.010001 | 1e-07 | 2.5137e-07 | 5e-07 | 500mHz10m10m | 2.3588e-06 | 0.0043624 | 3.6816e-06 | 0.006178
10 | m | 0.01 | 1 | 1 | Hz | 10 | m | 0.01 | 10.0027 | 0.0041566 | 0.0019839 | 0.00030826 | m | 0.001 | 0.010003 | 4.1566e-06 | 1.9839e-06 | 3.0826e-07 | 0.010003 | 1Hz10m | 61 | 0.010001 | 1e-07 | 8.7214e-07 | 5e-07 | 1Hz10m10m | -2.0327e-06 | 0.0083131 | -1.1117e-06 | 0.00061653
10 | m | 0.01 | 2 | 2 | Hz | 10 | m | 0.01 | 10.0001 | 0.0020338 | 0.010949 | 0.0029564 | m | 0.001 | 0.01 | 2.0338e-06 | 1.0949e-05 | 2.9564e-06 | 0.01 | 2Hz10m | 62 | 0.0099992 | 1e-07 | 2.9561e-06 | 5e-07 | 2Hz10m10m | -8.5536e-07 | 0.0040675 | -7.9933e-06 | 0.0059128
10 | m | 0.01 | 5 | 5 | Hz | 10 | m | 0.01 | 10.0032 | 0.0042902 | 0.0086725 | 0.00099903 | m | 0.001 | 0.010003 | 4.2902e-06 | 8.6725e-06 | 9.9903e-07 | 0.010003 | 5Hz10m | 63 | 0.0099994 | 1e-07 | 1.6976e-07 | 5e-07 | 5Hz10m10m | -3.7766e-06 | 0.0085804 | -8.5028e-06 | 0.0019981
10 | m | 0.01 | 10 | 10 | Hz | 10 | m | 0.01 | 10.0043 | 0.0019091 | -0.0044188 | 0.0047812 | m | 0.001 | 0.010004 | 1.9091e-06 | -4.4188e-06 | 4.7812e-06 | 0.010004 | 10Hz10m | 64 | 0.0099996 | 1e-07 | -3.3252e-07 | 5e-07 | 10Hz10m10m | -4.7765e-06 | 0.0038182 | 4.0862e-06 | 0.0095625
10 | m | 0.01 | 20 | 20 | Hz | 10 | m | 0.01 | 9.9914 | 0.0025581 | 0.0025671 | 0.0041909 | m | 0.001 | 0.0099914 | 2.5581e-06 | 2.5671e-06 | 4.1909e-06 | 0.0099914 | 20Hz10m | 65 | 0.010001 | 1e-07 | -1.3391e-07 | 5e-07 | 20Hz10m10m | 9.1946e-06 | 0.0051162 | -2.701e-06 | 0.0083818
10 | m | 0.01 | 50 | 50 | Hz | 10 | m | 0.01 | 9.9996 | 0.001441 | -0.0050911 | 0.00040394 | m | 0.001 | 0.0099996 | 1.441e-06 | -5.0911e-06 | 4.0394e-07 | 0.0099996 | 50Hz10m | 66 | 0.010001 | 1e-07 | 1.1725e-06 | 5e-07 | 50Hz10m10m | 1.7163e-06 | 0.0028821 | 6.2636e-06 | 0.00080787
10 | m | 0.01 | 100 | 100 | Hz | 10 | m | 0.01 | 9.9983 | 0.0040969 | 0.0037506 | 0.0015333 | m | 0.001 | 0.0099983 | 4.0969e-06 | 3.7506e-06 | 1.5333e-06 | 0.0099983 | 100Hz10m | 67 | 0.010003 | 1e-07 | 2.5917e-07 | 5e-07 | 100Hz10m10m | 4.8106e-06 | 0.0081937 | -3.4914e-06 | 0.0030666
10 | m | 0.01 | 200 | 200 | Hz | 10 | m | 0.01 | 10.0027 | 0.001696 | 0.0019076 | 0.0047513 | m | 0.001 | 0.010003 | 1.696e-06 | 1.9076e-06 | 4.7513e-06 | 0.010003 | 200Hz10m | 68 | 0.010007 | 1e-07 | 2.5996e-06 | 5e-07 | 200Hz10m10m | 3.9822e-06 | 0.0033919 | 6.9205e-07 | 0.0095026
10 | m | 0.01 | 500 | 500 | Hz | 10 | m | 0.01 | 10.0222 | 0.0025626 | 0.00049795 | 0.0036475 | m | 0.001 | 0.010022 | 2.5626e-06 | 4.9795e-07 | 3.6475e-06 | 0.010022 | 500Hz10m | 69 | 0.010024 | 1.5811e-07 | 6.4006e-06 | 5e-07 | 500Hz10m10m | 2.0292e-06 | 0.0051253 | 5.9026e-06 | 0.0072951
10 | m | 0.01 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.0798 | 0.0044286 | 0.013182 | 0.0024834 | m | 0.001 | 0.01008 | 4.4286e-06 | 1.3182e-05 | 2.4834e-06 | 0.01008 | 1kHz10m | 70 | 0.010074 | 4.4721e-07 | 1.2677e-05 | 1e-06 | 1kHz10m10m | -5.2726e-06 | 0.0088571 | -5.0497e-07 | 0.0049668
10 | m | 0.01 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.2116 | 0.0032016 | 0.028294 | 0.00081092 | m | 0.001 | 0.010212 | 3.2016e-06 | 2.8294e-05 | 8.1092e-07 | 0.010212 | 2kHz10m | 71 | 0.010208 | 1.2649e-06 | 2.6377e-05 | 2e-06 | 2kHz10m10m | -3.8097e-06 | 0.0064033 | -1.9173e-06 | 0.0016218
10 | m | 0.01 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.8313 | 0.0032429 | 0.072745 | 0.0046298 | m | 0.001 | 0.010831 | 3.2429e-06 | 7.2745e-05 | 4.6298e-06 | 0.010832 | 5kHz10m | 72 | 0.010827 | 5e-06 | 6.2068e-05 | 5e-06 | 5kHz10m10m | -4.3437e-06 | 0.0064858 | -1.0677e-05 | 0.0092597
100 | m | 0.1 | 0.01 | 10 | mHz | 10 | m | 0.01 | 10.0006 | 0.0018633 | 0.0074457 | 0.0046635 | m | 0.001 | 0.010001 | 1.8633e-06 | 7.4457e-06 | 4.6635e-06 | 0.010001 | 10mHz10m | 55 | 0.010001 | 1e-07 | 4.0086e-07 | 5e-07 | 10mHz100m10m | 3.5728e-07 | 0.0037267 | -7.0448e-06 | 0.0093271
100 | m | 0.1 | 0.02 | 20 | mHz | 10 | m | 0.01 | 10.001 | 0.0016499 | -0.0010097 | 0.00023338 | m | 0.001 | 0.010001 | 1.6499e-06 | -1.0097e-06 | 2.3338e-07 | 0.010001 | 20mHz10m | 56 | 0.0099999 | 1e-07 | 6.4265e-07 | 5e-07 | 20mHz100m10m | -1.0903e-06 | 0.0032999 | 1.6524e-06 | 0.00046677
100 | m | 0.1 | 0.05 | 50 | mHz | 10 | m | 0.01 | 10.004 | 0.0035923 | 0.00033942 | 0.0036536 | m | 0.001 | 0.010004 | 3.5923e-06 | 3.3942e-07 | 3.6536e-06 | 0.010004 | 50mHz10m | 57 | 0.010001 | 1e-07 | -3.2046e-07 | 5e-07 | 50mHz100m10m | -2.9309e-06 | 0.0071846 | -6.5988e-07 | 0.0073072
100 | m | 0.1 | 0.1 | 100 | mHz | 10 | m | 0.01 | 9.9983 | 0.0013298 | -0.0007204 | 0.0018912 | m | 0.001 | 0.0099983 | 1.3298e-06 | -7.204e-07 | 1.8912e-06 | 0.0099983 | 100mHz10m | 58 | 0.0099989 | 1e-07 | 2.8477e-07 | 5e-07 | 100mHz100m10m | 5.188e-07 | 0.0026596 | 1.0052e-06 | 0.0037823
100 | m | 0.1 | 0.2 | 200 | mHz | 10 | m | 0.01 | 10.0068 | 0.0011516 | -0.0021034 | 0.0043258 | m | 0.001 | 0.010007 | 1.1516e-06 | -2.1034e-06 | 4.3258e-06 | 0.010007 | 200mHz10m | 59 | 0.010001 | 1e-07 | -2.6436e-06 | 5e-07 | 200mHz100m10m | -5.4865e-06 | 0.0023032 | -5.4027e-07 | 0.0086516
100 | m | 0.1 | 0.5 | 500 | mHz | 10 | m | 0.01 | 9.9939 | 0.0026785 | 0.0049376 | 0.0033369 | m | 0.001 | 0.0099939 | 2.6785e-06 | 4.9376e-06 | 3.3369e-06 | 0.0099939 | 500mHz10m | 60 | 0.010001 | 1e-07 | 2.5137e-07 | 5e-07 | 500mHz100m10m | 7.4303e-06 | 0.0053571 | -4.6862e-06 | 0.0066737
100 | m | 0.1 | 1 | 1 | Hz | 10 | m | 0.01 | 9.9984 | 0.00027127 | 0.00029066 | 0.0014837 | m | 0.001 | 0.0099984 | 2.7127e-07 | 2.9066e-07 | 1.4837e-06 | 0.0099984 | 1Hz10m | 61 | 0.010001 | 1e-07 | 8.7214e-07 | 5e-07 | 1Hz100m10m | 2.3046e-06 | 0.00054254 | 5.8148e-07 | 0.0029675
100 | m | 0.1 | 2 | 2 | Hz | 10 | m | 0.01 | 9.9883 | 0.0041846 | 0.0076312 | 0.0038355 | m | 0.001 | 0.0099883 | 4.1846e-06 | 7.6312e-06 | 3.8355e-06 | 0.0099883 | 2Hz10m | 62 | 0.0099992 | 1e-07 | 2.9561e-06 | 5e-07 | 2Hz100m10m | 1.0865e-05 | 0.0083693 | -4.6751e-06 | 0.0076711
100 | m | 0.1 | 5 | 5 | Hz | 10 | m | 0.01 | 10.0034 | 0.0048785 | -0.0025884 | 0.0042826 | m | 0.001 | 0.010003 | 4.8785e-06 | -2.5884e-06 | 4.2826e-06 | 0.010003 | 5Hz10m | 63 | 0.0099994 | 1e-07 | 1.6976e-07 | 5e-07 | 5Hz100m10m | -4.0641e-06 | 0.009757 | 2.7582e-06 | 0.0085653
100 | m | 0.1 | 10 | 10 | Hz | 10 | m | 0.01 | 9.999 | 0.00012046 | -0.0057758 | 0.00327 | m | 0.001 | 0.009999 | 1.2046e-07 | -5.7758e-06 | 3.27e-06 | 0.009999 | 10Hz10m | 64 | 0.0099996 | 1e-07 | -3.3252e-07 | 5e-07 | 10Hz100m10m | 5.8756e-07 | 0.00024092 | 5.4433e-06 | 0.00654
100 | m | 0.1 | 20 | 20 | Hz | 10 | m | 0.01 | 9.9988 | 0.0042842 | 0.0019977 | 0.0016666 | m | 0.001 | 0.0099988 | 4.2842e-06 | 1.9977e-06 | 1.6666e-06 | 0.0099988 | 20Hz10m | 65 | 0.010001 | 1e-07 | -1.3391e-07 | 5e-07 | 20Hz100m10m | 1.7957e-06 | 0.0085685 | -2.1316e-06 | 0.0033332
100 | m | 0.1 | 50 | 50 | Hz | 10 | m | 0.01 | 10.0075 | 0.0040079 | 0.0012779 | 0.0042701 | m | 0.001 | 0.010008 | 4.0079e-06 | 1.2779e-06 | 4.2701e-06 | 0.010008 | 50Hz10m | 66 | 0.010001 | 1e-07 | 1.1725e-06 | 5e-07 | 50Hz100m10m | -6.2662e-06 | 0.0080157 | -1.0542e-07 | 0.0085401
100 | m | 0.1 | 100 | 100 | Hz | 10 | m | 0.01 | 10.0005 | 0.00031678 | 0.00418 | 0.0028063 | m | 0.001 | 0.010001 | 3.1678e-07 | 4.18e-06 | 2.8063e-06 | 0.010001 | 100Hz10m | 67 | 0.010003 | 1e-07 | 2.5917e-07 | 5e-07 | 100Hz100m10m | 2.5938e-06 | 0.00063357 | -3.9209e-06 | 0.0056125
100 | m | 0.1 | 200 | 200 | Hz | 10 | m | 0.01 | 10.0031 | 0.0021283 | 0.0043522 | 0.0012572 | m | 0.001 | 0.010003 | 2.1283e-06 | 4.3522e-06 | 1.2572e-06 | 0.010003 | 200Hz10m | 68 | 0.010007 | 1e-07 | 2.5996e-06 | 5e-07 | 200Hz100m10m | 3.5932e-06 | 0.0042566 | -1.7526e-06 | 0.0025144
100 | m | 0.1 | 500 | 500 | Hz | 10 | m | 0.01 | 10.0261 | 0.0049138 | 0.0061357 | 0.0040581 | m | 0.001 | 0.010026 | 4.9138e-06 | 6.1357e-06 | 4.0581e-06 | 0.010026 | 500Hz10m | 69 | 0.010024 | 1.5811e-07 | 6.4006e-06 | 5e-07 | 500Hz100m10m | -1.8758e-06 | 0.0098276 | 2.6486e-07 | 0.0081162
100 | m | 0.1 | 1000 | 1 | kHz | 10 | m | 0.01 | 10.077 | 0.0028037 | 0.0061492 | 0.0045902 | m | 0.001 | 0.010077 | 2.8037e-06 | 6.1492e-06 | 4.5902e-06 | 0.010077 | 1kHz10m | 70 | 0.010074 | 4.4721e-07 | 1.2677e-05 | 1e-06 | 1kHz100m10m | -2.5073e-06 | 0.0056074 | 6.5281e-06 | 0.0091804
100 | m | 0.1 | 2000 | 2 | kHz | 10 | m | 0.01 | 10.2071 | 0.0022304 | 0.017004 | 0.00023419 | m | 0.001 | 0.010207 | 2.2304e-06 | 1.7004e-05 | 2.3419e-07 | 0.010207 | 2kHz10m | 71 | 0.010208 | 1.2649e-06 | 2.6377e-05 | 2e-06 | 2kHz100m10m | 7.1951e-07 | 0.0044607 | 9.3729e-06 | 0.00046839
100 | m | 0.1 | 5000 | 5 | kHz | 10 | m | 0.01 | 10.8238 | 0.002715 | 0.062951 | 0.00081409 | m | 0.001 | 0.010824 | 2.715e-06 | 6.2951e-05 | 8.1409e-07 | 0.010824 | 5kHz10m | 72 | 0.010827 | 5e-06 | 6.2068e-05 | 5e-06 | 5kHz100m10m | 3.1651e-06 | 0.0054299 | -8.8348e-07 | 0.0016282
100 | m | 0.1 | 0.01 | 10 | mHz | 100 | m | 0.1 | 100.011 | 0.0027482 | 0.022506 | 0.001258 | m | 0.001 | 0.10001 | 2.7482e-06 | 2.2506e-05 | 1.258e-06 | 0.10001 | 10mHz100m | 73 | 0.10001 | 1e-06 | 2.1959e-05 | 5e-06 | 10mHz100m100m | -7.2913e-07 | 0.0054964 | -5.4688e-07 | 0.0025161
100 | m | 0.1 | 0.02 | 20 | mHz | 100 | m | 0.1 | 99.9781 | 0.0029606 | -0.014915 | 0.0043275 | m | 0.001 | 0.099978 | 2.9606e-06 | -1.4915e-05 | 4.3275e-06 | 0.099978 | 20mHz100m | 74 | 0.099991 | 1e-06 | -3.0977e-06 | 5e-06 | 20mHz100m100m | 1.3173e-05 | 0.0059211 | 1.1817e-05 | 0.0086549
100 | m | 0.1 | 0.05 | 50 | mHz | 100 | m | 0.1 | 100.0151 | 0.0015499 | -0.013869 | 0.0011627 | m | 0.001 | 0.10002 | 1.5499e-06 | -1.3869e-05 | 1.1627e-06 | 0.10002 | 50mHz100m | 75 | 0.099999 | 1e-06 | -1.193e-05 | 5e-06 | 50mHz100m100m | -1.5712e-05 | 0.0030998 | 1.9391e-06 | 0.0023254
100 | m | 0.1 | 0.1 | 100 | mHz | 100 | m | 0.1 | 100.0318 | 0.004033 | 0.0068751 | 0.0045699 | m | 0.001 | 0.10003 | 4.033e-06 | 6.8751e-06 | 4.5699e-06 | 0.10003 | 100mHz100m | 76 | 0.10002 | 1e-06 | 1.6399e-06 | 5e-06 | 100mHz100m100m | -1.2187e-05 | 0.0080661 | -5.2352e-06 | 0.0091398
100 | m | 0.1 | 0.2 | 200 | mHz | 100 | m | 0.1 | 99.9865 | 0.0016907 | -0.0033454 | 0.0038799 | m | 0.001 | 0.099987 | 1.6907e-06 | -3.3454e-06 | 3.8799e-06 | 0.099987 | 200mHz100m | 77 | 0.099993 | 1e-06 | -1.0265e-06 | 5e-06 | 200mHz100m100m | 6.0537e-06 | 0.0033814 | 2.319e-06 | 0.0077598
100 | m | 0.1 | 0.5 | 500 | mHz | 100 | m | 0.1 | 100.0088 | 0.0042711 | -0.01787 | 0.0011544 | m | 0.001 | 0.10001 | 4.2711e-06 | -1.787e-05 | 1.1544e-06 | 0.10001 | 500mHz100m | 78 | 0.10001 | 1e-06 | -1.3826e-05 | 5e-06 | 500mHz100m100m | 5.2662e-06 | 0.0085422 | 4.0437e-06 | 0.0023089
100 | m | 0.1 | 1 | 1 | Hz | 100 | m | 0.1 | 100.0095 | 0.0016923 | -0.010949 | 0.0032619 | m | 0.001 | 0.10001 | 1.6923e-06 | -1.0949e-05 | 3.2619e-06 | 0.10001 | 1Hz100m | 79 | 0.10001 | 1e-06 | -4.7865e-06 | 5e-06 | 1Hz100m100m | 2.0044e-06 | 0.0033847 | 6.1626e-06 | 0.0065239
100 | m | 0.1 | 2 | 2 | Hz | 100 | m | 0.1 | 99.9951 | 0.0051377 | -0.0071932 | 0.00039541 | m | 0.001 | 0.099995 | 5.1377e-06 | -7.1932e-06 | 3.9541e-07 | 0.099995 | 2Hz100m | 80 | 0.10001 | 1e-06 | -8.6121e-07 | 5e-06 | 2Hz100m100m | 1.5456e-05 | 0.010275 | 6.332e-06 | 0.00079084
100 | m | 0.1 | 5 | 5 | Hz | 100 | m | 0.1 | 100.0059 | 0.0035421 | -0.012735 | 0.00039836 | m | 0.001 | 0.10001 | 3.5421e-06 | -1.2735e-05 | 3.9836e-07 | 0.10001 | 5Hz100m | 81 | 0.10001 | 1e-06 | -1.1974e-05 | 5e-06 | 5Hz100m100m | 6.3364e-07 | 0.0070843 | 7.604e-07 | 0.00079673
100 | m | 0.1 | 10 | 10 | Hz | 100 | m | 0.1 | 100.0066 | 0.0036381 | -0.023156 | 0.0033161 | m | 0.001 | 0.10001 | 3.6381e-06 | -2.3156e-05 | 3.3161e-06 | 0.10001 | 10Hz100m | 82 | 0.1 | 1e-06 | -3.2911e-05 | 5e-06 | 10Hz100m100m | -6.4252e-06 | 0.0072761 | -9.7551e-06 | 0.0066323
100 | m | 0.1 | 20 | 20 | Hz | 100 | m | 0.1 | 100.0217 | 0.0048568 | -0.016951 | 0.0039288 | m | 0.001 | 0.10002 | 4.8568e-06 | -1.6951e-05 | 3.9288e-06 | 0.10002 | 20Hz100m | 83 | 0.10001 | 1e-06 | -1.1151e-05 | 5e-06 | 20Hz100m100m | -7.2992e-06 | 0.0097136 | 5.8002e-06 | 0.0078576
100 | m | 0.1 | 50 | 50 | Hz | 100 | m | 0.1 | 99.9768 | 0.0050216 | 0.00032917 | 0.001709 | m | 0.001 | 0.099977 | 5.0216e-06 | 3.2917e-07 | 1.709e-06 | 0.099977 | 50Hz100m | 84 | 0.099986 | 1e-06 | -3.4939e-06 | 5e-06 | 50Hz100m100m | 9.4487e-06 | 0.010043 | -3.8231e-06 | 0.003418
100 | m | 0.1 | 100 | 100 | Hz | 100 | m | 0.1 | 100.0158 | 0.0054863 | 0.0092629 | 0.003584 | m | 0.001 | 0.10002 | 5.4863e-06 | 9.2629e-06 | 3.584e-06 | 0.10002 | 100Hz100m | 85 | 0.10002 | 1e-06 | 7.221e-06 | 5e-06 | 100Hz100m100m | 4.0011e-06 | 0.010973 | -2.0418e-06 | 0.007168
100 | m | 0.1 | 200 | 200 | Hz | 100 | m | 0.1 | 100.0765 | 0.002938 | 0.018419 | 0.0007608 | m | 0.001 | 0.10008 | 2.938e-06 | 1.8419e-05 | 7.608e-07 | 0.10008 | 200Hz100m | 86 | 0.10007 | 1e-06 | 1.4824e-05 | 5e-06 | 200Hz100m100m | -2.277e-06 | 0.0058761 | -3.5953e-06 | 0.0015216
100 | m | 0.1 | 500 | 500 | Hz | 100 | m | 0.1 | 100.2749 | 0.0044716 | 0.064454 | 0.0036617 | m | 0.001 | 0.10027 | 4.4716e-06 | 6.4454e-05 | 3.6617e-06 | 0.10027 | 500Hz100m | 87 | 0.10027 | 1.5811e-06 | 5.9739e-05 | 5e-06 | 500Hz100m100m | -4.7634e-06 | 0.0089432 | -4.7152e-06 | 0.0073233
100 | m | 0.1 | 1000 | 1 | kHz | 100 | m | 0.1 | 100.7595 | 0.0057592 | 0.13147 | 2.4066e-05 | m | 0.001 | 0.10076 | 5.7592e-06 | 0.00013147 | 2.4066e-08 | 0.10076 | 1kHz100m | 88 | 0.10075 | 4.4721e-06 | 0.00012542 | 1e-05 | 1kHz100m100m | -5.0663e-06 | 0.011518 | -6.0504e-06 | 4.916e-05
100 | m | 0.1 | 2000 | 2 | kHz | 100 | m | 0.1 | 102.0785 | 0.0040308 | 0.24626 | 0.001454 | m | 0.001 | 0.10208 | 4.0308e-06 | 0.00024626 | 1.454e-06 | 0.10208 | 2kHz100m | 89 | 0.10209 | 1.2649e-05 | 0.00024727 | 2e-05 | 2kHz100m100m | 8.7809e-06 | 0.0080617 | 1.0075e-06 | 0.002908
100 | m | 0.1 | 5000 | 5 | kHz | 100 | m | 0.1 | 108.2593 | 0.0056654 | 0.61413 | 0.00076567 | m | 0.001 | 0.10826 | 5.6654e-06 | 0.00061413 | 7.6567e-07 | 0.10826 | 5kHz100m | 90 | 0.10827 | 5e-05 | 0.00061379 | 5e-05 | 5kHz100m100m | 1.0411e-05 | 0.011331 | -3.3759e-07 | 0.0015322

## Ref Z list
Ref. Z list
¦Z¦ |  | ¦Z¦ | Z tag
Ohm |  | Ohm | -
0 | m | 0 | 0m
1 | m | 0.001 | 1m
3 | m | 0.003 | 3m
10 | m | 0.01 | 10m
100 | m | 0.1 | 100m

## Ref Z
This is sheet containing calibration data of reference standards to be used for calibration of RLC meter.
f | f |  | Nominal Z |  | Nom Z | Rs | U(Rs) | Xs | U(Xs) | Rs-Xs unit | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | tag
Hz | - |  | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | -
0.01 | 10 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10mHz0m
0.02 | 20 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20mHz0m
0.05 | 50 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50mHz0m
0.1 | 100 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100mHz0m
0.2 | 200 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200mHz0m
0.5 | 500 | mHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500mHz0m
1 | 1 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1Hz0m
2 | 2 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2Hz0m
5 | 5 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5Hz0m
10 | 10 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 10Hz0m
20 | 20 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 20Hz0m
50 | 50 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 50Hz0m
100 | 100 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 100Hz0m
200 | 200 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 200Hz0m
500 | 500 | Hz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 500Hz0m
1000 | 1 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 1kHz0m
2000 | 2 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 2kHz0m
5000 | 5 | kHz | 0 | m | 0 | 0 | 0 | 0 | 0 | m | 0.001 | 0 | 0 | 0 | 0 | 5kHz0m
0.01 | 10 | mHz | 1 | m | 0.001 | 0.99993 | 1e-05 | -0.00011463 | 5e-05 | m | 0.001 | 0.00099993 | 1e-08 | -1.1463e-07 | 5e-08 | 10mHz1m
0.02 | 20 | mHz | 1 | m | 0.001 | 0.99987 | 1e-05 | 8.6308e-05 | 5e-05 | m | 0.001 | 0.00099987 | 1e-08 | 8.6308e-08 | 5e-08 | 20mHz1m
0.05 | 50 | mHz | 1 | m | 0.001 | 0.99979 | 1e-05 | 8.9699e-05 | 5e-05 | m | 0.001 | 0.00099979 | 1e-08 | 8.9699e-08 | 5e-08 | 50mHz1m
0.1 | 100 | mHz | 1 | m | 0.001 | 0.99986 | 1e-05 | -1.692e-05 | 5e-05 | m | 0.001 | 0.00099986 | 1e-08 | -1.692e-08 | 5e-08 | 100mHz1m
0.2 | 200 | mHz | 1 | m | 0.001 | 0.99986 | 1e-05 | -6.4602e-05 | 5e-05 | m | 0.001 | 0.00099986 | 1e-08 | -6.4602e-08 | 5e-08 | 200mHz1m
0.5 | 500 | mHz | 1 | m | 0.001 | 0.99999 | 1e-05 | -0.0001903 | 5e-05 | m | 0.001 | 0.00099999 | 1e-08 | -1.903e-07 | 5e-08 | 500mHz1m
1 | 1 | Hz | 1 | m | 0.001 | 1 | 1e-05 | 3.7586e-05 | 5e-05 | m | 0.001 | 0.001 | 1e-08 | 3.7586e-08 | 5e-08 | 1Hz1m
2 | 2 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | 1.4693e-05 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 1.4693e-08 | 5e-08 | 2Hz1m
5 | 5 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | -0.00020967 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | -2.0967e-07 | 5e-08 | 5Hz1m
10 | 10 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | -6.2055e-06 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | -6.2055e-09 | 5e-08 | 10Hz1m
20 | 20 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | -8.0442e-05 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | -8.0442e-08 | 5e-08 | 20Hz1m
50 | 50 | Hz | 1 | m | 0.001 | 1.0002 | 1e-05 | -4.9127e-05 | 5e-05 | m | 0.001 | 0.0010002 | 1e-08 | -4.9127e-08 | 5e-08 | 50Hz1m
100 | 100 | Hz | 1 | m | 0.001 | 1.0001 | 1e-05 | 0.00023187 | 5e-05 | m | 0.001 | 0.0010001 | 1e-08 | 2.3187e-07 | 5e-08 | 100Hz1m
200 | 200 | Hz | 1 | m | 0.001 | 1.0008 | 1e-05 | 4.2778e-05 | 5e-05 | m | 0.001 | 0.0010008 | 1e-08 | 4.2778e-08 | 5e-08 | 200Hz1m
500 | 500 | Hz | 1 | m | 0.001 | 1.0026 | 1.5811e-05 | 0.00081299 | 5e-05 | m | 0.001 | 0.0010026 | 1.5811e-08 | 8.1299e-07 | 5e-08 | 500Hz1m
1000 | 1 | kHz | 1 | m | 0.001 | 1.0073 | 4.4721e-05 | 0.0012147 | 0.0001 | m | 0.001 | 0.0010073 | 4.4721e-08 | 1.2147e-06 | 1e-07 | 1kHz1m
2000 | 2 | kHz | 1 | m | 0.001 | 1.0208 | 0.00012649 | 0.0025962 | 0.0002 | m | 0.001 | 0.0010208 | 1.2649e-07 | 2.5962e-06 | 2e-07 | 2kHz1m
5000 | 5 | kHz | 1 | m | 0.001 | 1.0827 | 0.0005 | 0.006277 | 0.0005 | m | 0.001 | 0.0010827 | 5e-07 | 6.277e-06 | 5e-07 | 5kHz1m
0.01 | 10 | mHz | 3 | m | 0.003 | 2.9999 | 3e-05 | 0.00018535 | 0.00015 | m | 0.001 | 0.0029999 | 3e-08 | 1.8535e-07 | 1.5e-07 | 10mHz3m
0.02 | 20 | mHz | 3 | m | 0.003 | 3 | 3e-05 | -0.00023489 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | -2.3489e-07 | 1.5e-07 | 20mHz3m
0.05 | 50 | mHz | 3 | m | 0.003 | 2.9997 | 3e-05 | -0.00011531 | 0.00015 | m | 0.001 | 0.0029997 | 3e-08 | -1.1531e-07 | 1.5e-07 | 50mHz3m
0.1 | 100 | mHz | 3 | m | 0.003 | 3.0003 | 3e-05 | 0.00011905 | 0.00015 | m | 0.001 | 0.0030003 | 3e-08 | 1.1905e-07 | 1.5e-07 | 100mHz3m
0.2 | 200 | mHz | 3 | m | 0.003 | 2.9996 | 3e-05 | -0.00022468 | 0.00015 | m | 0.001 | 0.0029996 | 3e-08 | -2.2468e-07 | 1.5e-07 | 200mHz3m
0.5 | 500 | mHz | 3 | m | 0.003 | 3 | 3e-05 | -1.4963e-05 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | -1.4963e-08 | 1.5e-07 | 500mHz3m
1 | 1 | Hz | 3 | m | 0.003 | 2.9996 | 3e-05 | -0.00052952 | 0.00015 | m | 0.001 | 0.0029996 | 3e-08 | -5.2952e-07 | 1.5e-07 | 1Hz3m
2 | 2 | Hz | 3 | m | 0.003 | 2.9999 | 3e-05 | 0.00035039 | 0.00015 | m | 0.001 | 0.0029999 | 3e-08 | 3.5039e-07 | 1.5e-07 | 2Hz3m
5 | 5 | Hz | 3 | m | 0.003 | 2.9998 | 3e-05 | -0.00018605 | 0.00015 | m | 0.001 | 0.0029998 | 3e-08 | -1.8605e-07 | 1.5e-07 | 5Hz3m
10 | 10 | Hz | 3 | m | 0.003 | 3 | 3e-05 | -0.00044661 | 0.00015 | m | 0.001 | 0.003 | 3e-08 | -4.4661e-07 | 1.5e-07 | 10Hz3m
20 | 20 | Hz | 3 | m | 0.003 | 3.0001 | 3e-05 | 0.00017986 | 0.00015 | m | 0.001 | 0.0030001 | 3e-08 | 1.7986e-07 | 1.5e-07 | 20Hz3m
50 | 50 | Hz | 3 | m | 0.003 | 2.9998 | 3e-05 | 0.00020252 | 0.00015 | m | 0.001 | 0.0029998 | 3e-08 | 2.0252e-07 | 1.5e-07 | 50Hz3m
100 | 100 | Hz | 3 | m | 0.003 | 3.0012 | 3e-05 | 0.00050344 | 0.00015 | m | 0.001 | 0.0030012 | 3e-08 | 5.0344e-07 | 1.5e-07 | 100Hz3m
200 | 200 | Hz | 3 | m | 0.003 | 3.0018 | 3e-05 | 0.0010905 | 0.00015 | m | 0.001 | 0.0030018 | 3e-08 | 1.0905e-06 | 1.5e-07 | 200Hz3m
500 | 500 | Hz | 3 | m | 0.003 | 3.0079 | 4.7434e-05 | 0.0016914 | 0.00015 | m | 0.001 | 0.0030079 | 4.7434e-08 | 1.6914e-06 | 1.5e-07 | 500Hz3m
1000 | 1 | kHz | 3 | m | 0.003 | 3.0214 | 0.00013416 | 0.0033908 | 0.0003 | m | 0.001 | 0.0030214 | 1.3416e-07 | 3.3908e-06 | 3e-07 | 1kHz3m
2000 | 2 | kHz | 3 | m | 0.003 | 3.0627 | 0.00037947 | 0.0073349 | 0.0006 | m | 0.001 | 0.0030627 | 3.7947e-07 | 7.3349e-06 | 6e-07 | 2kHz3m
5000 | 5 | kHz | 3 | m | 0.003 | 3.2478 | 0.0015 | 0.019038 | 0.0015 | m | 0.001 | 0.0032478 | 1.5e-06 | 1.9038e-05 | 1.5e-06 | 5kHz3m
0.01 | 10 | mHz | 10 | m | 0.01 | 10.001 | 0.0001 | 0.00040086 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 4.0086e-07 | 5e-07 | 10mHz10m
0.02 | 20 | mHz | 10 | m | 0.01 | 9.9999 | 0.0001 | 0.00064265 | 0.0005 | m | 0.001 | 0.0099999 | 1e-07 | 6.4265e-07 | 5e-07 | 20mHz10m
0.05 | 50 | mHz | 10 | m | 0.01 | 10.001 | 0.0001 | -0.00032046 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | -3.2046e-07 | 5e-07 | 50mHz10m
0.1 | 100 | mHz | 10 | m | 0.01 | 9.9989 | 0.0001 | 0.00028477 | 0.0005 | m | 0.001 | 0.0099989 | 1e-07 | 2.8477e-07 | 5e-07 | 100mHz10m
0.2 | 200 | mHz | 10 | m | 0.01 | 10.0013 | 0.0001 | -0.0026436 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | -2.6436e-06 | 5e-07 | 200mHz10m
0.5 | 500 | mHz | 10 | m | 0.01 | 10.0013 | 0.0001 | 0.00025137 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 2.5137e-07 | 5e-07 | 500mHz10m
1 | 1 | Hz | 10 | m | 0.01 | 10.0007 | 0.0001 | 0.00087214 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 8.7214e-07 | 5e-07 | 1Hz10m
2 | 2 | Hz | 10 | m | 0.01 | 9.9992 | 0.0001 | 0.0029561 | 0.0005 | m | 0.001 | 0.0099992 | 1e-07 | 2.9561e-06 | 5e-07 | 2Hz10m
5 | 5 | Hz | 10 | m | 0.01 | 9.9994 | 0.0001 | 0.00016976 | 0.0005 | m | 0.001 | 0.0099994 | 1e-07 | 1.6976e-07 | 5e-07 | 5Hz10m
10 | 10 | Hz | 10 | m | 0.01 | 9.9996 | 0.0001 | -0.00033252 | 0.0005 | m | 0.001 | 0.0099996 | 1e-07 | -3.3252e-07 | 5e-07 | 10Hz10m
20 | 20 | Hz | 10 | m | 0.01 | 10.0006 | 0.0001 | -0.00013391 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | -1.3391e-07 | 5e-07 | 20Hz10m
50 | 50 | Hz | 10 | m | 0.01 | 10.0013 | 0.0001 | 0.0011725 | 0.0005 | m | 0.001 | 0.010001 | 1e-07 | 1.1725e-06 | 5e-07 | 50Hz10m
100 | 100 | Hz | 10 | m | 0.01 | 10.0031 | 0.0001 | 0.00025917 | 0.0005 | m | 0.001 | 0.010003 | 1e-07 | 2.5917e-07 | 5e-07 | 100Hz10m
200 | 200 | Hz | 10 | m | 0.01 | 10.0067 | 0.0001 | 0.0025996 | 0.0005 | m | 0.001 | 0.010007 | 1e-07 | 2.5996e-06 | 5e-07 | 200Hz10m
500 | 500 | Hz | 10 | m | 0.01 | 10.0242 | 0.00015811 | 0.0064006 | 0.0005 | m | 0.001 | 0.010024 | 1.5811e-07 | 6.4006e-06 | 5e-07 | 500Hz10m
1000 | 1 | kHz | 10 | m | 0.01 | 10.0745 | 0.00044721 | 0.012677 | 0.001 | m | 0.001 | 0.010074 | 4.4721e-07 | 1.2677e-05 | 1e-06 | 1kHz10m
2000 | 2 | kHz | 10 | m | 0.01 | 10.2078 | 0.0012649 | 0.026377 | 0.002 | m | 0.001 | 0.010208 | 1.2649e-06 | 2.6377e-05 | 2e-06 | 2kHz10m
5000 | 5 | kHz | 10 | m | 0.01 | 10.8269 | 0.005 | 0.062068 | 0.005 | m | 0.001 | 0.010827 | 5e-06 | 6.2068e-05 | 5e-06 | 5kHz10m
0.01 | 10 | mHz | 100 | m | 0.1 | 100.0103 | 0.001 | 0.021959 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | 2.1959e-05 | 5e-06 | 10mHz100m
0.02 | 20 | mHz | 100 | m | 0.1 | 99.9913 | 0.001 | -0.0030977 | 0.005 | m | 0.001 | 0.099991 | 1e-06 | -3.0977e-06 | 5e-06 | 20mHz100m
0.05 | 50 | mHz | 100 | m | 0.1 | 99.9994 | 0.001 | -0.01193 | 0.005 | m | 0.001 | 0.099999 | 1e-06 | -1.193e-05 | 5e-06 | 50mHz100m
0.1 | 100 | mHz | 100 | m | 0.1 | 100.0196 | 0.001 | 0.0016399 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 1.6399e-06 | 5e-06 | 100mHz100m
0.2 | 200 | mHz | 100 | m | 0.1 | 99.9926 | 0.001 | -0.0010265 | 0.005 | m | 0.001 | 0.099993 | 1e-06 | -1.0265e-06 | 5e-06 | 200mHz100m
0.5 | 500 | mHz | 100 | m | 0.1 | 100.014 | 0.001 | -0.013826 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -1.3826e-05 | 5e-06 | 500mHz100m
1 | 1 | Hz | 100 | m | 0.1 | 100.0115 | 0.001 | -0.0047865 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -4.7865e-06 | 5e-06 | 1Hz100m
2 | 2 | Hz | 100 | m | 0.1 | 100.0105 | 0.001 | -0.00086121 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -8.6121e-07 | 5e-06 | 2Hz100m
5 | 5 | Hz | 100 | m | 0.1 | 100.0065 | 0.001 | -0.011974 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -1.1974e-05 | 5e-06 | 5Hz100m
10 | 10 | Hz | 100 | m | 0.1 | 100.0002 | 0.001 | -0.032911 | 0.005 | m | 0.001 | 0.1 | 1e-06 | -3.2911e-05 | 5e-06 | 10Hz100m
20 | 20 | Hz | 100 | m | 0.1 | 100.0144 | 0.001 | -0.011151 | 0.005 | m | 0.001 | 0.10001 | 1e-06 | -1.1151e-05 | 5e-06 | 20Hz100m
50 | 50 | Hz | 100 | m | 0.1 | 99.9862 | 0.001 | -0.0034939 | 0.005 | m | 0.001 | 0.099986 | 1e-06 | -3.4939e-06 | 5e-06 | 50Hz100m
100 | 100 | Hz | 100 | m | 0.1 | 100.0198 | 0.001 | 0.007221 | 0.005 | m | 0.001 | 0.10002 | 1e-06 | 7.221e-06 | 5e-06 | 100Hz100m
200 | 200 | Hz | 100 | m | 0.1 | 100.0743 | 0.001 | 0.014824 | 0.005 | m | 0.001 | 0.10007 | 1e-06 | 1.4824e-05 | 5e-06 | 200Hz100m
500 | 500 | Hz | 100 | m | 0.1 | 100.2702 | 0.0015811 | 0.059739 | 0.005 | m | 0.001 | 0.10027 | 1.5811e-06 | 5.9739e-05 | 5e-06 | 500Hz100m
1000 | 1 | kHz | 100 | m | 0.1 | 100.7545 | 0.0044721 | 0.12542 | 0.01 | m | 0.001 | 0.10075 | 4.4721e-06 | 0.00012542 | 1e-05 | 1kHz100m
2000 | 2 | kHz | 100 | m | 0.1 | 102.0873 | 0.012649 | 0.24727 | 0.02 | m | 0.001 | 0.10209 | 1.2649e-05 | 0.00024727 | 2e-05 | 2kHz100m
5000 | 5 | kHz | 100 | m | 0.1 | 108.2697 | 0.05 | 0.61379 | 0.05 | m | 0.001 | 0.10827 | 5e-05 | 0.00061379 | 5e-05 | 5kHz100m