## Info
RLC corrections for LiBforSecUse project
V1.0, 2021-05-27, Czech Metrology Institute, Stanislav Maslan
EMPIR project url: | https://www.ptb.de/empir2018/de/libforsecuse/project/overview/
This project Git: | https://github.com/smaslan/RLC-bridge-corrections
Description:
This XLS sheet was made as an example how to correct reading errors of RLC bridge at low impedances. The calibration concepts is to measure RLC bridge error using idealy two standards per range (top and bottom of each range) in to the sheet "Cal Data". The sheet "Cal Data" then takes calibration data of the used impedances from sheet "Ref Z" that has to be filled in. The "Cal Data" sheet then calculates gain-phase errors for each row (combination of range, frequency and nominal value of calibration standard). The gain-phase corrections from the "Cal Data" sheet are then loaded by sheet "Measurement" and used to correct the raw RLC meter readings.
Validation:
The sheet functionality was validated by Octave scripts that are part of this Git project. The Octave scripts generate proceduraly some correction data and measurement data with simulated errors of the RLC bridge and user can than verify this XLS sheet corrected the values as expected. Note the validation does not cover uncertainty propagation yet!

## Measurement
This sheet uses calibration data to fix bridge errors for main measurements.
1) Select measurement frequency and enable manual ranging of the instrument and select range appropriate for measured impedance following manual. If the manual ranging is not possible, it will be necessary to estimate which range RLS/EIS meter is using. Fill in selected range and frequency to the yellow columns.
2) Measure Rs-Xs of the standard preferably using averaging and write results to columns Rs, Xs and type A uncertainties to ua(Rs) and ua(Xs). You can change the suffix "m" which indicates milliOhms to nothing for Ohms.
3) Connect SHORT standard and measure residual Rs-Xs at the SAME RANGE and frequency as in step 2). Fill in measured Rs-Xs and associated type A uncertainties to columns "sh Rs", "sh Xs", "sh ua(Rs)" and "sh ua(Xs)".
4) Read the corrected result from blue columns "Fixed …". Note the expanded uncertainties "U(…)" are being calculated from reference standards uncertainties, type A uncertainties of calibration sheet and type A uncertainties of this sheet. The uncertainty does not contain effect of RLC/EIS meter non-linearity! Always check the "Status" column value. If this sheet found calibration data for the measurement spot, it will show "OK". If the measured value is outside range of calibration sheet data, it will show "Extrapolated". This happens e.g. if range is 3 mOhm, calibration standards are 1 mOhm and 3 mOhm and the measure value is outside interval of <1 mOhm ; 3 mOhm>, e.g. 3.07 mOhm or 0.997 mOhm. The correction will still work reasonaby well if the distance from calibration spots is not too big. If the calibration data are not found at all, it will show red error message. This possibly indicates wrong selection of range or frequency that was not measured in calibration sheet.
Notes: Use Rs-Xs suffix "m" for milliOhms or nothing for Ohms. The sheet will recalculate. The same applies for nominal Z and range. Add or remove lines to/from bottom of the table, but always copy formulas in whole width of the sheet.
|  |  | Measured Z |  |  |  |  | Measured residual SHORT Z: |  |  |  |  | Corrected Z (results) | [5 empty] | Validation: ref data |  |  |  |  | Deviation of XLS calculated from Octave calculated: |  |  |  |  | Deviation with no correction: | [6 empty] | Measured Z (basic units) | [8 empty] | Correction data search for interpolation | [25 empty] | Interpolated correction |  |  |  | corrected Z | [13 empty] | Interpolated correction (gain-phase) |  |  |  | Corrected Z (gain-phase method)
Range |  | f | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Fixed Rs | Fixed U(Rs) | Fixed Xs | Fixed U(Xs) |  | Status | Valid Rs |  | Valid Xs |  |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | Dev Rs | unc(Dev Rs) | Dev Xs | unc(Dev Xs) |  | f |  | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | ua(¦Z¦) | phi | ua(phi) | lower Z ID' | lower ID' | lower Z' | lower Z ID' | lower Z ID' | lower ID'' | lower Z ID'' | lower Z ID'' | lower Z tag | higher Z tag | lower Z | higher Z | lower tag | higher tag | lower ID | higher ID | low check | low cor(Rs) | low U(cor(Rs)) | low cor(Xs) | low U(cor(Xs)) | high check | high cor(Rs) | high U(cor(Rs)) | high cor(Xs) | high U(cor(Xs)) | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | Rs | U(Rs) | Xs | U(Xs) | low cor(gain) | low U(cor(gain)) | low cor(phi) | low U(cor(phi)) | low status | hi cor(gain) | hi U(cor(gain)) | hi cor(phi) | hi U(cor(phi)) | low status | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi)) | ¦Z¦ | U(¦Z¦) | phi | U(phi) | Rs | U(Rs) | Xs | U(Xs) | #octave_header#
Ohm |  | Hz | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm |  | - | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | Ohm | Ohm | Ohm | Ohm |  | - |  | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | Ohm |  | - |  |  |  |  | - | Ohm | Ohm | - | - | - | - | - | Ohm | Ohm | Ohm | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | - | rad | rad | - | - | - | rad | rad | - | - | - | rad | rad | Ohm | Ohm | rad | rad | Ohm | Ohm | Ohm | Ohm
100 | m | 100 | -37.3981 | -5.6656e-08 | 82.7057 | 0.0012628 | m | -0.0011044 | 0.0015225 | 0.00064602 | 0.0013223 | m | -37.4 | 0.0086741 | 82.7 | 0.0050565 | m | OK | -37.4 |  | 82.7 |  | m | 2.0243e-06 | 0.0086741 | -1.8385e-06 | 0.0050565 | m | 0.0030235 |  | 0.0050959 |  | m | 100 | Hz | 0.001 | -0.037397 | 1.5225e-06 | 0.082705 | 1.8284e-06 | 0.090767 | 1.7802e-06 | 1.9955 | 1.7391e-05 | 4 | 10m | 0.01 | 121 | 4 | 100m | 139 | 4 | 10m | 100m | 0.01 | 0.1 | 100Hz100m10m | 100Hz100m100m | 121 | 139 | 100Hz100m10m | 7.5569e-07 | 0.0023218 | 1.0009e-06 | 0.0032991 | 100Hz100m100m | -5.0441e-06 | 0.0027683 | 4.8126e-06 | 0.0038891 | -4.4491e-06 | 0.0027683 | 4.4215e-06 | 0.0038286 | -0.037401 | 0.0027683 | 0.08271 | 0.0038286 | 1.0001 | 4.0981e-06 | 9.9998e-05 | 0.00036047 | OK | 0.99995 | 3.8403e-06 | 4.8155e-05 | 6.7252e-05 | OK | 0.99996 | 3.8403e-06 | 5.3473e-05 | 9.7333e-05 | 0.090764 | 3.5774e-06 | 1.9955 | 0.00010336 | -0.0374 | 8.6741e-06 | 0.0827 | 5.0565e-06
3 | m | 0.01 | -2.7797 | 0.0012698 | 0.77591 | 2.5832e-05 | m | 0.00023041 | 0.0015336 | 0.00070774 | 0.0001581 | m | -2.78 | 0.0037549 | 0.775 | 0.0025939 | m | OK | -2.78 |  | 0.775 |  | m | -2.802e-08 | 0.0037549 | -2.7849e-09 | 0.0025939 | m | 0.00010185 |  | 0.00020722 |  | m | 10 | mHz | 0.001 | -0.0027799 | 1.9911e-06 | 0.00077521 | 1.602e-07 | 0.002886 | 1.9184e-06 | 2.8696 | 0.00019288 | 2 | 1m | 0.001 | 37 | 2 | 3m | 55 | 2 | 1m | 3m | 0.001 | 0.003 | 10mHz3m1m | 10mHz3m3m | 37 | 55 | 10mHz3m1m | 6.8576e-08 | 0.0002082 | 1e-07 | 0.00039678 | 10mHz3m3m | 3.4393e-08 | 0.0019573 | 2.3204e-07 | 0.001534 | 3.6342e-08 | 0.0019573 | 2.2451e-07 | 0.0014692 | -0.0027799 | 0.0019573 | 0.00077543 | 0.0014692 | 1.0001 | 1.1188e-06 | 9.9999e-05 | 0.0036046 | OK | 1 | 2.2411e-06 | 7.7357e-05 | 0.00059275 | OK | 1 | 2.2411e-06 | 7.8648e-05 | 0.00076448 | 0.002886 | 3.8367e-06 | 2.8697 | 0.00085629 | -0.00278 | 3.7549e-06 | 0.000775 | 2.5939e-06
3 | m | 0.05 | -1.552 | 0.0016715 | -2.5317 | 0.0017757 | m | -0.0016794 | 0.00183 | -0.0015999 | 0.0001872 | m | -1.55 | 0.0051922 | -2.53 | 0.0044833 | m | OK | -1.55 |  | -2.53 |  | m | 1.1935e-08 | 0.0051922 | 3.0726e-08 | 0.0044833 | m | -0.00032389 |  | -0.00013215 |  | m | 50 | mHz | 0.001 | -0.0015503 | 2.4785e-06 | -0.0025301 | 1.7855e-06 | 0.0029673 | 1.9987e-06 | -2.1205 | 0.0007785 | 2 | 1m | 0.001 | 39 | 2 | 3m | 57 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 7.9425e-08 | 0.0031084 | 9.998e-08 | 0.0014049 | 50mHz3m3m | -2.9377e-07 | 0.0027194 | 2.0792e-07 | 0.0030941 | -2.8768e-07 | 0.0027194 | 2.0615e-07 | 0.0030665 | -0.0015506 | 0.0027194 | -0.0025299 | 0.0030665 | 1.0001 | 3.3965e-06 | 0.0001 | 0.0017695 | OK | 0.9999 | 2.8003e-06 | 6.9306e-05 | 0.0010399 | OK | 0.9999 | 2.8003e-06 | 6.9807e-05 | 0.0010518 | 0.0029671 | 3.9973e-06 | -2.1205 | 0.001879 | -0.00155 | 5.1922e-06 | -0.00253 | 4.4833e-06
10 | m | 5000 | 3.1581 | 0.0014167 | 0.83964 | 0.001395 | m | -0.0019786 | 0.00016512 | 0.0019495 | 0.0019314 | m | 3.16 | 0.0032291 | 0.83799 | 0.0052965 | m | OK | 3.16 |  | 0.838 |  | m | 1.5845e-05 | 0.0032291 | -6.8502e-06 | 0.0052965 | m | 7.3631e-05 |  | -0.000309 |  | m | 5 | kHz | 0.001 | 0.0031601 | 1.4263e-06 | 0.00083769 | 2.3825e-06 | 0.0032692 | 1.5078e-06 | 0.25913 | 0.00071326 | 3 | 3m | 0.003 | 90 | 3 | 10m | 108 | 3 | 3m | 10m | 0.003 | 0.01 | 5kHz10m3m | 5kHz10m10m | 90 | 108 | 5kHz10m3m | -4.7338e-09 | 0.0034816 | 3.1839e-07 | 0.0020409 | 5kHz10m10m | 1.1528e-06 | 0.0034956 | 2.1712e-07 | 0.0002864 | 3.9784e-08 | 0.0034956 | 3.1449e-07 | 0.0019734 | 0.0031601 | 0.0034956 | 0.00083801 | 0.0019734 | 1 | 3.581e-06 | 9.7001e-05 | 0.00085617 | OK | 1.0001 | 4.4657e-06 | 1.544e-05 | 0.00057344 | OK | 1 | 4.4657e-06 | 9.3864e-05 | 0.00084529 | 0.0032692 | 3.0156e-06 | 0.25922 | 0.0016582 | 0.00316 | 3.2291e-06 | 0.00083799 | 5.2965e-06
3 | m | 2000 | 0.86186 | 0.0008256 | -0.30724 | 0.0012871 | m | 0.001959 | 0.00018473 | 0.0008243 | 9.3202e-06 | m | 0.86 | 0.0020997 | -0.308 | 0.0034646 | m | Extrapolated | 0.86 |  | -0.308 |  | m | -5.4384e-07 | 0.0020997 | -8.9285e-07 | 0.0034646 | m | -0.00010396 |  | -6.4134e-05 |  | m | 2 | kHz | 0.001 | 0.0008599 | 8.4602e-07 | -0.00030806 | 1.2872e-06 | 0.00091341 | 9.0708e-07 | -0.34401 | 0.0013629 | 1 | 0m | 0 | 0 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 8.2417e-08 | 0.0022211 | 1.0261e-07 | 0.0022059 | 2kHz3m3m | 1.8611e-07 | 0.0015749 | -1.8404e-08 | 0.00086881 | 7.7927e-08 | 0.0015749 | 1.0784e-07 | 0.0022638 | 0.00085997 | 0.0015749 | -0.00030796 | 0.0022638 | 1.0001 | 2.3327e-06 | 9.8773e-05 | 0.0028087 | OK | 1.0001 | 3.026e-06 | -7.077e-06 | 0.0011652 | OK | 1.0001 | 3.026e-06 | 0.00010336 | 0.0028799 | 0.00091349 | 1.8142e-06 | -0.34391 | 0.0039653 | 0.00086 | 2.0997e-06 | -0.000308 | 3.4646e-06
3 | m | 10 | 1.3588 | 0.001805 | -0.53037 | 0.0012107 | m | -0.0010935 | 0.0013643 | 0.0017414 | 0.00072937 | m | 1.36 | 0.004447 | -0.532 | 0.0050068 | m | OK | 1.36 |  | -0.532 |  | m | -1.8023e-09 | 0.004447 | -7.9985e-10 | 0.0050068 | m | -9.9504e-05 |  | -0.00011376 |  | m | 10 | Hz | 0.001 | 0.0013599 | 2.2626e-06 | -0.00053211 | 1.4134e-06 | 0.0014603 | 2.1691e-06 | -0.37297 | 0.0010636 | 2 | 1m | 0.001 | 46 | 2 | 3m | 64 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 3.7737e-08 | 0.0024236 | 1.0002e-07 | 0.0016399 | 10Hz3m3m | 7.8446e-08 | 0.0025183 | 2.6572e-07 | 0.0036641 | 4.7106e-08 | 0.0025183 | 1.3816e-07 | 0.0021058 | 0.0013599 | 0.0025184 | -0.00053198 | 0.0021058 | 1 | 3.7848e-06 | 9.9999e-05 | 0.0032292 | OK | 1 | 3.3627e-06 | 8.8572e-05 | 0.0012379 | OK | 1 | 3.3627e-06 | 9.7369e-05 | 0.0027709 | 0.0014604 | 4.3381e-06 | -0.37288 | 0.0034932 | 0.00136 | 4.447e-06 | -0.000532 | 5.0068e-06
3 | m | 0.02 | 2.1895 | 0.0011623 | 0.95892 | 0.0019378 | m | -0.00039507 | 0.00061804 | -0.0011253 | 0.00091771 | m | 2.19 | 0.0033662 | 0.96 | 0.0046972 | m | OK | 2.19 |  | 0.96 |  | m | -1.9555e-08 | 0.0033662 | 2.5705e-08 | 0.0046972 | m | -8.6554e-05 |  | 4.2771e-05 |  | m | 20 | mHz | 0.001 | 0.0021899 | 1.3164e-06 | 0.00096004 | 2.1441e-06 | 0.0023911 | 1.4814e-06 | 0.41316 | 0.00085048 | 2 | 1m | 0.001 | 38 | 2 | 3m | 56 | 2 | 1m | 3m | 0.001 | 0.003 | 20mHz3m1m | 20mHz3m3m | 38 | 56 | 20mHz3m1m | -1.1939e-08 | 0.0028051 | 1e-07 | 0.00048298 | 20mHz3m3m | 1.277e-07 | 0.0037817 | -2.6457e-07 | 0.0016759 | 8.5187e-08 | 0.0037817 | -1.5358e-07 | 0.0013127 | 0.00219 | 0.0037817 | 0.00095989 | 0.0013127 | 0.99999 | 3.0066e-06 | 0.0001 | 0.0022748 | OK | 1 | 4.443e-06 | -8.8199e-05 | 0.00071272 | OK | 1 | 4.443e-06 | -3.0902e-05 | 0.0011883 | 0.0023912 | 2.9629e-06 | 0.41313 | 0.0020749 | 0.00219 | 3.3662e-06 | 0.00096 | 4.6972e-06
10 | m | 500 | 4.8408 | 0.0013678 | -3.2393 | 0.00027386 | m | 0.0014654 | 0.0013657 | 0.00094909 | 0.00033527 | m | 4.84 | 0.0030477 | -3.24 | 0.002781 | m | OK | 4.84 |  | -3.24 |  | m | -4.5091e-08 | 0.0030477 | -4.6675e-08 | 0.002781 | m | -0.00065064 |  | -0.00025323 |  | m | 500 | Hz | 0.001 | 0.0048393 | 1.9328e-06 | -0.0032403 | 4.3291e-07 | 0.005824 | 1.624e-06 | -0.59001 | 0.0001947 | 3 | 3m | 0.003 | 87 | 3 | 10m | 105 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 2.0609e-07 | 0.0020148 | 3.0181e-07 | 0.00012312 | 500Hz10m10m | 6.7942e-07 | 0.0037438 | 9.6332e-07 | 0.0015915 | 3.9704e-07 | 0.0037438 | 5.6868e-07 | 0.00071549 | 0.0048397 | 0.0037438 | -0.0032397 | 0.00071549 | 1.0001 | 4.173e-06 | 9.9995e-05 | 0.00012568 | OK | 1.0001 | 3.7612e-06 | 9.5716e-05 | 0.00030464 | OK | 1.0001 | 3.7612e-06 | 9.8268e-05 | 0.00019788 | 0.0058244 | 3.2481e-06 | -0.58991 | 0.00043679 | 0.00484 | 3.0477e-06 | -0.00324 | 2.781e-06
3 | m | 0.2 | 0.14172 | 0.0017401 | -2.3407 | 0.0014531 | m | -0.0011968 | 8.4627e-05 | -0.00068716 | 1.2781e-05 | m | 0.143 | 0.0060996 | -2.34 | 0.0029272 | m | OK | 0.143 |  | -2.34 |  | m | 3.2365e-09 | 0.0060996 | -1.0021e-08 | 0.0029272 | m | -8.5462e-05 |  | 3.082e-05 |  | m | 200 | mHz | 0.001 | 0.00014291 | 1.7421e-06 | -0.00234 | 1.4532e-06 | 0.0023443 | 1.4544e-06 | -1.5098 | 0.00074271 | 2 | 1m | 0.001 | 41 | 2 | 3m | 59 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | 9.3668e-08 | 0.0035594 | 9.9998e-08 | 0.0029736 | 200mHz3m3m | -6.8559e-08 | 0.0011882 | 1.2511e-08 | 0.0026845 | -1.5375e-08 | 0.0011882 | 4.1192e-08 | 0.0027793 | 0.0001429 | 0.0011882 | -0.0023399 | 0.0027793 | 1.0001 | 3.7266e-06 | 9.9999e-05 | 0.0046473 | OK | 0.99998 | 1.8904e-06 | 4.1703e-06 | 0.00091812 | OK | 1 | 1.8904e-06 | 3.5586e-05 | 0.0021407 | 0.0023444 | 2.9088e-06 | -1.5098 | 0.0026056 | 0.000143 | 6.0996e-06 | -0.00234 | 2.9272e-06
10 | m | 0.1 | -6.1289 | 0.00024196 | -2.979 | 0.0001048 | m | 0.0012139 | 0.0019188 | 0.00053143 | 0.00062502 | m | -6.13 | 0.0037854 | -2.98 | 0.0045309 | m | OK | -6.13 |  | -2.98 |  | m | -9.4224e-09 | 0.0037854 | -1.3471e-08 | 0.0045309 | m | -0.00016195 |  | 0.00043596 |  | m | 100 | mHz | 0.001 | -0.0061302 | 1.934e-06 | -0.0029796 | 6.3375e-07 | 0.0068159 | 1.7614e-06 | -2.6892 | 0.0001496 | 3 | 3m | 0.003 | 76 | 3 | 10m | 94 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | 1.1845e-07 | 0.00015425 | 2.9999e-07 | 0.0034865 | 100mHz10m10m | -2.0872e-07 | 0.00018693 | 4.114e-07 | 0.0013083 | -5.9903e-08 | 0.00018693 | 3.6072e-07 | 0.0022991 | -0.0061302 | 0.00018697 | -0.0029792 | 0.0022991 | 1 | 1.6336e-06 | 0.0001 | 0.0011898 | OK | 0.99998 | 3.2915e-06 | 4.1139e-05 | 0.00015822 | OK | 1 | 3.2915e-06 | 6.7915e-05 | 0.00062744 | 0.006816 | 3.5228e-06 | -2.6891 | 0.00069512 | -0.00613 | 3.7854e-06 | -0.00298 | 4.5309e-06
1 | m | 200 | -0.23176 | 0.00018324 | 0.46035 | 0.00028819 | m | 0.00024315 | 0.0017811 | -0.0016462 | 0.0009589 | m | -0.23199 | 0.0031812 | 0.462 | 0.0026226 | m | OK | -0.232 |  | 0.462 |  | m | 5.6275e-06 | 0.0031812 | 2.8226e-06 | 0.0026226 | m | -7.8903e-06 |  | 7.1853e-07 |  | m | 200 | Hz | 0.001 | -0.00023201 | 1.7905e-06 | 0.000462 | 1.0013e-06 | 0.00051698 | 1.2026e-06 | 2.0362 | 0.0032148 | 1 | 0m | 0 | 14 | 1 | 1m | 32 | 1 | 0m | 1m | 0 | 0.001 | 200Hz1m0m | 200Hz1m1m | 14 | 32 | 200Hz1m0m | 0 | 3.2533e-05 | 0 | 0.00022199 | 200Hz1m1m | -1.5628e-08 | 0.00063582 | -2.5239e-08 | 0.00069892 | -8.0794e-09 | 0.00063582 | -2.5239e-08 | 0.00046856 | -0.00023202 | 0.00063583 | 0.00046198 | 0.00046856 | 1 | 2.2436e-07 | -2.5193e-05 | 0.00078379 | OK | 0.99998 | 3.4936e-06 | -2.5193e-05 | 0.00078379 | OK | 0.99999 | 3.4936e-06 | -2.5193e-05 | 0.00078379 | 0.00051698 | 2.4052e-06 | 2.0362 | 0.0064772 | -0.00023199 | 3.1812e-06 | 0.000462 | 2.6226e-06
100 | m | 10 | -76.9971 | -0.00065378 | 50.1904 | 0.001878 | m | 0.0016876 | 7.7204e-06 | -0.00032148 | 0.0012447 | m | -77 | 0.0046975 | 50.2 | 0.0064737 | m | OK | -77 |  | 50.2 |  | m | -7.9935e-07 | 0.0046975 | -9.0643e-07 | 0.0064737 | m | 0.00122 |  | -0.0092857 |  | m | 10 | Hz | 0.001 | -0.076999 | 6.5383e-07 | 0.050191 | 2.2531e-06 | 0.091913 | 1.3468e-06 | 2.5639 | 2.09e-05 | 4 | 10m | 0.01 | 118 | 4 | 100m | 136 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | 9.1346e-07 | 0.00024891 | 9.9995e-07 | 0.00018397 | 10Hz100m100m | 6.3813e-06 | 0.0038509 | -9.4869e-06 | 0.0017151 | 5.89e-06 | 0.0038509 | -8.5446e-06 | 0.0015775 | -0.076993 | 0.0038509 | 0.050182 | 0.0015775 | 1.0001 | 1.6713e-06 | 9.9999e-05 | 0.00024636 | OK | 1.0001 | 4.1119e-06 | -9.488e-05 | 5.2973e-05 | OK | 1.0001 | 4.1119e-06 | -7.7368e-05 | 7.035e-05 | 0.091919 | 2.7199e-06 | 2.5638 | 8.1831e-05 | -0.077 | 4.6975e-06 | 0.0502 | 6.4737e-06
3 | m | 10 | -0.87029 | 0.00093137 | -0.39888 | 0.0015687 | m | -0.00028659 | 0.00057557 | -0.000984 | 0.00084069 | m | -0.87 | 0.0029572 | -0.398 | 0.0043038 | m | Extrapolated | -0.87 |  | -0.398 |  | m | 5.2467e-10 | 0.0029572 | 1.3734e-09 | 0.0043038 | m | -6.8401e-06 |  | 0.00010233 |  | m | 10 | Hz | 0.001 | -0.00087001 | 1.0949e-06 | -0.0003979 | 1.7798e-06 | 0.00095668 | 1.2407e-06 | -2.7126 | 0.0017575 | 1 | 0m | 0 | 0 | 2 | 3m | 64 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 3.7737e-08 | 0.0024236 | 1.0002e-07 | 0.0016399 | 10Hz3m3m | 7.8446e-08 | 0.0025183 | 2.6572e-07 | 0.0036641 | 3.6855e-08 | 0.0025183 | 9.643e-08 | 0.001596 | -0.00086997 | 0.0025183 | -0.0003978 | 0.0015961 | 1 | 3.7848e-06 | 9.9999e-05 | 0.0032292 | OK | 1 | 3.3627e-06 | 8.8572e-05 | 0.0012379 | OK | 1 | 3.3627e-06 | 0.00010025 | 0.0032724 | 0.00095672 | 2.4814e-06 | -2.7125 | 0.0048025 | -0.00087 | 2.9572e-06 | -0.000398 | 4.3038e-06
3 | m | 200 | -0.83258 | 0.0014782 | -0.78929 | 0.00035808 | m | -0.0014709 | 0.00022643 | 0.00065893 | 0.00067715 | m | -0.831 | 0.0026045 | -0.79 | 0.0026242 | m | OK | -0.831 |  | -0.79 |  | m | -5.2015e-09 | 0.0026045 | -2.5519e-09 | 0.0026242 | m | -0.0001141 |  | 4.9299e-05 |  | m | 200 | Hz | 0.001 | -0.00083111 | 1.4955e-06 | -0.00078995 | 7.66e-07 | 0.0011466 | 1.2056e-06 | -2.3816 | 0.0010207 | 2 | 1m | 0.001 | 50 | 2 | 3m | 68 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | -4.3583e-08 | 0.002558 | 9.9996e-08 | 0.00064257 | 200Hz3m3m | -9.4686e-08 | 0.00067678 | 2.887e-07 | 7.157e-05 | -4.7329e-08 | 0.00067678 | 1.1383e-07 | 0.0006007 | -0.00083116 | 0.00067678 | -0.00078984 | 0.0006007 | 0.99996 | 4.4907e-06 | 9.9999e-05 | 0.0014006 | OK | 0.99997 | 2.5292e-06 | 9.6222e-05 | 0.00063872 | OK | 0.99996 | 2.5292e-06 | 9.9722e-05 | 0.0013448 | 0.0011466 | 2.4112e-06 | -2.3815 | 0.0024445 | -0.000831 | 2.6045e-06 | -0.00079 | 2.6242e-06
3 | m | 0.1 | 2.0317 | 0.001008 | 1.2386 | 0.00048316 | m | 0.0016378 | 0.0011454 | -0.0013336 | 0.00023609 | m | 2.03 | 0.0033954 | 1.24 | 0.0043562 | m | OK | 2.03 |  | 1.24 |  | m | -3.919e-08 | 0.0033954 | 5.7445e-09 | 0.0043562 | m | 6.9584e-05 |  | -8.2308e-05 |  | m | 100 | mHz | 0.001 | 0.0020301 | 1.5258e-06 | 0.0012399 | 5.3776e-07 | 0.0023788 | 1.3319e-06 | 0.54831 | 0.000386 | 2 | 1m | 0.001 | 40 | 2 | 3m | 58 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | -8.9053e-08 | 0.00091244 | 1.0002e-07 | 0.00073157 | 100mHz3m3m | 9.0138e-08 | 0.0031681 | 5.9749e-08 | 0.00024521 | 3.4479e-08 | 0.0031681 | 7.2258e-08 | 0.00039628 | 0.0020301 | 0.0031681 | 0.00124 | 0.00039628 | 0.99991 | 1.8792e-06 | 9.9997e-05 | 0.0032159 | OK | 1 | 4.5862e-06 | 1.9912e-05 | 0.0012806 | OK | 0.99999 | 4.5862e-06 | 4.4787e-05 | 0.0018817 | 0.0023788 | 2.6639e-06 | 0.54835 | 0.0020339 | 0.00203 | 3.3954e-06 | 0.00124 | 4.3562e-06
3 | m | 5000 | 1.1892 | 0.0017758 | -1.5201 | 4.2037e-06 | m | -0.00080767 | 0.0014637 | -7.6513e-05 | 0.0019103 | m | 1.19 | 0.0045834 | -1.52 | 0.0044153 | m | OK | 1.19 |  | -1.52 |  | m | -3.9832e-05 | 0.0045834 | 4.2833e-06 | 0.0044153 | m | -3.358e-05 |  | -3.0949e-05 |  | m | 5 | kHz | 0.001 | 0.00119 | 2.3012e-06 | -0.00152 | 1.9103e-06 | 0.0019304 | 2.0676e-06 | -0.9066 | 0.0011195 | 2 | 1m | 0.001 | 54 | 2 | 3m | 72 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | 6.676e-08 | 0.0014337 | 1.0383e-07 | 0.00037658 | 5kHz3m3m | -3.4814e-07 | 0.00028055 | -3.0166e-07 | 0.00090262 | -1.2625e-07 | 0.00028055 | -8.4807e-08 | 0.0006213 | 0.0011898 | 0.00028059 | -0.0015201 | 0.00062131 | 1.0001 | 3.2539e-06 | 9.229e-05 | 0.0011818 | OK | 0.99989 | 1.0015e-06 | -8.7391e-05 | 0.0010617 | OK | 0.99998 | 1.0015e-06 | 8.7005e-06 | 0.0011259 | 0.0019304 | 4.1351e-06 | -0.90659 | 0.0025061 | 0.00119 | 4.5834e-06 | -0.00152 | 4.4153e-06
100 | m | 1 | -12.0012 | 0.00027169 | -3.7108 | 0.0018257 | m | -0.00037035 | 0.0014336 | -0.0018657 | 0.00033145 | m | -12 | 0.0031604 | -3.71 | 0.0044241 | m | OK | -12 |  | -3.71 |  | m | -3.6623e-09 | 0.0031604 | -2.3115e-10 | 0.0044241 | m | -0.00080859 |  | 0.0010604 |  | m | 1 | Hz | 0.001 | -0.012001 | 1.4591e-06 | -0.0037089 | 1.8556e-06 | 0.012561 | 1.4978e-06 | -2.8418 | 0.00014525 | 4 | 10m | 0.01 | 115 | 4 | 100m | 133 | 4 | 10m | 100m | 0.01 | 0.1 | 1Hz100m10m | 1Hz100m100m | 115 | 133 | 1Hz100m10m | -3.7732e-07 | 0.0039589 | 1e-06 | 0.0021272 | 1Hz100m100m | 3.0133e-07 | 0.0017307 | 8.828e-06 | 9.3545e-05 | -3.5801e-07 | 0.0017307 | 1.2228e-06 | 0.0020693 | -0.012001 | 0.0017307 | -0.0037077 | 0.0020693 | 0.99996 | 3.9633e-06 | 0.0001 | 0.00021924 | OK | 1 | 3.0339e-06 | 8.8269e-05 | 5.9552e-05 | OK | 0.99996 | 3.0339e-06 | 9.9666e-05 | 0.0002147 | 0.01256 | 2.9959e-06 | -2.8417 | 0.00036122 | -0.012 | 3.1604e-06 | -0.00371 | 4.4241e-06
10 | m | 20 | -4.0887 | 0.0012652 | 1.682 | 8.0151e-05 | m | 0.001087 | 0.00072792 | 0.0016535 | 0.0011396 | m | -4.09 | 0.0034721 | 1.68 | 0.0056402 | m | OK | -4.09 |  | 1.68 |  | m | 1.8937e-09 | 0.0034721 | 3.87e-09 | 0.0056402 | m | 0.00018092 |  | 0.00033123 |  | m | 20 | Hz | 0.001 | -0.0040898 | 1.4597e-06 | 0.0016803 | 1.1424e-06 | 0.0044216 | 1.4182e-06 | 2.7518 | 0.00026992 | 3 | 3m | 0.003 | 83 | 3 | 10m | 101 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | 7.3898e-08 | 0.0013404 | 3.0002e-07 | 0.0026209 | 20Hz10m10m | -5.06e-07 | 0.003568 | 2.5343e-07 | 0.0011237 | -4.3868e-08 | 0.003568 | 2.9056e-07 | 0.0023169 | -0.0040899 | 0.003568 | 0.0016806 | 0.0023169 | 1 | 2.651e-06 | 0.0001 | 0.0014598 | OK | 0.99995 | 3.7879e-06 | 2.5347e-05 | 0.00038391 | OK | 1 | 3.7879e-06 | 8.484e-05 | 0.0012413 | 0.0044216 | 2.8365e-06 | 2.7518 | 0.0013536 | -0.00409 | 3.4721e-06 | 0.00168 | 5.6402e-06
1 | m | 0.02 | 0.33408 | 0.00023249 | -0.61737 | 0.0017852 | m | 0.001108 | 0.00034037 | -0.00040018 | 0.00096048 | m | 0.333 | 0.0034231 | -0.617 | 0.0035417 | m | OK | 0.333 |  | -0.617 |  | m | 3.5406e-06 | 0.0034231 | 1.9102e-06 | 0.0035417 | m | -2.5136e-05 |  | 2.6706e-05 |  | m | 20 | mHz | 0.001 | 0.00033297 | 4.122e-07 | -0.00061697 | 2.0272e-06 | 0.00070109 | 1.7947e-06 | -1.0759 | 0.0014675 | 1 | 0m | 0 | 2 | 1 | 1m | 20 | 1 | 0m | 1m | 0 | 0.001 | 20mHz1m0m | 20mHz1m1m | 2 | 20 | 20mHz1m0m | 0 | 0.002644 | 0 | 0.0016373 | 20mHz1m1m | 7.21e-08 | 0.002925 | 1.9186e-08 | 0.0018052 | 5.0549e-08 | 0.002925 | 1.9186e-08 | 0.001755 | 0.00033303 | 0.002925 | -0.00061695 | 0.001755 | 1 | 3.1099e-06 | 1.9197e-05 | 0.003812 | OK | 1.0001 | 3.8738e-06 | 1.9197e-05 | 0.003812 | OK | 1.0001 | 3.8738e-06 | 1.9197e-05 | 0.003812 | 0.00070113 | 3.5894e-06 | -1.0759 | 0.004811 | 0.000333 | 3.4231e-06 | -0.000617 | 3.5417e-06
10 | m | 2000 | -9.6418 | 0.0015795 | -2.9074 | 0.00035264 | m | -0.0019606 | 0.0013472 | 0.0018061 | 0.00041966 | m | -9.64 | 0.0040685 | -2.91 | 0.0048014 | m | Extrapolated | -9.64 |  | -2.91 |  | m | -6.1551e-06 | 0.0040685 | 9.2229e-06 | 0.0048014 | m | 0.000147 |  | 0.00076604 |  | m | 2 | kHz | 0.001 | -0.0096399 | 2.076e-06 | -0.0029092 | 5.4815e-07 | 0.010069 | 1.9937e-06 | -2.8485 | 7.9148e-05 | 4 | 10m | 0.01 | 107 | 4 | 100m | 0 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | 7.5924e-08 | 0.0020259 | 3.0754e-07 | 0.0017468 | 2kHz10m10m | 3.5784e-07 | 0.0017525 | 7.0124e-07 | 0.0035239 | 3.6063e-07 | 0.0017525 | 7.0514e-07 | 0.0035415 | -0.0096395 | 0.0017525 | -0.0029085 | 0.0035415 | 1 | 2.6566e-06 | 9.9684e-05 | 0.0011987 | OK | 1 | 3.1544e-06 | 6.7873e-05 | 0.00046414 | OK | 1 | 3.1544e-06 | 6.7559e-05 | 0.00045687 | 0.01007 | 3.9876e-06 | -2.8484 | 0.00048352 | -0.00964 | 4.0685e-06 | -0.00291 | 4.8014e-06
3 | m | 10 | 0.96645 | 0.0011866 | 1.4287 | 5.0459e-05 | m | 0.0003428 | 0.00093235 | -0.0011674 | 0.0012838 | m | 0.966 | 0.0045695 | 1.43 | 0.0036804 | m | OK | 0.966 |  | 1.43 |  | m | 4.3395e-10 | 0.0045695 | -2.1251e-09 | 0.0036804 | m | 0.00010467 |  | -0.00014055 |  | m | 10 | Hz | 0.001 | 0.0009661 | 1.5091e-06 | 0.0014299 | 1.2848e-06 | 0.0017256 | 1.3591e-06 | 0.97659 | 0.00083594 | 2 | 1m | 0.001 | 46 | 2 | 3m | 64 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 3.7737e-08 | 0.0024236 | 1.0002e-07 | 0.0016399 | 10Hz3m3m | 7.8446e-08 | 0.0025183 | 2.6572e-07 | 0.0036641 | 5.2507e-08 | 0.0025183 | 1.6014e-07 | 0.0023743 | 0.00096616 | 0.0025183 | 0.00143 | 0.0023743 | 1 | 3.7848e-06 | 9.9999e-05 | 0.0032292 | OK | 1 | 3.3627e-06 | 8.8572e-05 | 0.0012379 | OK | 1 | 3.3627e-06 | 9.5853e-05 | 0.0025067 | 0.0017257 | 2.7181e-06 | 0.97669 | 0.0030131 | 0.000966 | 4.5695e-06 | 0.00143 | 3.6804e-06
10 | m | 50 | 0.72705 | 0.0013545 | -3.272 | 0.00075719 | m | 0.0013258 | 0.0002333 | -0.0017041 | 0.0019675 | m | 0.726 | 0.0060415 | -3.27 | 0.0042713 | m | OK | 0.726 |  | -3.27 |  | m | -1.5773e-09 | 0.0060415 | -1.2664e-08 | 0.0042713 | m | -0.00027229 |  | -0.00029919 |  | m | 50 | Hz | 0.001 | 0.00072573 | 1.3745e-06 | -0.0032703 | 2.1082e-06 | 0.0033499 | 2.0795e-06 | -1.3524 | 0.00042313 | 3 | 3m | 0.003 | 84 | 3 | 10m | 102 | 3 | 3m | 10m | 0.003 | 0.01 | 50Hz10m3m | 50Hz10m10m | 84 | 102 | 50Hz10m3m | -2.1913e-07 | 0.0010099 | 2.9991e-07 | 0.0034198 | 50Hz10m10m | -5.033e-08 | 0.00016567 | 7.4204e-07 | 0.00029584 | -2.1069e-07 | 0.00016567 | 3.2201e-07 | 0.0032636 | 0.00072552 | 0.00016569 | -0.00327 | 0.0032636 | 0.99993 | 1.1888e-06 | 9.9999e-05 | 0.001689 | OK | 0.99999 | 8.1371e-07 | 7.4197e-05 | 0.00028901 | OK | 0.99993 | 8.1371e-07 | 9.8709e-05 | 0.0016191 | 0.0033496 | 4.159e-06 | -1.3523 | 0.0018269 | 0.000726 | 6.0415e-06 | -0.00327 | 4.2713e-06
10 | m | 0.01 | 1.2307 | 0.0008937 | 9.8986 | 0.00092125 | m | 8.5719e-05 | 0.00082167 | -0.00037517 | 0.0016339 | m | 1.23 | 0.003095 | 9.9 | 0.0037258 | m | OK | 1.23 |  | 9.9 |  | m | 6.0949e-08 | 0.003095 | 2.0345e-07 | 0.0037258 | m | 0.00065609 |  | -0.00098565 |  | m | 10 | mHz | 0.001 | 0.0012307 | 1.214e-06 | 0.009899 | 1.8757e-06 | 0.0099752 | 1.8674e-06 | 1.4471 | 0.00012298 | 3 | 3m | 0.003 | 73 | 3 | 10m | 91 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | -1.2909e-07 | 0.0017511 | 2.9998e-07 | 0.00286 | 10mHz10m10m | 9.0451e-07 | 0.0029623 | 7.7359e-07 | 0.00028384 | 9.0085e-07 | 0.0029623 | 7.7192e-07 | 0.00029296 | 0.0012316 | 0.0029623 | 0.0098998 | 0.00029298 | 0.99996 | 4.2448e-06 | 0.0001 | 0.0010851 | OK | 1.0001 | 3.4123e-06 | 7.7369e-05 | 0.00018409 | OK | 1.0001 | 3.4123e-06 | 7.7449e-05 | 0.00018728 | 0.0099761 | 3.7349e-06 | 1.4472 | 0.00030914 | 0.00123 | 3.095e-06 | 0.0099 | 3.7258e-06
3 | m | 0.1 | -0.23925 | 0.00016982 | -1.0408 | 0.00030382 | m | -0.0011331 | 0.00047678 | -0.00078332 | 0.0015434 | m | -0.238 | 0.0035501 | -1.04 | 0.0031018 | m | OK | -0.238 |  | -1.04 |  | m | -3.3305e-09 | 0.0035501 | -3.5018e-09 | 0.0031018 | m | -0.00012147 |  | -6.531e-05 |  | m | 100 | mHz | 0.001 | -0.00023812 | 5.0612e-07 | -0.0010401 | 1.5731e-06 | 0.001067 | 1.5375e-06 | -1.7959 | 0.00056751 | 2 | 1m | 0.001 | 40 | 2 | 3m | 58 | 2 | 1m | 3m | 0.001 | 0.003 | 100mHz3m1m | 100mHz3m3m | 40 | 58 | 100mHz3m1m | -8.9053e-08 | 0.00091244 | 1.0002e-07 | 0.00073157 | 100mHz3m3m | 9.0138e-08 | 0.0031681 | 5.9749e-08 | 0.00024521 | -8.3052e-08 | 0.0031681 | 9.8673e-08 | 0.00071529 | -0.0002382 | 0.0031681 | -0.00104 | 0.00071529 | 0.99991 | 1.8792e-06 | 9.9997e-05 | 0.0032159 | OK | 1 | 4.5862e-06 | 1.9912e-05 | 0.0012806 | OK | 0.99991 | 4.5862e-06 | 9.7315e-05 | 0.0031511 | 0.0010669 | 3.0751e-06 | -1.7958 | 0.0033493 | -0.000238 | 3.5501e-06 | -0.00104 | 3.1018e-06
1 | m | 10 | 0.47269 | 0.00064365 | -0.44352 | 0.0016318 | m | -0.0013554 | 0.00071222 | -0.0015359 | 0.00094652 | m | 0.47398 | 0.0031974 | -0.44202 | 0.0032354 | m | OK | 0.474 |  | -0.442 |  | m | -1.553e-05 | 0.0031974 | -1.6648e-05 | 0.0032354 | m | 4.4933e-05 |  | 1.541e-05 |  | m | 10 | Hz | 0.001 | 0.00047404 | 9.5997e-07 | -0.00044198 | 1.8864e-06 | 0.00064813 | 1.4656e-06 | -0.75041 | 0.0023563 | 1 | 0m | 0 | 10 | 1 | 1m | 28 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0029676 | 0 | 0.0030462 | 10Hz1m1m | -5.3218e-08 | 0.0021012 | -9.9829e-08 | 0.0005903 | -3.4492e-08 | 0.0021012 | -9.9829e-08 | 0.0014545 | 0.00047401 | 0.0021012 | -0.00044208 | 0.0014545 | 1 | 4.2528e-06 | -9.9797e-05 | 0.0025687 | OK | 0.99995 | 4.19e-06 | -9.9797e-05 | 0.0025687 | OK | 0.99997 | 4.19e-06 | -9.9797e-05 | 0.0025687 | 0.0006481 | 2.9311e-06 | -0.75051 | 0.0053672 | 0.00047398 | 3.1974e-06 | -0.00044202 | 3.2354e-06
1 | m | 0.1 | -0.019451 | 0.0010262 | -0.0085767 | 0.00057205 | m | -0.0015514 | 0.00075464 | -0.0019767 | 0.00061163 | m | -0.0179 | 0.0023905 | -0.0066004 | 0.0018933 | m | OK | -0.0179 |  | -0.0066 |  | m | 1.4005e-07 | 0.0023905 | -3.7983e-07 | 0.0018933 | m | -3.2206e-09 |  | 7.2035e-09 |  | m | 100 | mHz | 0.001 | -1.79e-05 | 1.2738e-06 | -6.6e-06 | 8.3746e-07 | 1.9078e-05 | 1.2298e-06 | -2.7883 | 0.047221 | 1 | 0m | 0 | 4 | 1 | 1m | 22 | 1 | 0m | 1m | 0 | 0.001 | 100mHz1m0m | 100mHz1m1m | 4 | 22 | 100mHz1m0m | 0 | 0.00024824 | 0 | 0.0026528 | 100mHz1m1m | -1.4534e-09 | 0.001512 | 2.1634e-08 | 0.002673 | -2.7729e-11 | 0.001512 | 2.1634e-08 | 0.0026532 | -1.79e-05 | 0.001512 | -6.5784e-06 | 0.0026532 | 1 | 2.6644e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | 1.9078e-05 | 2.4596e-06 | -2.7883 | 0.094489 | -1.79e-05 | 2.3905e-06 | -6.6004e-06 | 1.8933e-06
100 | m | 0.2 | 14.8058 | 0.0011267 | 42.8966 | 0.00091712 | m | 0.0018544 | 6.2663e-05 | -0.0017876 | 0.001617 | m | 14.8 | 0.0084268 | 42.9 | 0.0044568 | m | OK | 14.8 |  | 42.9 |  | m | 4.6687e-09 | 0.0084268 | -6.736e-08 | 0.0044568 | m | 0.003932 |  | -0.0016497 |  | m | 200 | mHz | 0.001 | 0.014804 | 1.1285e-06 | 0.042898 | 1.859e-06 | 0.045381 | 1.7954e-06 | 1.2385 | 2.7039e-05 | 4 | 10m | 0.01 | 113 | 4 | 100m | 131 | 4 | 10m | 100m | 0.01 | 0.1 | 200mHz100m10m | 200mHz100m100m | 113 | 131 | 200mHz100m10m | 3.1379e-07 | 0.00044148 | 9.9993e-07 | 0.0023903 | 200mHz100m100m | -3.2931e-06 | 0.0042524 | 8.4146e-06 | 0.00042289 | -1.1042e-06 | 0.0042524 | 3.9148e-06 | 0.0016169 | 0.014803 | 0.0042524 | 0.042902 | 0.0016169 | 1 | 2.9627e-06 | 0.0001 | 0.00027525 | OK | 0.99997 | 5.327e-06 | 8.4134e-05 | 5.0378e-05 | OK | 1 | 5.327e-06 | 9.3763e-05 | 0.00018685 | 0.045381 | 3.599e-06 | 1.2386 | 0.00019451 | 0.0148 | 8.4268e-06 | 0.0429 | 4.4568e-06
10 | m | 10 | -1.7798 | 0.0014291 | 2.1317 | 0.0018698 | m | -4.8257e-05 | 0.0012525 | 0.0015084 | 0.00048592 | m | -1.78 | 0.0048189 | 2.13 | 0.0045454 | m | Extrapolated | -1.78 |  | 2.13 |  | m | 5.1371e-09 | 0.0048189 | 3.1412e-09 | 0.0045454 | m | 0.00022082 |  | 0.00019231 |  | m | 10 | Hz | 0.001 | -0.0017798 | 1.9003e-06 | 0.0021302 | 1.9319e-06 | 0.0027758 | 1.919e-06 | 2.2668 | 0.00068928 | 2 | 1m | 0.001 | 0 | 3 | 10m | 100 | 3 | 3m | 10m | 0.003 | 0.01 | 10Hz10m3m | 10Hz10m10m | 82 | 100 | 10Hz10m3m | -8.7454e-09 | 0.0014289 | 3e-07 | 0.0019201 | 10Hz10m10m | -2.676e-07 | 0.0015115 | -7.0775e-07 | 0.0015076 | -4.5662e-10 | 0.0015115 | 3.3227e-07 | 0.0019333 | -0.0017798 | 0.0015115 | 0.0021305 | 0.0019333 | 1 | 3.3748e-06 | 9.9997e-05 | 0.0013368 | OK | 0.99997 | 4.093e-06 | -7.0769e-05 | 0.00023114 | OK | 1 | 4.093e-06 | 0.00010547 | 0.0013722 | 0.0027758 | 3.8379e-06 | 2.2669 | 0.0019451 | -0.00178 | 4.8189e-06 | 0.00213 | 4.5454e-06
1 | m | 10 | -0.77618 | 0.00076027 | 0.12913 | 0.00081627 | m | 0.0018614 | 0.001847 | -0.00081219 | 0.00088055 | m | -0.778 | 0.0039409 | 0.13002 | 0.0032091 | m | OK | -0.778 |  | 0.13 |  | m | 2.7476e-06 | 0.0039409 | 1.6407e-05 | 0.0032091 | m | -4.289e-05 |  | -5.578e-05 |  | m | 10 | Hz | 0.001 | -0.00077804 | 1.9974e-06 | 0.00012994 | 1.2007e-06 | 0.00078882 | 1.98e-06 | 2.9761 | 0.0015582 | 1 | 0m | 0 | 10 | 1 | 1m | 28 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0029676 | 0 | 0.0030462 | 10Hz1m1m | -5.3218e-08 | 0.0021012 | -9.9829e-08 | 0.0005903 | -4.1979e-08 | 0.0021012 | -9.9829e-08 | 0.0011089 | -0.00077808 | 0.0021012 | 0.00012984 | 0.0011089 | 1 | 4.2528e-06 | -9.9797e-05 | 0.0025687 | OK | 0.99995 | 4.19e-06 | -9.9797e-05 | 0.0025687 | OK | 0.99996 | 4.19e-06 | -9.9797e-05 | 0.0025687 | 0.00078879 | 3.96e-06 | 2.976 | 0.0040386 | -0.000778 | 3.9409e-06 | 0.00013002 | 3.2091e-06
3 | m | 2000 | -1.1899 | 8.2823e-05 | 1.0103 | 0.0010368 | m | -5.0597e-05 | 0.0017024 | 0.00025945 | 0.00088367 | m | -1.19 | 0.0038989 | 1.01 | 0.0041562 | m | OK | -1.19 |  | 1.01 |  | m | 2.4016e-06 | 0.0038989 | 1.8791e-06 | 0.0041562 | m | 0.00016284 |  | 7.159e-06 |  | m | 2 | kHz | 0.001 | -0.0011898 | 1.7044e-06 | 0.00101 | 1.3623e-06 | 0.0015607 | 1.5702e-06 | 2.4378 | 0.0009707 | 2 | 1m | 0.001 | 53 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 8.2417e-08 | 0.0022211 | 1.0261e-07 | 0.0022059 | 2kHz3m3m | 1.8611e-07 | 0.0015749 | -1.8404e-08 | 0.00086881 | 1.1149e-07 | 0.0015749 | 6.8679e-08 | 0.0018311 | -0.0011897 | 0.0015749 | 0.0010101 | 0.0018311 | 1.0001 | 2.3327e-06 | 9.8773e-05 | 0.0028087 | OK | 1.0001 | 3.026e-06 | -7.077e-06 | 0.0011652 | OK | 1.0001 | 3.026e-06 | 6.9097e-05 | 0.002348 | 0.0015608 | 3.1404e-06 | 2.4378 | 0.0030466 | -0.00119 | 3.8989e-06 | 0.00101 | 4.1562e-06
1 | m | 1 | 0.19239 | 0.00022429 | 0.0058421 | 0.0011752 | m | 0.0013885 | 0.00034441 | 0.00017528 | 0.001349 | m | 0.191 | 0.00083517 | 0.0056831 | 0.0036438 | m | OK | 0.191 |  | 0.00567 |  | m | -3.875e-07 | 0.00083517 | 1.3057e-05 | 0.0036438 | m | -2.2356e-06 |  | -3.1526e-06 |  | m | 1 | Hz | 0.001 | 0.000191 | 4.1101e-07 | 5.6668e-06 | 1.7891e-06 | 0.00019108 | 4.1424e-07 | 0.029661 | 0.009359 | 1 | 0m | 0 | 7 | 1 | 1m | 25 | 1 | 0m | 1m | 0 | 0.001 | 1Hz1m0m | 1Hz1m1m | 7 | 25 | 1Hz1m0m | 0 | 0.0028866 | 0 | 0.0034903 | 1Hz1m1m | 6.3792e-08 | 0.0022335 | 8.4515e-08 | 0.0028142 | 1.219e-08 | 0.0022335 | 8.4515e-08 | 0.0033611 | 0.00019101 | 0.0022335 | 5.7514e-06 | 0.0033611 | 1 | 4.5293e-06 | 8.4505e-05 | 0.0036842 | OK | 1.0001 | 2.4001e-06 | 8.4505e-05 | 0.0036842 | OK | 1 | 2.4001e-06 | 8.4505e-05 | 0.0036842 | 0.00019108 | 8.2847e-07 | 0.029746 | 0.019077 | 0.000191 | 8.3517e-07 | 5.6831e-06 | 3.6438e-06
100 | m | 0.05 | -40.9965 | 0.0013298 | 49.9037 | 0.0021562 | m | -0.00057111 | 0.00082966 | -3.5727e-05 | 0.00071966 | m | -41 | 0.0091635 | 49.9 | 0.007874 | m | OK | -41 |  | 49.9 |  | m | -3.6846e-08 | 0.0091635 | -4.6444e-09 | 0.007874 | m | 0.0040594 |  | 0.0037707 |  | m | 50 | mHz | 0.001 | -0.040996 | 1.5674e-06 | 0.049904 | 2.2732e-06 | 0.064584 | 2.0187e-06 | 2.2585 | 2.9169e-05 | 4 | 10m | 0.01 | 111 | 4 | 100m | 129 | 4 | 10m | 100m | 0.01 | 0.1 | 50mHz100m10m | 50mHz100m100m | 111 | 129 | 50mHz100m10m | -6.5259e-07 | 0.0010456 | 1.0001e-06 | 0.002612 | 50mHz100m100m | 3.3757e-06 | 0.0034795 | 7.6298e-06 | 0.0010886 | 1.7905e-06 | 0.0034795 | 5.0209e-06 | 0.0016881 | -0.040994 | 0.0034795 | 0.049909 | 0.0016881 | 0.99993 | 3.6773e-06 | 0.0001 | 0.00032398 | OK | 1 | 3.5375e-06 | 7.6307e-05 | 6.407e-05 | OK | 0.99999 | 3.5375e-06 | 8.563e-05 | 0.00016635 | 0.064583 | 4.0438e-06 | 2.2586 | 0.00017628 | -0.041 | 9.1635e-06 | 0.0499 | 7.874e-06
100 | m | 50 | 13.7991 | 0.00080556 | -43.5031 | -0.00022527 | m | 0.0016287 | 0.0014977 | -0.00012585 | 0.00020238 | m | 13.8 | 0.010273 | -43.5 | 0.0034526 | m | OK | 13.8 |  | -43.5 |  | m | -2.2174e-07 | 0.010273 | -3.1511e-07 | 0.0034526 | m | -0.0024999 |  | -0.0029609 |  | m | 50 | Hz | 0.001 | 0.013798 | 1.7006e-06 | -0.043503 | 3.0283e-07 | 0.045639 | 5.8962e-07 | -1.2637 | 3.5576e-05 | 4 | 10m | 0.01 | 120 | 4 | 100m | 138 | 4 | 10m | 100m | 0.01 | 0.1 | 50Hz100m10m | 50Hz100m100m | 120 | 138 | 50Hz100m10m | -6.4421e-07 | 0.0023701 | 9.9992e-07 | 0.0032992 | 50Hz100m100m | -1.6148e-06 | 0.0033878 | 2.8843e-06 | 0.00057296 | -1.0285e-06 | 0.0033878 | 1.7461e-06 | 0.0022197 | 0.013796 | 0.0033878 | -0.043501 | 0.0022197 | 0.99994 | 2.7126e-06 | 9.9999e-05 | 0.00033955 | OK | 0.99998 | 5.0604e-06 | 2.8843e-05 | 5.0361e-05 | OK | 0.99995 | 5.0604e-06 | 7.1823e-05 | 0.00022504 | 0.045636 | 1.2016e-06 | -1.2636 | 0.00023602 | 0.0138 | 1.0273e-05 | -0.0435 | 3.4526e-06
10 | m | 0.01 | 8.9203 | 0.00071158 | -1.3396 | 0.0015722 | m | 0.0010298 | 0.0017956 | 0.0010456 | 0.0016476 | m | 8.92 | 0.0039175 | -1.34 | 0.0053065 | m | OK | 8.92 |  | -1.34 |  | m | 1.6843e-07 | 0.0039175 | -5.1865e-08 | 0.0053065 | m | -0.00074765 |  | -0.00062228 |  | m | 10 | mHz | 0.001 | 0.0089193 | 1.9314e-06 | -0.0013406 | 2.2774e-06 | 0.0090194 | 1.9397e-06 | -0.14919 | 0.00025172 | 3 | 3m | 0.003 | 73 | 3 | 10m | 91 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | -1.2909e-07 | 0.0017511 | 2.9998e-07 | 0.00286 | 10mHz10m10m | 9.0451e-07 | 0.0029623 | 7.7359e-07 | 0.00028384 | 7.5973e-07 | 0.0029623 | 7.0725e-07 | 0.0006447 | 0.00892 | 0.0029623 | -0.0013399 | 0.00064472 | 0.99996 | 4.2448e-06 | 0.0001 | 0.0010851 | OK | 1.0001 | 3.4123e-06 | 7.7369e-05 | 0.00018409 | OK | 1.0001 | 3.4123e-06 | 8.0539e-05 | 0.0003103 | 0.0090201 | 3.8796e-06 | -0.14911 | 0.00059138 | 0.00892 | 3.9175e-06 | -0.00134 | 5.3065e-06
1 | m | 2000 | 0.019576 | 3.0722e-05 | -0.15773 | 0.0019503 | m | 0.00037474 | 0.00031477 | -0.00073319 | 0.0018174 | m | 0.019191 | 0.0013128 | -0.157 | 0.0052555 | m | OK | 0.0192 |  | -0.157 |  | m | -8.6624e-06 | 0.0013128 | -1.104e-06 | 0.0052555 | m | 1.3568e-06 |  | 2.0502e-06 |  | m | 2 | kHz | 0.001 | 1.9201e-05 | 3.1626e-07 | -0.000157 | 2.6658e-06 | 0.00015817 | 2.6464e-06 | -1.4491 | 0.0028506 | 1 | 0m | 0 | 17 | 1 | 1m | 35 | 1 | 0m | 1m | 0 | 0.001 | 2kHz1m0m | 2kHz1m1m | 17 | 35 | 2kHz1m0m | 0 | 0.001595 | 0 | 0.0031824 | 2kHz1m1m | 7.954e-08 | 0.0024715 | -6.5424e-08 | 0.00399 | 1.2581e-08 | 0.0024715 | -6.5424e-08 | 0.0033101 | 1.9214e-05 | 0.0024715 | -0.00015706 | 0.0033101 | 1 | 3.5597e-06 | -6.5297e-05 | 0.0045475 | OK | 1.0001 | 3.5594e-06 | -6.5297e-05 | 0.0045475 | OK | 1 | 3.5594e-06 | -6.5297e-05 | 0.0045475 | 0.00015817 | 5.2927e-06 | -1.4492 | 0.0072927 | 1.9191e-05 | 1.3128e-06 | -0.000157 | 5.2555e-06
10 | m | 2 | -6.349 | 0.0010975 | 3.7517 | 0.0001586 | m | 0.0010872 | 0.00038611 | 0.0013021 | 0.00031403 | m | -6.35 | 0.0025401 | 3.75 | 0.003281 | m | OK | -6.35 |  | 3.75 |  | m | -8.7883e-08 | 0.0025401 | 3.5956e-09 | 0.003281 | m | -0.00012297 |  | 0.00039737 |  | m | 2 | Hz | 0.001 | -0.0063501 | 1.1635e-06 | 0.0037504 | 3.5181e-07 | 0.0073749 | 1.0177e-06 | 2.6081 | 9.0131e-05 | 3 | 3m | 0.003 | 80 | 3 | 10m | 98 | 3 | 3m | 10m | 0.003 | 0.01 | 2Hz10m3m | 2Hz10m10m | 80 | 98 | 2Hz10m3m | 1.0505e-07 | 0.0024012 | 2.9999e-07 | 0.0022815 | 2Hz10m10m | -8.7805e-07 | 0.00035334 | 6.7498e-09 | 0.0022509 | -5.0938e-07 | 0.00035334 | 1.1672e-07 | 0.0022624 | -0.0063506 | 0.00035335 | 0.0037505 | 0.0022624 | 1 | 2.4128e-06 | 0.0001 | 0.00080955 | OK | 0.99991 | 7.4977e-07 | 6.7391e-07 | 0.00024383 | OK | 0.99996 | 7.4977e-07 | 3.7922e-05 | 0.00045598 | 0.0073746 | 2.0353e-06 | 2.6081 | 0.00049032 | -0.00635 | 2.5401e-06 | 0.00375 | 3.281e-06
3 | m | 2000 | 1.449 | 0.00087854 | -0.12293 | 0.0005154 | m | -0.00084203 | 0.00057781 | -0.00082767 | 1.0559e-05 | m | 1.45 | 0.0021128 | -0.122 | 0.0036843 | m | OK | 1.45 |  | -0.122 |  | m | -7.5046e-07 | 0.0021128 | -2.5421e-06 | 0.0036843 | m | -0.00012199 |  | -0.0001014 |  | m | 2 | kHz | 0.001 | 0.0014499 | 1.0515e-06 | -0.0001221 | 5.1551e-07 | 0.001455 | 1.0487e-06 | -0.084017 | 0.00035822 | 2 | 1m | 0.001 | 53 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 8.2417e-08 | 0.0022211 | 1.0261e-07 | 0.0022059 | 2kHz3m3m | 1.8611e-07 | 0.0015749 | -1.8404e-08 | 0.00086881 | 1.0601e-07 | 0.0015749 | 7.5075e-08 | 0.0019017 | 0.00145 | 0.0015749 | -0.00012203 | 0.0019017 | 1.0001 | 2.3327e-06 | 9.8773e-05 | 0.0028087 | OK | 1.0001 | 3.026e-06 | -7.077e-06 | 0.0011652 | OK | 1.0001 | 3.026e-06 | 7.4692e-05 | 0.0024348 | 0.0014551 | 2.0974e-06 | -0.083942 | 0.002538 | 0.00145 | 2.1128e-06 | -0.000122 | 3.6843e-06
3 | m | 0.5 | -1.6416 | 0.00062115 | -1.9395 | 0.00015008 | m | -0.0015249 | 0.0014468 | 0.00044011 | 0.0014658 | m | -1.64 | 0.0044926 | -1.94 | 0.004129 | m | OK | -1.64 |  | -1.94 |  | m | 7.9063e-09 | 0.0044926 | -9.1499e-09 | 0.004129 | m | -7.9866e-05 |  | 5.2522e-05 |  | m | 500 | mHz | 0.001 | -0.0016401 | 1.5745e-06 | -0.0019399 | 1.4734e-06 | 0.0025403 | 1.5164e-06 | -2.2726 | 0.00060355 | 2 | 1m | 0.001 | 42 | 2 | 3m | 60 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | 4.9982e-09 | 0.0024467 | 1e-07 | 0.0004964 | 500mHz3m3m | -2.2045e-08 | 0.001603 | 5.6006e-08 | 0.0015108 | -1.5829e-08 | 0.001603 | 6.6118e-08 | 0.0012777 | -0.0016401 | 0.0016031 | -0.0019399 | 0.0012777 | 1 | 2.595e-06 | 9.9998e-05 | 0.0030775 | OK | 0.99999 | 3.9172e-06 | 1.867e-05 | 0.0012878 | OK | 1 | 3.9172e-06 | 3.7362e-05 | 0.0016991 | 0.0025403 | 3.0328e-06 | -2.2726 | 0.0020843 | -0.00164 | 4.4926e-06 | -0.00194 | 4.129e-06
100 | m | 200 | -32.0972 | 0.00086742 | 5.6913 | 0.0016419 | m | 0.0015387 | 0.00121 | -0.00027186 | 0.00073464 | m | -32.1 | 0.0035268 | 5.69 | 0.010858 | m | OK | -32.1 |  | 5.69 |  | m | 2.7731e-07 | 0.0035268 | 1.5248e-06 | 0.010858 | m | 0.0012452 |  | 0.0015326 |  | m | 200 | Hz | 0.001 | -0.032099 | 1.4888e-06 | 0.0056915 | 1.7987e-06 | 0.032599 | 1.4992e-06 | 2.9661 | 5.4911e-05 | 4 | 10m | 0.01 | 122 | 4 | 100m | 140 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | 2.9465e-07 | 0.0011917 | 1.0011e-06 | 0.0039887 | 200Hz100m100m | 2.8943e-06 | 0.0047552 | -8.7494e-06 | 0.0030253 | 9.4742e-07 | 0.0047552 | -1.4473e-06 | 0.0037468 | -0.032098 | 0.0047552 | 0.0056901 | 0.0037468 | 1 | 2.5925e-06 | 9.9984e-05 | 0.00040413 | OK | 1 | 4.9761e-06 | -8.7485e-05 | 6.7212e-05 | OK | 1 | 4.9761e-06 | 5.291e-05 | 0.00031953 | 0.0326 | 3.0028e-06 | 2.9662 | 0.00033787 | -0.0321 | 3.5268e-06 | 0.00569 | 1.0858e-05
1 | m | 0.2 | 0.32796 | 8.3736e-05 | 0.87844 | 0.0014835 | m | 0.0019034 | 0.00062412 | 0.0005176 | 0.001611 | m | 0.32599 | 0.0036468 | 0.878 | 0.0040661 | m | OK | 0.326 |  | 0.878 |  | m | -5.1748e-06 | 0.0036468 | 1.9231e-06 | 0.0040661 | m | 5.6437e-05 |  | -8.1849e-05 |  | m | 200 | mHz | 0.001 | 0.00032606 | 6.2971e-07 | 0.00087792 | 2.19e-06 | 0.00093651 | 2.0646e-06 | 1.2152 | 0.0010296 | 1 | 0m | 0 | 5 | 1 | 1m | 23 | 1 | 0m | 1m | 0 | 0.001 | 200mHz1m0m | 200mHz1m1m | 5 | 23 | 200mHz1m0m | 0 | 0.003922 | 0 | 0.0028784 | 200mHz1m1m | 6.5082e-08 | 0.00069181 | 9.281e-08 | 0.0025517 | 6.095e-08 | 0.00069181 | 9.281e-08 | 0.0025724 | 0.00032612 | 0.00069181 | 0.00087801 | 0.0025724 | 1 | 4.8649e-06 | 9.281e-05 | 0.0032143 | OK | 1.0001 | 3.4064e-06 | 9.281e-05 | 0.0032143 | OK | 1.0001 | 3.4064e-06 | 9.281e-05 | 0.0032143 | 0.00093657 | 4.1292e-06 | 1.2153 | 0.0038174 | 0.00032599 | 3.6468e-06 | 0.000878 | 4.0661e-06
1 | m | 5 | 0.24758 | 0.0015616 | -0.47643 | 0.0013243 | m | -0.00042406 | 0.00093369 | -0.001437 | 0.0019065 | m | 0.24799 | 0.0043183 | -0.475 | 0.0044122 | m | OK | 0.248 |  | -0.475 |  | m | -7.7601e-06 | 0.0043183 | -4.0518e-06 | 0.0044122 | m | 7.0905e-06 |  | 8.2519e-06 |  | m | 5 | Hz | 0.001 | 0.00024801 | 1.8195e-06 | -0.00047499 | 2.3214e-06 | 0.00053584 | 2.2234e-06 | -1.0896 | 0.0036167 | 1 | 0m | 0 | 9 | 1 | 1m | 27 | 1 | 0m | 1m | 0 | 0.001 | 5Hz1m0m | 5Hz1m1m | 9 | 27 | 5Hz1m0m | 0 | 0.0012574 | 0 | 0.0033072 | 5Hz1m1m | 1.4056e-08 | 0.0025719 | -3.5194e-08 | 0.0033732 | 7.5319e-09 | 0.0025719 | -3.5194e-08 | 0.0033425 | 0.00024801 | 0.0025719 | -0.00047503 | 0.0033425 | 1 | 3.5381e-06 | -3.5195e-05 | 0.0033992 | OK | 1 | 3.684e-06 | -3.5195e-05 | 0.0033992 | OK | 1 | 3.684e-06 | -3.5195e-05 | 0.0033992 | 0.00053584 | 4.4468e-06 | -1.0896 | 0.0079923 | 0.00024799 | 4.3183e-06 | -0.000475 | 4.4122e-06
10 | m | 2000 | -1.4292 | 0.0019285 | 2.3287 | 0.0010065 | m | 0.00052577 | 0.001284 | -0.0014221 | 0.0002044 | m | -1.43 | 0.0047731 | 2.33 | 0.003762 | m | Extrapolated | -1.43 |  | 2.33 |  | m | 5.4558e-07 | 0.0047731 | 6.1576e-07 | 0.003762 | m | 0.00027301 |  | 8.3978e-05 |  | m | 2 | kHz | 0.001 | -0.0014297 | 2.3168e-06 | 0.0023301 | 1.027e-06 | 0.0027338 | 1.4948e-06 | 2.1212 | 0.00074859 | 2 | 1m | 0.001 | 0 | 3 | 10m | 107 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | 7.5924e-08 | 0.0020259 | 3.0754e-07 | 0.0017468 | 2kHz10m10m | 3.5784e-07 | 0.0017525 | 7.0124e-07 | 0.0035239 | 6.5201e-08 | 0.0017525 | 2.9257e-07 | 0.0016792 | -0.0014297 | 0.0017525 | 0.0023304 | 0.0016792 | 1 | 2.6566e-06 | 9.9684e-05 | 0.0011987 | OK | 1 | 3.1544e-06 | 6.7873e-05 | 0.00046414 | OK | 1 | 3.1544e-06 | 0.00010089 | 0.0012266 | 0.0027338 | 2.9896e-06 | 2.1213 | 0.0019355 | -0.00143 | 4.7731e-06 | 0.00233 | 3.762e-06
10 | m | 0.5 | -1.8309 | 0.0011742 | -6.9801 | 0.0015165 | m | -0.00054032 | 0.00021042 | 0.00012099 | 0.0009285 | m | -1.83 | 0.0054069 | -6.98 | 0.0036567 | m | OK | -1.83 |  | -6.98 |  | m | 8.949e-09 | 0.0054069 | -3.2625e-08 | 0.0036567 | m | -0.00031421 |  | -0.00026642 |  | m | 500 | mHz | 0.001 | -0.0018303 | 1.1929e-06 | -0.0069803 | 1.7781e-06 | 0.0072162 | 1.7464e-06 | -1.8272 | 0.00017169 | 3 | 3m | 0.003 | 78 | 3 | 10m | 96 | 3 | 3m | 10m | 0.003 | 0.01 | 500mHz10m3m | 500mHz10m10m | 78 | 96 | 500mHz10m3m | 2.8682e-08 | 0.0017611 | 2.9999e-07 | 0.00026644 | 500mHz10m10m | -8.3937e-07 | 0.0012203 | -1.1645e-07 | 0.0033582 | -4.9416e-07 | 0.0012203 | 4.9161e-08 | 0.0021287 | -0.0018308 | 0.0012203 | -0.0069802 | 0.0021287 | 1 | 3.7079e-06 | 0.0001 | 0.0010654 | OK | 0.99992 | 3.4581e-06 | -1.164e-05 | 0.00042994 | OK | 0.99995 | 3.4581e-06 | 3.2758e-05 | 0.00068267 | 0.0072159 | 3.4929e-06 | -1.8272 | 0.00076416 | -0.00183 | 5.4069e-06 | -0.00698 | 3.6567e-06
10 | m | 20 | 0.11823 | 0.0017425 | -4.7709 | 0.00067266 | m | 0.0016126 | 0.001772 | -0.00093331 | 0.00055434 | m | 0.117 | 0.0075344 | -4.77 | 0.0017563 | m | OK | 0.117 |  | -4.77 |  | m | -6.0328e-09 | 0.0075344 | 8.042e-10 | 0.0017563 | m | -0.00038755 |  | 1.7339e-05 |  | m | 20 | Hz | 0.001 | 0.00011661 | 2.4852e-06 | -0.00477 | 8.7165e-07 | 0.0047714 | 8.735e-07 | -1.5464 | 0.00052072 | 3 | 3m | 0.003 | 83 | 3 | 10m | 101 | 3 | 3m | 10m | 0.003 | 0.01 | 20Hz10m3m | 20Hz10m10m | 83 | 101 | 20Hz10m3m | 7.3898e-08 | 0.0013404 | 3.0002e-07 | 0.0026209 | 20Hz10m10m | -5.06e-07 | 0.003568 | 2.5343e-07 | 0.0011237 | -7.2851e-08 | 0.003568 | 2.8823e-07 | 0.002242 | 0.00011654 | 0.003568 | -0.0047697 | 0.002242 | 1 | 2.651e-06 | 0.0001 | 0.0014598 | OK | 0.99995 | 3.7879e-06 | 2.5347e-05 | 0.00038391 | OK | 1 | 3.7879e-06 | 8.1109e-05 | 0.0011875 | 0.0047714 | 1.7471e-06 | -1.5463 | 0.0015795 | 0.000117 | 7.5344e-06 | -0.00477 | 1.7563e-06
1 | m | 1000 | -0.53689 | 0.00062441 | -0.12301 | 0.0012338 | m | -0.00089701 | 0.0014929 | -0.0010084 | 0.00038828 | m | -0.536 | 0.0032003 | -0.122 | 0.0030657 | m | OK | -0.536 |  | -0.122 |  | m | -9.6802e-07 | 0.0032003 | 4.1785e-06 | 0.0030657 | m | 3.0625e-06 |  | -4.5748e-06 |  | m | 1 | kHz | 0.001 | -0.000536 | 1.6182e-06 | -0.000122 | 1.2934e-06 | 0.00054971 | 1.6038e-06 | -2.9178 | 0.0023855 | 1 | 0m | 0 | 16 | 1 | 1m | 34 | 1 | 0m | 1m | 0 | 0.001 | 1kHz1m0m | 1kHz1m1m | 16 | 34 | 1kHz1m0m | 0 | 0.003178 | 0 | 0.0012633 | 1kHz1m1m | 6.7902e-09 | 0.0014218 | -1.7236e-08 | 0.0028123 | 3.7326e-09 | 0.0014218 | -1.7236e-08 | 0.0021148 | -0.00053599 | 0.0014218 | -0.00012202 | 0.0021148 | 1 | 3.4199e-06 | -1.7154e-05 | 0.0028616 | OK | 1 | 2.8584e-06 | -1.7154e-05 | 0.0028616 | OK | 1 | 2.8584e-06 | -1.7154e-05 | 0.0028616 | 0.00054971 | 3.2075e-06 | -2.9178 | 0.0055633 | -0.000536 | 3.2003e-06 | -0.000122 | 3.0657e-06
1 | m | 0.2 | -0.25763 | 0.001735 | -0.31204 | 0.00011731 | m | -0.0016299 | 0.00044711 | -5.7745e-05 | 0.0010018 | m | -0.25598 | 0.003103 | -0.31201 | 0.0029941 | m | OK | -0.256 |  | -0.312 |  | m | 1.7271e-05 | 0.003103 | -1.4171e-05 | 0.0029941 | m | -4.9616e-06 |  | 1.7783e-05 |  | m | 200 | mHz | 0.001 | -0.000256 | 1.7917e-06 | -0.00031198 | 1.0086e-06 | 0.00040357 | 1.3783e-06 | -2.258 | 0.0037805 | 1 | 0m | 0 | 5 | 1 | 1m | 23 | 1 | 0m | 1m | 0 | 0.001 | 200mHz1m0m | 200mHz1m1m | 5 | 23 | 200mHz1m0m | 0 | 0.003922 | 0 | 0.0028784 | 200mHz1m1m | 6.5082e-08 | 0.00069181 | 9.281e-08 | 0.0025517 | 2.6265e-08 | 0.00069181 | 9.281e-08 | 0.0027465 | -0.00025598 | 0.00069182 | -0.00031189 | 0.0027465 | 1 | 4.8649e-06 | 9.281e-05 | 0.0032143 | OK | 1.0001 | 3.4064e-06 | 9.281e-05 | 0.0032143 | OK | 1 | 3.4064e-06 | 9.281e-05 | 0.0032143 | 0.00040358 | 2.7566e-06 | -2.2579 | 0.0082159 | -0.00025598 | 3.103e-06 | -0.00031201 | 2.9941e-06
3 | m | 20 | 0.14142 | 0.00043409 | -0.983 | 0.00084983 | m | 0.0005086 | 6.9767e-05 | -0.0019001 | 0.0010436 | m | 0.141 | 0.0037764 | -0.981 | 0.0026948 | m | Extrapolated | 0.141 |  | -0.981 |  | m | -5.2749e-09 | 0.0037764 | -2.8216e-09 | 0.0026948 | m | -8.6696e-05 |  | -9.9561e-05 |  | m | 20 | Hz | 0.001 | 0.00014091 | 4.3966e-07 | -0.0009811 | 1.3458e-06 | 0.00099117 | 1.3336e-06 | -1.4281 | 0.00047964 | 1 | 0m | 0 | 0 | 2 | 3m | 65 | 2 | 1m | 3m | 0.001 | 0.003 | 20Hz3m1m | 20Hz3m3m | 47 | 65 | 20Hz3m1m | -8.6689e-08 | 0.0001908 | 9.9995e-08 | 0.0026694 | 20Hz3m3m | -4.3006e-08 | 0.0034367 | -2.9044e-07 | 0.0024599 | -8.6882e-08 | 0.0034367 | 1.0172e-07 | 0.0026703 | 0.00014083 | 0.0034367 | -0.000981 | 0.0026703 | 0.99991 | 2.0905e-07 | 9.9994e-05 | 0.0036971 | OK | 0.99999 | 3.6132e-06 | -9.6807e-05 | 0.00123 | OK | 0.99991 | 3.6132e-06 | 0.00010086 | 0.003708 | 0.00099108 | 2.6673e-06 | -1.428 | 0.00383 | 0.000141 | 3.7764e-06 | -0.000981 | 2.6948e-06
10 | m | 0.05 | 1.9901 | 0.0014315 | -3.1408 | 0.0012416 | m | 0.0002508 | 0.00096034 | -0.00042551 | 0.00093689 | m | 1.99 | 0.0047434 | -3.14 | 0.0038995 | m | OK | 1.99 |  | -3.14 |  | m | -4.6014e-09 | 0.0047434 | -2.2852e-09 | 0.0038995 | m | -0.00017105 |  | -0.00038132 |  | m | 50 | mHz | 0.001 | 0.0019898 | 1.7238e-06 | -0.0031404 | 1.5554e-06 | 0.0037177 | 1.6054e-06 | -1.006 | 0.00045116 | 3 | 3m | 0.003 | 75 | 3 | 10m | 93 | 3 | 3m | 10m | 0.003 | 0.01 | 50mHz10m3m | 50mHz10m10m | 75 | 93 | 50mHz10m3m | -1.8406e-07 | 0.0002001 | 3.0001e-07 | 0.0032907 | 50mHz10m10m | -6.7861e-07 | 0.00056139 | 3.9218e-07 | 0.0014465 | -2.3477e-07 | 0.00056139 | 3.0946e-07 | 0.0031016 | 0.0019896 | 0.0005614 | -0.0031401 | 0.0031016 | 0.99994 | 2.21e-06 | 9.9999e-05 | 0.0011863 | OK | 0.99993 | 2.4301e-06 | 3.9214e-05 | 0.00015768 | OK | 0.99994 | 2.4301e-06 | 9.3767e-05 | 0.0010808 | 0.0037175 | 3.2109e-06 | -1.0059 | 0.001408 | 0.00199 | 4.7434e-06 | -0.00314 | 3.8995e-06
3 | m | 0.5 | -1.1488 | 0.00058288 | -1.0284 | 0.0011385 | m | 0.0013045 | 0.0016978 | 0.0014932 | 0.0008454 | m | -1.15 | 0.0041972 | -1.03 | 0.0043948 | m | OK | -1.15 |  | -1.03 |  | m | 2.3535e-10 | 0.0041972 | -4.4374e-10 | 0.0043948 | m | -7.832e-05 |  | 9.1267e-05 |  | m | 500 | mHz | 0.001 | -0.0011501 | 1.795e-06 | -0.0010299 | 1.4181e-06 | 0.0015438 | 1.638e-06 | -2.4113 | 0.0010344 | 2 | 1m | 0.001 | 42 | 2 | 3m | 60 | 2 | 1m | 3m | 0.001 | 0.003 | 500mHz3m1m | 500mHz3m3m | 42 | 60 | 500mHz3m1m | 4.9982e-09 | 0.0024467 | 1e-07 | 0.0004964 | 500mHz3m3m | -2.2045e-08 | 0.001603 | 5.6006e-08 | 0.0015108 | -2.3551e-09 | 0.001603 | 8.804e-08 | 0.00077224 | -0.0011501 | 0.0016031 | -0.0010298 | 0.00077225 | 1 | 2.595e-06 | 9.9998e-05 | 0.0030775 | OK | 0.99999 | 3.9172e-06 | 1.867e-05 | 0.0012878 | OK | 1 | 3.9172e-06 | 7.7884e-05 | 0.0025908 | 0.0015438 | 3.2761e-06 | -2.4112 | 0.0033154 | -0.00115 | 4.1972e-06 | -0.00103 | 4.3948e-06
3 | m | 100 | 2.0706 | 0.00071073 | -2.0693 | 0.00090182 | m | 0.00076244 | 0.001853 | 0.00075422 | 0.00028246 | m | 2.07 | 0.0041182 | -2.07 | 0.0041182 | m | OK | 2.07 |  | -2.07 |  | m | -2.8748e-08 | 0.0041182 | -1.0192e-07 | 0.0041182 | m | -0.00015067 |  | -5.4928e-05 |  | m | 100 | Hz | 0.001 | 0.0020698 | 1.9846e-06 | -0.0020701 | 9.4502e-07 | 0.0029274 | 1.5543e-06 | -0.78545 | 0.00053098 | 2 | 1m | 0.001 | 49 | 2 | 3m | 67 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | -5.0985e-09 | 0.0019123 | 1.0002e-07 | 0.0010444 | 100Hz3m3m | 7.2517e-08 | 0.0037581 | 1.4332e-07 | 0.0024724 | 6.9698e-08 | 0.0037581 | 1.4174e-07 | 0.0024205 | 0.0020699 | 0.0037581 | -0.0020699 | 0.0024205 | 0.99999 | 2.2442e-06 | 9.9993e-05 | 0.003877 | OK | 1 | 4.6302e-06 | 4.7732e-05 | 0.0012079 | OK | 1 | 4.6302e-06 | 4.9631e-05 | 0.0013049 | 0.0029274 | 3.1086e-06 | -0.7854 | 0.0016824 | 0.00207 | 4.1182e-06 | -0.00207 | 4.1182e-06
3 | m | 100 | -0.71609 | 0.00064973 | -1.639 | 0.0011175 | m | -0.00096453 | 0.00072197 | 0.0009241 | 0.00046454 | m | -0.715 | 0.0050776 | -1.64 | 0.003062 | m | OK | -0.715 |  | -1.64 |  | m | -3.4851e-08 | 0.0050776 | -3.768e-09 | 0.003062 | m | -0.00012555 |  | 6.7438e-05 |  | m | 100 | Hz | 0.001 | -0.00071513 | 9.7128e-07 | -0.0016399 | 1.2102e-06 | 0.0017891 | 1.1753e-06 | -1.982 | 0.00056635 | 2 | 1m | 0.001 | 49 | 2 | 3m | 67 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | -5.0985e-09 | 0.0019123 | 1.0002e-07 | 0.0010444 | 100Hz3m3m | 7.2517e-08 | 0.0037581 | 1.4332e-07 | 0.0024724 | 2.5524e-08 | 0.0037581 | 1.171e-07 | 0.0016078 | -0.0007151 | 0.0037581 | -0.0016398 | 0.0016078 | 0.99999 | 2.2442e-06 | 9.9993e-05 | 0.003877 | OK | 1 | 4.6302e-06 | 4.7732e-05 | 0.0012079 | OK | 1 | 4.6302e-06 | 7.9374e-05 | 0.0028239 | 0.0017891 | 2.3506e-06 | -1.9819 | 0.0030426 | -0.000715 | 5.0776e-06 | -0.00164 | 3.062e-06
10 | m | 0.01 | 5.9515 | 0.0011492 | 7.4481 | 0.00067173 | m | 0.0014224 | 0.00036981 | -0.00081298 | 8.4524e-05 | m | 5.95 | 0.0026903 | 7.45 | 0.0024176 | m | OK | 5.95 |  | 7.45 |  | m | 1.4504e-07 | 0.0026903 | 1.3188e-07 | 0.0024176 | m | 0.00010256 |  | -0.0010766 |  | m | 10 | mHz | 0.001 | 0.0059501 | 1.2073e-06 | 0.0074489 | 6.7702e-07 | 0.0095336 | 9.2062e-07 | 0.8968 | 0.00010842 | 3 | 3m | 0.003 | 73 | 3 | 10m | 91 | 3 | 3m | 10m | 0.003 | 0.01 | 10mHz10m3m | 10mHz10m10m | 73 | 91 | 10mHz10m3m | -1.2909e-07 | 0.0017511 | 2.9998e-07 | 0.00286 | 10mHz10m10m | 9.0451e-07 | 0.0029623 | 7.7359e-07 | 0.00028384 | 8.3565e-07 | 0.0029623 | 7.4204e-07 | 0.00045547 | 0.0059509 | 0.0029623 | 0.0074497 | 0.00045547 | 0.99996 | 4.2448e-06 | 0.0001 | 0.0010851 | OK | 1.0001 | 3.4123e-06 | 7.7369e-05 | 0.00018409 | OK | 1.0001 | 3.4123e-06 | 7.8877e-05 | 0.00024411 | 0.0095344 | 1.8415e-06 | 0.89687 | 0.00032651 | 0.00595 | 2.6903e-06 | 0.00745 | 2.4176e-06
100 | m | 0.5 | -4.5836 | 0.0016701 | -49.5001 | -0.0002626 | m | -0.0013895 | 0.0002613 | 0.0011394 | 0.00039473 | m | -4.58 | 0.013128 | -49.5 | 0.0015834 | m | OK | -4.58 |  | -49.5 |  | m | 1.3604e-07 | 0.013128 | -3.2146e-07 | 0.0015834 | m | -0.0022491 |  | -0.0012509 |  | m | 500 | mHz | 0.001 | -0.0045822 | 1.6904e-06 | -0.049501 | 4.741e-07 | 0.049713 | 4.9713e-07 | -1.6631 | 3.387e-05 | 4 | 10m | 0.01 | 114 | 4 | 100m | 132 | 4 | 10m | 100m | 0.01 | 0.1 | 500mHz100m10m | 500mHz100m100m | 114 | 132 | 500mHz100m10m | 2.1434e-07 | 0.0037621 | 9.9998e-07 | 0.0039641 | 500mHz100m100m | -9.3361e-06 | 0.004472 | -2.9789e-06 | 0.0014135 | -3.9998e-06 | 0.004472 | -7.5574e-07 | 0.0028386 | -0.0045862 | 0.004472 | -0.049502 | 0.0028386 | 1 | 4.0958e-06 | 0.0001 | 0.00041135 | OK | 0.99991 | 4.5855e-06 | -2.9777e-05 | 6.0205e-05 | OK | 0.99997 | 4.5855e-06 | 4.2735e-05 | 0.00025641 | 0.049711 | 1.0201e-06 | -1.6631 | 0.0002652 | -0.00458 | 1.3128e-05 | -0.0495 | 1.5834e-06
10 | m | 0.1 | 2.6502 | 0.00058349 | 0.74536 | 0.00090686 | m | 0.00021155 | 0.0013535 | -0.0013382 | 0.00028503 | m | 2.65 | 0.0029727 | 0.747 | 0.0038555 | m | Extrapolated | 2.65 |  | 0.747 |  | m | 6.8944e-09 | 0.0029727 | 1.2459e-08 | 0.0038555 | m | -3.4099e-05 |  | -0.00030159 |  | m | 100 | mHz | 0.001 | 0.00265 | 1.4739e-06 | 0.0007467 | 9.506e-07 | 0.0027532 | 1.4419e-06 | 0.27466 | 0.00036267 | 2 | 1m | 0.001 | 0 | 3 | 10m | 94 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | 1.1845e-07 | 0.00015425 | 2.9999e-07 | 0.0034865 | 100mHz10m10m | -2.0872e-07 | 0.00018693 | 4.114e-07 | 0.0013083 | 1.2998e-07 | 0.00018693 | 2.9606e-07 | 0.0035633 | 0.0026501 | 0.00018696 | 0.00074699 | 0.0035633 | 1 | 1.6336e-06 | 0.0001 | 0.0011898 | OK | 0.99998 | 3.2915e-06 | 4.1139e-05 | 0.00015822 | OK | 1 | 3.2915e-06 | 0.00010208 | 0.0012261 | 0.0027533 | 2.8839e-06 | 0.27476 | 0.0014246 | 0.00265 | 2.9727e-06 | 0.000747 | 3.8555e-06
10 | m | 0.02 | 2.3704 | 0.001543 | -2.4217 | 0.0010755 | m | 0.00039649 | 0.00031567 | -0.0012185 | 0.0018116 | m | 2.37 | 0.0041355 | -2.42 | 0.0041197 | m | OK | 2.37 |  | -2.42 |  | m | -5.4259e-09 | 0.0041355 | 7.2059e-09 | 0.0041197 | m | -3.0014e-05 |  | -0.00044974 |  | m | 20 | mHz | 0.001 | 0.00237 | 1.575e-06 | -0.0024204 | 2.1068e-06 | 0.0033875 | 1.8655e-06 | -0.79594 | 0.00054742 | 3 | 3m | 0.003 | 74 | 3 | 10m | 92 | 3 | 3m | 10m | 0.003 | 0.01 | 20mHz10m3m | 20mHz10m10m | 74 | 92 | 20mHz10m3m | -2.9375e-07 | 9.3614e-05 | 3.0001e-07 | 0.0019409 | 20mHz10m10m | 6.9304e-07 | 0.0022429 | 8.5965e-07 | 0.0012619 | -2.3912e-07 | 0.0022429 | 3.3099e-07 | 0.0019033 | 0.0023697 | 0.0022429 | -0.0024201 | 0.0019033 | 0.9999 | 2.9689e-06 | 0.0001 | 0.00076645 | OK | 1.0001 | 3.1926e-06 | 8.5979e-05 | 0.00041325 | OK | 0.99991 | 3.1926e-06 | 9.9224e-05 | 0.0007469 | 0.0033872 | 3.731e-06 | -0.79584 | 0.0013253 | 0.00237 | 4.1355e-06 | -0.00242 | 4.1197e-06
3 | m | 100 | 1.1304 | 0.001616 | -0.38436 | 0.0016705 | m | 0.00047512 | 0.00079057 | 0.0017508 | 0.0019273 | m | 1.13 | 0.0041645 | -0.386 | 0.0063472 | m | OK | 1.13 |  | -0.386 |  | m | 2.3969e-09 | 0.0041645 | -1.2695e-08 | 0.0063472 | m | -3.4183e-05 |  | -0.00010811 |  | m | 100 | Hz | 0.001 | 0.00113 | 1.799e-06 | -0.00038611 | 2.5506e-06 | 0.0011941 | 1.8916e-06 | -0.32926 | 0.0020791 | 2 | 1m | 0.001 | 49 | 2 | 3m | 67 | 2 | 1m | 3m | 0.001 | 0.003 | 100Hz3m1m | 100Hz3m3m | 49 | 67 | 100Hz3m1m | -5.0985e-09 | 0.0019123 | 1.0002e-07 | 0.0010444 | 100Hz3m3m | 7.2517e-08 | 0.0037581 | 1.4332e-07 | 0.0024724 | 2.4345e-09 | 0.0037581 | 1.0422e-07 | 0.001183 | 0.00113 | 0.0037581 | -0.000386 | 0.001183 | 0.99999 | 2.2442e-06 | 9.9993e-05 | 0.003877 | OK | 1 | 4.6302e-06 | 4.7732e-05 | 0.0012079 | OK | 1 | 4.6302e-06 | 9.4921e-05 | 0.0036179 | 0.0011941 | 3.7832e-06 | -0.32917 | 0.0055118 | 0.00113 | 4.1645e-06 | -0.000386 | 6.3472e-06
10 | m | 200 | 2.8484 | 0.0010429 | -2.0497 | 0.00094232 | m | -0.0013109 | 0.0010649 | 0.0005312 | 0.00069322 | m | 2.85 | 0.0037964 | -2.05 | 0.0045446 | m | OK | 2.85 |  | -2.05 |  | m | -1.1957e-07 | 0.0037964 | 7.8059e-08 | 0.0045446 | m | -0.00030291 |  | -0.00021294 |  | m | 200 | Hz | 0.001 | 0.0028497 | 1.4905e-06 | -0.0020502 | 1.1698e-06 | 0.0035106 | 1.3895e-06 | -0.62366 | 0.00036695 | 3 | 3m | 0.003 | 86 | 3 | 10m | 104 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | 1.2877e-07 | 0.0013144 | 3.0043e-07 | 0.0037855 | 200Hz10m10m | -7.3614e-07 | 0.0025075 | 9.4798e-07 | 0.00019174 | 6.5686e-08 | 0.0025075 | 3.4766e-07 | 0.0035234 | 0.0028498 | 0.0025075 | -0.0020499 | 0.0035234 | 1 | 1.8789e-06 | 9.9998e-05 | 0.001369 | OK | 0.99993 | 3.7485e-06 | 9.4847e-05 | 0.00036911 | OK | 1 | 3.7485e-06 | 9.9623e-05 | 0.0012961 | 0.0035107 | 2.779e-06 | -0.62356 | 0.0014895 | 0.00285 | 3.7964e-06 | -0.00205 | 4.5446e-06
3 | m | 5000 | -0.96867 | 0.0016487 | -1.2904 | 0.0018781 | m | -0.00063296 | 0.0015517 | -0.00045566 | 0.0010974 | m | -0.968 | 0.0046861 | -1.29 | 0.0045698 | m | OK | -0.968 |  | -1.29 |  | m | -4.3364e-06 | 0.0046861 | 2.7626e-05 | 0.0045698 | m | -4.1344e-05 |  | 7.8371e-05 |  | m | 5 | kHz | 0.001 | -0.00096804 | 2.2641e-06 | -0.0012899 | 2.1753e-06 | 0.0016128 | 2.2077e-06 | -2.2146 | 0.0013843 | 2 | 1m | 0.001 | 54 | 2 | 3m | 72 | 2 | 1m | 3m | 0.001 | 0.003 | 5kHz3m1m | 5kHz3m3m | 54 | 72 | 5kHz3m1m | 6.676e-08 | 0.0014337 | 1.0383e-07 | 0.00037658 | 5kHz3m3m | -3.4814e-07 | 0.00028055 | -3.0166e-07 | 0.00090262 | -6.0356e-08 | 0.00028055 | -2.0405e-08 | 0.00053775 | -0.0009681 | 0.00028059 | -0.0012899 | 0.00053776 | 1.0001 | 3.2539e-06 | 9.229e-05 | 0.0011818 | OK | 0.99989 | 1.0015e-06 | -8.7391e-05 | 0.0010617 | OK | 1 | 1.0015e-06 | 3.7239e-05 | 0.001145 | 0.0016128 | 4.4153e-06 | -2.2146 | 0.002996 | -0.000968 | 4.6861e-06 | -0.00129 | 4.5698e-06
3 | m | 0.05 | 0.73446 | 0.0008617 | -2.3721 | 0.0016144 | m | 0.0016064 | 0.00022218 | -0.0019488 | 0.00071778 | m | 0.733 | 0.0036248 | -2.37 | 0.003437 | m | OK | 0.733 |  | -2.37 |  | m | -3.7606e-09 | 0.0036248 | 1.7471e-09 | 0.003437 | m | -0.00014512 |  | -0.00017959 |  | m | 50 | mHz | 0.001 | 0.00073285 | 8.8988e-07 | -0.0023702 | 1.7667e-06 | 0.0024809 | 1.7083e-06 | -1.2709 | 0.0004021 | 2 | 1m | 0.001 | 39 | 2 | 3m | 57 | 2 | 1m | 3m | 0.001 | 0.003 | 50mHz3m1m | 50mHz3m3m | 39 | 57 | 50mHz3m1m | 7.9425e-08 | 0.0031084 | 9.998e-08 | 0.0014049 | 50mHz3m3m | -2.9377e-07 | 0.0027194 | 2.0792e-07 | 0.0030941 | -1.9691e-07 | 0.0027194 | 1.799e-07 | 0.0026556 | 0.00073266 | 0.0027194 | -0.00237 | 0.0026556 | 1.0001 | 3.3965e-06 | 0.0001 | 0.0017695 | OK | 0.9999 | 2.8003e-06 | 6.9306e-05 | 0.0010399 | OK | 0.99995 | 2.8003e-06 | 7.7273e-05 | 0.0012292 | 0.0024808 | 3.4165e-06 | -1.2708 | 0.0014689 | 0.000733 | 3.6248e-06 | -0.00237 | 3.437e-06
1 | m | 10 | -0.0069438 | 0.0011744 | 0.0029453 | 0.0019472 | m | 0.00032621 | 0.0014951 | -0.0016847 | 5.5455e-05 | m | -0.0072695 | 0.0038412 | 0.0046307 | 0.0038579 | m | OK | -0.00727 |  | 0.00463 |  | m | 4.5812e-07 | 0.0038412 | 7.1925e-07 | 0.0038579 | m | -7.3165e-09 |  | -4.1285e-09 |  | m | 10 | Hz | 0.001 | -7.27e-06 | 1.9012e-06 | 4.63e-06 | 1.948e-06 | 8.6192e-06 | 1.9148e-06 | 2.5745 | 0.22446 | 1 | 0m | 0 | 10 | 1 | 1m | 28 | 1 | 0m | 1m | 0 | 0.001 | 10Hz1m0m | 10Hz1m1m | 10 | 28 | 10Hz1m0m | 0 | 0.0029676 | 0 | 0.0030462 | 10Hz1m1m | -5.3218e-08 | 0.0021012 | -9.9829e-08 | 0.0005903 | -4.587e-10 | 0.0021012 | -9.9829e-08 | 0.003025 | -7.2705e-06 | 0.0021012 | 4.5302e-06 | 0.003025 | 1 | 4.2528e-06 | -9.9797e-05 | 0.0025687 | OK | 0.99995 | 4.19e-06 | -9.9797e-05 | 0.0025687 | OK | 1 | 4.19e-06 | -9.9797e-05 | 0.0025687 | 8.6192e-06 | 3.8297e-06 | 2.5744 | 0.44892 | -7.2695e-06 | 3.8412e-06 | 4.6307e-06 | 3.8579e-06
10 | m | 0.2 | -2.811 | 0.00078446 | -2.6 | 0.0017817 | m | -0.00082085 | 0.0018107 | -0.00024535 | 0.00072026 | m | -2.81 | 0.0046495 | -2.6 | 0.004764 | m | OK | -2.81 |  | -2.6 |  | m | 4.4107e-09 | 0.0046495 | -3.3148e-09 | 0.004764 | m | -0.00019693 |  | 0.00029169 |  | m | 200 | mHz | 0.001 | -0.0028102 | 1.9733e-06 | -0.0025997 | 1.9218e-06 | 0.0038283 | 1.9497e-06 | -2.3951 | 0.00050825 | 3 | 3m | 0.003 | 77 | 3 | 10m | 95 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | 3.6511e-08 | 0.0027181 | 2.9999e-07 | 0.0017498 | 200mHz10m10m | 2.7533e-07 | 0.0036655 | 2.278e-07 | 0.00043877 | 6.4769e-08 | 0.0036655 | 2.9145e-07 | 0.0015947 | -0.0028101 | 0.0036655 | -0.0025994 | 0.0015947 | 1 | 2.9826e-06 | 0.0001 | 0.0010669 | OK | 1 | 3.7536e-06 | 2.2779e-05 | 0.00029937 | OK | 1 | 3.7536e-06 | 9.0863e-05 | 0.00097607 | 0.0038283 | 3.8995e-06 | -2.395 | 0.0014093 | -0.00281 | 4.6495e-06 | -0.0026 | 4.764e-06
100 | m | 1000 | -37.4981 | 6.817e-05 | 81.5961 | 0.0015553 | m | 0.00046107 | 0.0019237 | 0.0017308 | 0.00064711 | m | -37.4999 | 0.012246 | 81.6001 | 0.00642 | m | OK | -37.5 |  | 81.6 |  | m | 5.2795e-05 | 0.012246 | 0.00010878 | 0.00642 | m | 0.0014704 |  | -0.005607 |  | m | 1 | kHz | 0.001 | -0.037499 | 1.9249e-06 | 0.081594 | 1.6846e-06 | 0.089799 | 1.7289e-06 | 2.0016 | 2.0994e-05 | 4 | 10m | 0.01 | 124 | 4 | 100m | 142 | 4 | 10m | 100m | 0.01 | 0.1 | 1kHz100m10m | 1kHz100m100m | 124 | 142 | 1kHz100m10m | -2.8772e-07 | 0.00043023 | 1.0046e-06 | 0.0036529 | 1kHz100m100m | 7.7068e-06 | 0.0040572 | -2.6087e-06 | 0.00017136 | 6.8006e-06 | 0.0040572 | -2.1991e-06 | 0.00056599 | -0.037492 | 0.0040572 | 0.081592 | 0.000566 | 0.99997 | 4.4875e-07 | 9.9884e-05 | 0.000453 | OK | 1.0001 | 4.2632e-06 | -2.6568e-05 | 0.00010334 | OK | 1.0001 | 4.2632e-06 | -1.2235e-05 | 0.00014297 | 0.089804 | 3.4789e-06 | 2.0016 | 0.00014901 | -0.0375 | 1.2246e-05 | 0.0816 | 6.42e-06
10 | m | 500 | 2.92 | 0.0014171 | -1.4906 | 0.0004849 | m | 0.00036764 | 0.00060568 | -0.00045425 | 0.0016793 | m | 2.92 | 0.0032298 | -1.49 | 0.0033883 | m | OK | 2.92 |  | -1.49 |  | m | -1.2091e-08 | 0.0032298 | -1.5902e-08 | 0.0033883 | m | -0.00035 |  | -0.00018879 |  | m | 500 | Hz | 0.001 | 0.0029197 | 1.5411e-06 | -0.0014902 | 1.7479e-06 | 0.003278 | 1.5861e-06 | -0.47193 | 0.00052083 | 3 | 3m | 0.003 | 87 | 3 | 10m | 105 | 3 | 3m | 10m | 0.003 | 0.01 | 500Hz10m3m | 500Hz10m10m | 87 | 105 | 500Hz10m3m | 2.0609e-07 | 0.0020148 | 3.0181e-07 | 0.00012312 | 500Hz10m10m | 6.7942e-07 | 0.0037438 | 9.6332e-07 | 0.0015915 | 2.2488e-07 | 0.0037438 | 3.2808e-07 | 0.00018143 | 0.0029199 | 0.0037438 | -0.0014899 | 0.00018146 | 1.0001 | 4.173e-06 | 9.9995e-05 | 0.00012568 | OK | 1.0001 | 3.7612e-06 | 9.5716e-05 | 0.00030464 | OK | 1.0001 | 3.7612e-06 | 9.9825e-05 | 0.00013279 | 0.0032782 | 3.1721e-06 | -0.47183 | 0.0010501 | 0.00292 | 3.2298e-06 | -0.00149 | 3.3883e-06
3 | m | 0.01 | -0.20481 | 0.00084954 | 0.97646 | 0.00079322 | m | -0.00092262 | 0.00094561 | -0.0014963 | 0.00062045 | m | -0.204 | 0.0043255 | 0.978 | 0.0021886 | m | Extrapolated | -0.204 |  | 0.978 |  | m | 2.2222e-09 | 0.0043255 | -3.3694e-09 | 0.0021886 | m | 0.0001118 |  | -4.6682e-05 |  | m | 10 | mHz | 0.001 | -0.00020389 | 1.2712e-06 | 0.00097795 | 1.0071e-06 | 0.00099898 | 1.0194e-06 | 1.7763 | 0.0012626 | 1 | 0m | 0 | 0 | 2 | 3m | 55 | 2 | 1m | 3m | 0.001 | 0.003 | 10mHz3m1m | 10mHz3m3m | 37 | 55 | 10mHz3m1m | 6.8576e-08 | 0.0002082 | 1e-07 | 0.00039678 | 10mHz3m3m | 3.4393e-08 | 0.0019573 | 2.3204e-07 | 0.001534 | 6.8593e-08 | 0.0019573 | 9.9933e-08 | 0.00039621 | -0.00020382 | 0.0019573 | 0.00097805 | 0.00039621 | 1.0001 | 1.1188e-06 | 9.9999e-05 | 0.0036046 | OK | 1 | 2.2411e-06 | 7.7357e-05 | 0.00059275 | OK | 1.0001 | 2.2411e-06 | 0.00010001 | 0.0036061 | 0.00099905 | 2.0388e-06 | 1.7764 | 0.0044023 | -0.000204 | 4.3255e-06 | 0.000978 | 2.1886e-06
1 | m | 500 | 0.16425 | 0.00070466 | -0.19017 | 0.0019444 | m | 0.0012454 | 0.00077135 | -0.00017896 | 0.00090821 | m | 0.16299 | 0.0033742 | -0.19001 | 0.0034161 | m | OK | 0.163 |  | -0.19 |  | m | -1.31e-05 | 0.0033742 | -1.1255e-05 | 0.0034161 | m | 1.5488e-06 |  | 7.0173e-06 |  | m | 500 | Hz | 0.001 | 0.000163 | 1.0448e-06 | -0.00018999 | 2.1461e-06 | 0.00025033 | 1.7652e-06 | -0.86171 | 0.0064182 | 1 | 0m | 0 | 15 | 1 | 1m | 33 | 1 | 0m | 1m | 0 | 0.001 | 500Hz1m0m | 500Hz1m1m | 15 | 33 | 500Hz1m0m | 0 | 0.0016014 | 0 | 0.00068694 | 500Hz1m1m | 6.9724e-08 | 0.0037316 | -9.188e-08 | 0.00077467 | 1.7454e-08 | 0.0037316 | -9.188e-08 | 0.0007089 | 0.00016302 | 0.0037316 | -0.00019008 | 0.00070891 | 1 | 1.7425e-06 | -9.1937e-05 | 0.002059 | OK | 1.0001 | 4.5382e-06 | -9.1937e-05 | 0.002059 | OK | 1 | 4.5382e-06 | -9.1937e-05 | 0.002059 | 0.00025034 | 3.5303e-06 | -0.86181 | 0.013001 | 0.00016299 | 3.3742e-06 | -0.00019001 | 3.4161e-06
10 | m | 0.2 | -3.8288 | 0.00094842 | -6.8714 | 4.9439e-05 | m | 0.0014726 | 0.00068078 | -0.0017438 | 0.00018155 | m | -3.83 | 0.0041183 | -6.87 | 0.0024975 | m | OK | -3.83 |  | -6.87 |  | m | 4.4223e-08 | 0.0041183 | -1.4674e-08 | 0.0024975 | m | -0.00023071 |  | 0.0003344 |  | m | 200 | mHz | 0.001 | -0.0038302 | 1.1675e-06 | -0.0068697 | 1.8816e-07 | 0.0078653 | 5.918e-07 | -2.0794 | 0.00013016 | 3 | 3m | 0.003 | 77 | 3 | 10m | 95 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | 3.6511e-08 | 0.0027181 | 2.9999e-07 | 0.0017498 | 200mHz10m10m | 2.7533e-07 | 0.0036655 | 2.278e-07 | 0.00043877 | 2.025e-07 | 0.0036655 | 2.4982e-07 | 0.00083859 | -0.00383 | 0.0036655 | -0.0068694 | 0.00083859 | 1 | 2.9826e-06 | 0.0001 | 0.0010669 | OK | 1 | 3.7536e-06 | 2.2779e-05 | 0.00029937 | OK | 1 | 3.7536e-06 | 4.6328e-05 | 0.00053343 | 0.0078655 | 1.184e-06 | -2.0794 | 0.00059356 | -0.00383 | 4.1183e-06 | -0.00687 | 2.4975e-06
100 | m | 1 | -38.3982 | -0.00019683 | 26.7042 | 0.0006613 | m | 0.00011676 | 9.3415e-05 | -1.0996e-05 | 0.0011276 | m | -38.4 | 0.0044754 | 26.7 | 0.0062362 | m | OK | -38.4 |  | 26.7 |  | m | -8.349e-08 | 0.0044754 | 9.4401e-08 | 0.0062362 | m | 0.0017326 |  | 0.0042189 |  | m | 1 | Hz | 0.001 | -0.038398 | 2.1787e-07 | 0.026704 | 1.3072e-06 | 0.046771 | 7.6749e-07 | 2.5339 | 2.3099e-05 | 4 | 10m | 0.01 | 115 | 4 | 100m | 133 | 4 | 10m | 100m | 0.01 | 0.1 | 1Hz100m10m | 1Hz100m100m | 115 | 133 | 1Hz100m10m | -3.7732e-07 | 0.0039589 | 1e-06 | 0.0021272 | 1Hz100m100m | 3.0133e-07 | 0.0017307 | 8.828e-06 | 9.3545e-05 | -1.0005e-07 | 0.0017307 | 4.1983e-06 | 0.0012963 | -0.038398 | 0.0017307 | 0.026708 | 0.0012963 | 0.99996 | 3.9633e-06 | 0.0001 | 0.00021924 | OK | 1 | 3.0339e-06 | 8.8269e-05 | 5.9552e-05 | OK | 0.99998 | 3.0339e-06 | 9.5207e-05 | 0.000154 | 0.04677 | 1.5415e-06 | 2.534 | 0.00016078 | -0.0384 | 4.4754e-06 | 0.0267 | 6.2362e-06
3 | m | 500 | -0.71959 | 1.7565e-05 | -2.7919 | 0.0012934 | m | -0.0016994 | 0.001899 | -0.0019881 | 0.0010501 | m | -0.718 | 0.0053705 | -2.79 | 0.0035313 | m | OK | -0.718 |  | -2.79 |  | m | -1.4665e-06 | 0.0053705 | 1.3891e-06 | 0.0035313 | m | 0.00011342 |  | 4.5379e-05 |  | m | 500 | Hz | 0.001 | -0.00071789 | 1.899e-06 | -0.00279 | 1.666e-06 | 0.0028808 | 1.6814e-06 | -1.8226 | 0.00065447 | 2 | 1m | 0.001 | 51 | 2 | 3m | 69 | 2 | 1m | 3m | 0.001 | 0.003 | 500Hz3m1m | 500Hz3m3m | 51 | 69 | 500Hz3m1m | 9.9295e-08 | 0.0017077 | 1.0053e-07 | 0.0029239 | 500Hz3m3m | 6.0754e-08 | 0.00067813 | -1.3013e-07 | 0.002561 | 6.305e-08 | 0.00067813 | -1.1639e-07 | 0.0025826 | -0.00071782 | 0.00067814 | -0.0027901 | 0.0025826 | 1.0001 | 3.904e-06 | 9.9792e-05 | 0.003246 | OK | 1 | 2.3519e-06 | -4.3343e-05 | 0.0012608 | OK | 1 | 2.3519e-06 | -3.4814e-05 | 0.0013791 | 0.0028809 | 3.3628e-06 | -1.8227 | 0.0019013 | -0.000718 | 5.3705e-06 | -0.00279 | 3.5313e-06
100 | m | 200 | 6.1006 | 0.00064937 | 4.549 | 0.00059169 | m | 0.00033392 | 4.9164e-05 | -0.00026978 | 0.00058858 | m | 6.1 | 0.0023938 | 4.55 | 0.0029386 | m | Extrapolated | 6.1 |  | 4.55 |  | m | 5.7625e-08 | 0.0023938 | -7.7981e-08 | 0.0029386 | m | 0.00029682 |  | -0.00077519 |  | m | 200 | Hz | 0.001 | 0.0061003 | 6.5123e-07 | 0.0045492 | 8.3458e-07 | 0.0076098 | 7.2212e-07 | 0.64077 | 0.00010172 | 3 | 3m | 0.003 | 0 | 4 | 100m | 140 | 4 | 10m | 100m | 0.01 | 0.1 | 200Hz100m10m | 200Hz100m100m | 122 | 140 | 200Hz100m10m | 2.9465e-07 | 0.0011917 | 1.0011e-06 | 0.0039887 | 200Hz100m100m | 2.8943e-06 | 0.0047552 | -8.7494e-06 | 0.0030253 | 2.2561e-07 | 0.0047552 | 1.2601e-06 | 0.0040143 | 0.0061005 | 0.0047552 | 0.0045505 | 0.0040143 | 1 | 2.5925e-06 | 9.9984e-05 | 0.00040413 | OK | 1 | 4.9761e-06 | -8.7485e-05 | 6.7212e-05 | OK | 1 | 4.9761e-06 | 0.00010496 | 0.00041308 | 0.00761 | 1.4447e-06 | 0.64087 | 0.00046046 | 0.0061 | 2.3938e-06 | 0.00455 | 2.9386e-06
3 | m | 50 | -0.31707 | 0.0013444 | 2.5817 | 0.00079152 | m | -0.0010886 | 8.2306e-05 | 0.0017365 | 4.8983e-05 | m | -0.316 | 0.0048527 | 2.58 | 0.001703 | m | OK | -0.316 |  | 2.58 |  | m | 3.4781e-08 | 0.0048527 | 4.7425e-08 | 0.001703 | m | 1.4193e-05 |  | -3.6261e-05 |  | m | 50 | Hz | 0.001 | -0.00031599 | 1.3469e-06 | 0.00258 | 7.9303e-07 | 0.0025992 | 8.04e-07 | 1.6927 | 0.00051568 | 2 | 1m | 0.001 | 48 | 2 | 3m | 66 | 2 | 1m | 3m | 0.001 | 0.003 | 50Hz3m1m | 50Hz3m3m | 48 | 66 | 50Hz3m1m | -7.5343e-08 | 0.00016799 | 9.9984e-08 | 0.0018841 | 50Hz3m3m | 1.1116e-07 | 0.0039934 | -6.1194e-08 | 0.0025636 | 7.379e-08 | 0.0039934 | -2.8898e-08 | 0.0024274 | -0.00031591 | 0.0039934 | 0.0025799 | 0.0024274 | 0.99992 | 2.0935e-06 | 9.9996e-05 | 0.0040936 | OK | 1 | 4.0218e-06 | -2.0421e-05 | 0.00093897 | OK | 1 | 4.0218e-06 | 3.7083e-06 | 0.0015711 | 0.0025993 | 1.608e-06 | 1.6927 | 0.0018794 | -0.000316 | 4.8527e-06 | 0.00258 | 1.703e-06
3 | m | 2 | 1.7101 | 0.0016926 | -1.6 | 0.00033782 | m | 0.00024468 | 0.001114 | 0.00022445 | 0.0013133 | m | 1.71 | 0.0040767 | -1.6 | 0.0041533 | m | OK | 1.71 |  | -1.6 |  | m | -2.4963e-10 | 0.0040767 | 4.0424e-10 | 0.0041533 | m | -0.00013713 |  | -0.00018605 |  | m | 2 | Hz | 0.001 | 0.0017099 | 2.0263e-06 | -0.0016002 | 1.3561e-06 | 0.0023418 | 1.7457e-06 | -0.75228 | 0.00072686 | 2 | 1m | 0.001 | 44 | 2 | 3m | 62 | 2 | 1m | 3m | 0.001 | 0.003 | 2Hz3m1m | 2Hz3m3m | 44 | 62 | 2Hz3m1m | -5.7652e-08 | 0.0018895 | 1.0001e-07 | 0.0014003 | 2Hz3m3m | 3.3348e-08 | 0.001191 | 2.9114e-07 | 0.0019746 | 3.4016e-09 | 0.001191 | 2.2824e-07 | 0.0017856 | 0.0017099 | 0.001191 | -0.0016 | 0.0017856 | 0.99994 | 4.3386e-06 | 0.0001 | 0.0015196 | OK | 1 | 2.0754e-06 | 9.7049e-05 | 0.0012763 | OK | 0.99999 | 2.0754e-06 | 9.802e-05 | 0.0013564 | 0.0023418 | 3.4914e-06 | -0.75218 | 0.0019882 | 0.00171 | 4.0767e-06 | -0.0016 | 4.1533e-06
3 | m | 0.2 | 0.86528 | 0.0012431 | 1.1699 | 0.00044288 | m | -0.0017512 | 5.078e-05 | -2.6252e-06 | 0.0019688 | m | 0.867 | 0.005527 | 1.17 | 0.0047439 | m | OK | 0.867 |  | 1.17 |  | m | 2.1187e-09 | 0.005527 | 1.8828e-09 | 0.0047439 | m | 3.325e-05 |  | -0.00014625 |  | m | 200 | mHz | 0.001 | 0.00086703 | 1.2441e-06 | 0.0011699 | 2.018e-06 | 0.0014561 | 1.7825e-06 | 0.93299 | 0.0010734 | 2 | 1m | 0.001 | 41 | 2 | 3m | 59 | 2 | 1m | 3m | 0.001 | 0.003 | 200mHz3m1m | 200mHz3m3m | 41 | 59 | 200mHz3m1m | 9.3668e-08 | 0.0035594 | 9.9998e-08 | 0.0029736 | 200mHz3m3m | -6.8559e-08 | 0.0011882 | 1.2511e-08 | 0.0026845 | 5.667e-08 | 0.0011882 | 8.0045e-08 | 0.0029077 | 0.00086709 | 0.0011882 | 0.0011699 | 0.0029077 | 1.0001 | 3.7266e-06 | 9.9999e-05 | 0.0046473 | OK | 0.99998 | 1.8904e-06 | 4.1703e-06 | 0.00091812 | OK | 1.0001 | 1.8904e-06 | 7.8144e-05 | 0.0037968 | 0.0014562 | 3.565e-06 | 0.93306 | 0.0043617 | 0.000867 | 5.527e-06 | 0.00117 | 4.7439e-06
10 | m | 0.1 | 2.4698 | 0.001681 | 3.3694 | 0.00083833 | m | -0.00040291 | 0.0011269 | -0.00032427 | 0.001078 | m | 2.47 | 0.0049043 | 3.37 | 0.004221 | m | OK | 2.47 |  | 3.37 |  | m | -1.0796e-09 | 0.0049043 | 1.8305e-08 | 0.004221 | m | 0.00023112 |  | -0.00032143 |  | m | 100 | mHz | 0.001 | 0.0024702 | 2.0238e-06 | 0.0033697 | 1.3656e-06 | 0.0041781 | 1.6262e-06 | 0.93821 | 0.00043583 | 3 | 3m | 0.003 | 76 | 3 | 10m | 94 | 3 | 3m | 10m | 0.003 | 0.01 | 100mHz10m3m | 100mHz10m10m | 76 | 94 | 100mHz10m3m | 1.1845e-07 | 0.00015425 | 2.9999e-07 | 0.0034865 | 100mHz10m10m | -2.0872e-07 | 0.00018693 | 4.114e-07 | 0.0013083 | 6.3383e-08 | 0.00018693 | 3.1874e-07 | 0.0031199 | 0.0024703 | 0.00018698 | 0.00337 | 0.0031199 | 1 | 1.6336e-06 | 0.0001 | 0.0011898 | OK | 0.99998 | 3.2915e-06 | 4.1139e-05 | 0.00015822 | OK | 1 | 3.2915e-06 | 9.0096e-05 | 0.0010161 | 0.0041783 | 3.2525e-06 | 0.9383 | 0.0013388 | 0.00247 | 4.9043e-06 | 0.00337 | 4.221e-06
3 | m | 1 | 1.1916 | 0.001276 | 0.9303 | 0.00093691 | m | 0.0015427 | 0.00107 | 0.0013676 | 0.0013175 | m | 1.19 | 0.0043413 | 0.929 | 0.004896 | m | OK | 1.19 |  | 0.929 |  | m | 1.63e-09 | 0.0043413 | -1.5784e-08 | 0.004896 | m | 7.9465e-05 |  | -6.8063e-05 |  | m | 1 | Hz | 0.001 | 0.0011901 | 1.6652e-06 | 0.00092893 | 1.6167e-06 | 0.0015097 | 1.647e-06 | 0.66278 | 0.0010832 | 2 | 1m | 0.001 | 43 | 2 | 3m | 61 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | 1.5144e-08 | 0.00095295 | 1.0001e-07 | 0.0031072 | 1Hz3m3m | -2.9468e-07 | 0.0018178 | -7.7678e-08 | 0.002125 | -6.3814e-08 | 0.0018178 | 5.4722e-08 | 0.0028569 | 0.00119 | 0.0018178 | 0.00092899 | 0.0028569 | 1 | 3.8175e-06 | 9.9988e-05 | 0.0038541 | OK | 0.9999 | 1.9897e-06 | -2.585e-05 | 0.00071823 | OK | 0.99999 | 1.9897e-06 | 6.7918e-05 | 0.0030549 | 0.0015097 | 3.294e-06 | 0.66284 | 0.0037451 | 0.00119 | 4.3413e-06 | 0.000929 | 4.896e-06
1 | m | 0.05 | 0.023183 | 0.0011323 | 0.011233 | 0.00075629 | m | -0.0017166 | 0.0004606 | 0.0018435 | 0.0015973 | m | 0.0249 | 0.00272 | 0.0093898 | 0.0033287 | m | OK | 0.0249 |  | 0.00939 |  | m | 9.3666e-08 | 0.00272 | -2.4837e-07 | 0.0033287 | m | 3.063e-08 |  | 1.9307e-08 |  | m | 50 | mHz | 0.001 | 2.49e-05 | 1.2224e-06 | 9.39e-06 | 1.7673e-06 | 2.6612e-05 | 1.3027e-06 | 0.36062 | 0.064218 | 1 | 0m | 0 | 3 | 1 | 1m | 21 | 1 | 0m | 1m | 0 | 0.001 | 50mHz1m0m | 50mHz1m1m | 3 | 21 | 50mHz1m0m | 0 | 0.00073944 | 0 | 0.0020477 | 50mHz1m1m | -5.0084e-08 | 0.0015495 | -1.0243e-08 | 0.00059645 | -1.3328e-09 | 0.0015495 | -1.0243e-08 | 0.0020091 | 2.4899e-05 | 0.0015495 | 9.3798e-06 | 0.0020091 | 1 | 2.1771e-06 | -1.0247e-05 | 0.0035261 | OK | 0.99995 | 2.5111e-06 | -1.0247e-05 | 0.0035261 | OK | 1 | 2.5111e-06 | -1.0247e-05 | 0.0035261 | 2.6612e-05 | 2.6054e-06 | 0.36061 | 0.12848 | 2.49e-05 | 2.72e-06 | 9.3898e-06 | 3.3287e-06
1 | m | 2000 | -0.28724 | 0.00099515 | 0.70059 | 0.0017545 | m | 0.0017725 | 0.0019783 | 0.00063995 | 0.00084333 | m | -0.28899 | 0.0053518 | 0.70001 | 0.0042421 | m | OK | -0.289 |  | 0.7 |  | m | 1.1509e-05 | 0.0053518 | 5.853e-06 | 0.0042421 | m | -1.7448e-05 |  | -5.3588e-05 |  | m | 2 | kHz | 0.001 | -0.00028902 | 2.2145e-06 | 0.00069995 | 1.9467e-06 | 0.00075727 | 1.9879e-06 | 1.9624 | 0.0028755 | 1 | 0m | 0 | 17 | 1 | 1m | 35 | 1 | 0m | 1m | 0 | 0.001 | 2kHz1m0m | 2kHz1m1m | 17 | 35 | 2kHz1m0m | 0 | 0.001595 | 0 | 0.0031824 | 2kHz1m1m | 7.954e-08 | 0.0024715 | -6.5424e-08 | 0.00399 | 6.0233e-08 | 0.0024715 | -6.5424e-08 | 0.003794 | -0.00028896 | 0.0024715 | 0.00069988 | 0.003794 | 1 | 3.5597e-06 | -6.5297e-05 | 0.0045475 | OK | 1.0001 | 3.5594e-06 | -6.5297e-05 | 0.0045475 | OK | 1.0001 | 3.5594e-06 | -6.5297e-05 | 0.0045475 | 0.00075731 | 3.9759e-06 | 1.9623 | 0.0073318 | -0.00028899 | 5.3518e-06 | 0.00070001 | 4.2421e-06
10 | m | 1000 | -1.5104 | 0.0016817 | 10.1011 | 0.0010693 | m | 0.00020213 | 0.00012528 | 0.0011675 | 0.00099094 | m | -1.51 | 0.0058187 | 10.1 | 0.0030216 | m | Extrapolated | -1.51 |  | 10.1 |  | m | 1.7355e-05 | 0.0058187 | 1.444e-05 | 0.0030216 | m | -0.00060987 |  | -0.00011156 |  | m | 1 | kHz | 0.001 | -0.0015106 | 1.6864e-06 | 0.0101 | 1.4579e-06 | 0.010212 | 1.4633e-06 | 1.7193 | 0.00016468 | 4 | 10m | 0.01 | 106 | 4 | 100m | 0 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.908e-07 | 0.0034509 | 2.9819e-07 | 0.0014362 | 1kHz10m10m | 7.4889e-09 | 0.002247 | -5.8276e-07 | 0.003049 | 1.6533e-08 | 0.002247 | -6.0947e-07 | 0.0030979 | -0.0015106 | 0.002247 | 0.010099 | 0.0030979 | 0.9999 | 3.9777e-06 | 9.9453e-05 | 0.00062772 | OK | 1 | 2.29e-06 | -5.78e-05 | 0.00047536 | OK | 1 | 2.29e-06 | -6.2568e-05 | 0.00047074 | 0.010212 | 2.9266e-06 | 1.7192 | 0.00057452 | -0.00151 | 5.8187e-06 | 0.0101 | 3.0216e-06
1 | m | 0.1 | 0.043262 | 0.0012993 | 0.42555 | 0.0013359 | m | -0.00024164 | 0.0006767 | 0.00055235 | 0.0011533 | m | 0.043495 | 0.003202 | 0.425 | 0.0035207 | m | OK | 0.0435 |  | 0.425 |  | m | -5.2662e-06 | 0.003202 | 5.3896e-07 | 0.0035207 | m | 3.9545e-06 |  | -1.3774e-07 |  | m | 100 | mHz | 0.001 | 4.3504e-05 | 1.4649e-06 | 0.000425 | 1.7648e-06 | 0.00042722 | 1.762e-06 | 1.4688 | 0.003437 | 1 | 0m | 0 | 4 | 1 | 1m | 22 | 1 | 0m | 1m | 0 | 0.001 | 100mHz1m0m | 100mHz1m1m | 4 | 22 | 100mHz1m0m | 0 | 0.00024824 | 0 | 0.0026528 | 100mHz1m1m | -1.4534e-09 | 0.001512 | 2.1634e-08 | 0.002673 | -6.2094e-10 | 0.001512 | 2.1634e-08 | 0.0026614 | 4.3503e-05 | 0.001512 | 0.00042502 | 0.0026614 | 1 | 2.6644e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | 0.00042722 | 3.5239e-06 | 1.4688 | 0.0074866 | 4.3495e-05 | 3.202e-06 | 0.000425 | 3.5207e-06
10 | m | 2000 | -5.6984 | 0.0002518 | -1.2002 | 0.00079644 | m | 0.0015773 | 0.0001932 | -0.00070873 | 0.0015416 | m | -5.7 | 0.0015878 | -1.2 | 0.0061276 | m | OK | -5.7 |  | -1.2 |  | m | -1.7913e-06 | 0.0015878 | 3.257e-06 | 0.0061276 | m | 6.7359e-05 |  | 0.00053477 |  | m | 2 | kHz | 0.001 | -0.0056999 | 3.1737e-07 | -0.0011995 | 1.7352e-06 | 0.0058248 | 4.7342e-07 | -2.9342 | 0.00029173 | 3 | 3m | 0.003 | 89 | 3 | 10m | 107 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | 7.5924e-08 | 0.0020259 | 3.0754e-07 | 0.0017468 | 2kHz10m10m | 3.5784e-07 | 0.0017525 | 7.0124e-07 | 0.0035239 | 1.8969e-07 | 0.0017525 | 4.6642e-07 | 0.0024639 | -0.0056997 | 0.0017525 | -0.001199 | 0.0024639 | 1 | 2.6566e-06 | 9.9684e-05 | 0.0011987 | OK | 1 | 3.1544e-06 | 6.7873e-05 | 0.00046414 | OK | 1 | 3.1544e-06 | 8.6847e-05 | 0.00090225 | 0.0058249 | 9.4702e-07 | -2.9341 | 0.0010745 | -0.0057 | 1.5878e-06 | -0.0012 | 6.1276e-06
3 | m | 200 | 1.1589 | 0.0010839 | -0.96906 | 0.00065608 | m | -0.0010764 | 0.0019075 | 9.5035e-05 | 0.00014787 | m | 1.16 | 0.0034889 | -0.969 | 0.0034945 | m | OK | 1.16 |  | -0.969 |  | m | 5.4485e-09 | 0.0034889 | -8.4718e-09 | 0.0034945 | m | -4.921e-05 |  | -0.00015394 |  | m | 200 | Hz | 0.001 | 0.00116 | 2.194e-06 | -0.00096915 | 6.7254e-07 | 0.0015115 | 1.738e-06 | -0.69602 | 0.00099131 | 2 | 1m | 0.001 | 50 | 2 | 3m | 68 | 2 | 1m | 3m | 0.001 | 0.003 | 200Hz3m1m | 200Hz3m3m | 50 | 68 | 200Hz3m1m | -4.3583e-08 | 0.002558 | 9.9996e-08 | 0.00064257 | 200Hz3m3m | -9.4686e-08 | 0.00067678 | 2.887e-07 | 7.157e-05 | -5.6653e-08 | 0.00067678 | 1.4826e-07 | 0.00049652 | 0.0011599 | 0.00067679 | -0.00096901 | 0.00049653 | 0.99996 | 4.4907e-06 | 9.9999e-05 | 0.0014006 | OK | 0.99997 | 2.5292e-06 | 9.6222e-05 | 0.00063872 | OK | 0.99996 | 2.5292e-06 | 9.9033e-05 | 0.0012058 | 0.0015115 | 3.476e-06 | -0.69592 | 0.0023205 | 0.00116 | 3.4889e-06 | -0.000969 | 3.4945e-06
1 | m | 0.02 | -0.39102 | 0.0015485 | 0.1417 | 0.00042553 | m | -0.0010324 | 0.00063422 | -0.0012994 | 0.0016502 | m | -0.39 | 0.0034036 | 0.143 | 0.0037069 | m | OK | -0.39 |  | 0.143 |  | m | -1.6053e-06 | 0.0034036 | -4.3767e-06 | 0.0037069 | m | 1.282e-05 |  | -1.1728e-06 |  | m | 20 | mHz | 0.001 | -0.00038999 | 1.6733e-06 | 0.000143 | 1.7042e-06 | 0.00041538 | 1.677e-06 | 2.7901 | 0.004094 | 1 | 0m | 0 | 2 | 1 | 1m | 20 | 1 | 0m | 1m | 0 | 0.001 | 20mHz1m0m | 20mHz1m1m | 2 | 20 | 20mHz1m0m | 0 | 0.002644 | 0 | 0.0016373 | 20mHz1m1m | 7.21e-08 | 0.002925 | 1.9186e-08 | 0.0018052 | 2.9949e-08 | 0.002925 | 1.9186e-08 | 0.0017071 | -0.00038996 | 0.002925 | 0.00014302 | 0.0017071 | 1 | 3.1099e-06 | 1.9197e-05 | 0.003812 | OK | 1.0001 | 3.8738e-06 | 1.9197e-05 | 0.003812 | OK | 1 | 3.8738e-06 | 1.9197e-05 | 0.003812 | 0.00041539 | 3.3541e-06 | 2.7902 | 0.0090319 | -0.00039 | 3.4036e-06 | 0.000143 | 3.7069e-06
100 | m | 2 | -1.9806 | 1.3202e-05 | -6.4687 | -2.9523e-05 | m | 0.00027607 | 0.00041721 | 0.0017271 | 0.0010821 | m | -1.98 | 0.0019545 | -6.47 | 0.0020728 | m | Extrapolated | -1.98 |  | -6.47 |  | m | 1.1841e-08 | 0.0019545 | -2.3173e-10 | 0.0020728 | m | -0.00087205 |  | -0.00040107 |  | m | 2 | Hz | 0.001 | -0.0019809 | 4.1742e-07 | -0.0064704 | 1.0825e-06 | 0.0067668 | 1.0423e-06 | -1.8679 | 7.5313e-05 | 3 | 3m | 0.003 | 0 | 4 | 100m | 134 | 4 | 10m | 100m | 0.01 | 0.1 | 2Hz100m10m | 2Hz100m100m | 116 | 134 | 2Hz100m10m | -9.2929e-07 | 0.00091562 | 1.0001e-06 | 0.0019441 | 2Hz100m100m | -5.2275e-06 | 0.0026094 | -6.3874e-06 | 0.0033282 | -7.7487e-07 | 0.0026094 | 1.2654e-06 | 0.0018944 | -0.0019816 | 0.0026094 | -0.0064691 | 0.0018944 | 0.99991 | 2.7079e-06 | 0.0001 | 0.00023795 | OK | 0.99995 | 2.663e-06 | -6.3872e-05 | 6.1379e-05 | OK | 0.99991 | 2.663e-06 | 0.00010589 | 0.00024429 | 0.0067662 | 2.0846e-06 | -1.8678 | 0.000287 | -0.00198 | 1.9545e-06 | -0.00647 | 2.0728e-06
100 | m | 1 | 65.1046 | 0.0020095 | 53.7939 | 0.0016098 | m | -0.0005543 | 0.00098936 | -0.00045676 | 0.0013076 | m | 65.1 | 0.0063798 | 53.8 | 0.0071266 | m | OK | 65.1 |  | 53.8 |  | m | 3.6971e-07 | 0.0063798 | 1.6689e-07 | 0.0071266 | m | 0.0051197 |  | -0.0056621 |  | m | 1 | Hz | 0.001 | 0.065105 | 2.2398e-06 | 0.053794 | 2.0739e-06 | 0.084454 | 2.174e-06 | 0.69055 | 2.5372e-05 | 4 | 10m | 0.01 | 115 | 4 | 100m | 133 | 4 | 10m | 100m | 0.01 | 0.1 | 1Hz100m10m | 1Hz100m100m | 115 | 133 | 1Hz100m10m | -3.7732e-07 | 0.0039589 | 1e-06 | 0.0021272 | 1Hz100m100m | 3.0133e-07 | 0.0017307 | 8.828e-06 | 9.3545e-05 | 1.841e-07 | 0.0017307 | 7.4758e-06 | 0.00044482 | 0.065105 | 0.0017307 | 0.053802 | 0.00044484 | 0.99996 | 3.9633e-06 | 0.0001 | 0.00021924 | OK | 1 | 3.0339e-06 | 8.8269e-05 | 5.9552e-05 | OK | 1 | 3.0339e-06 | 9.0295e-05 | 8.7135e-05 | 0.084454 | 4.3556e-06 | 0.69064 | 0.00010083 | 0.0651 | 6.3798e-06 | 0.0538 | 7.1266e-06
100 | m | 2 | -44.103 | 0.0010565 | 8.6419 | 0.00014321 | m | 9.7522e-05 | 0.0019443 | -0.00034891 | 0.0019237 | m | -44.1 | 0.0046268 | 8.64 | 0.0084276 | m | OK | -44.1 |  | 8.64 |  | m | -2.8121e-08 | 0.0046268 | 2.1997e-07 | 0.0084276 | m | -0.0030879 |  | 0.0022712 |  | m | 2 | Hz | 0.001 | -0.044103 | 2.2128e-06 | 0.0086423 | 1.929e-06 | 0.044942 | 2.203e-06 | 2.9481 | 4.3173e-05 | 4 | 10m | 0.01 | 116 | 4 | 100m | 134 | 4 | 10m | 100m | 0.01 | 0.1 | 2Hz100m10m | 2Hz100m100m | 116 | 134 | 2Hz100m10m | -9.2929e-07 | 0.00091562 | 1.0001e-06 | 0.0019441 | 2Hz100m100m | -5.2275e-06 | 0.0026094 | -6.3874e-06 | 0.0033282 | -2.5981e-06 | 0.0026094 | -1.8681e-06 | 0.0024815 | -0.044106 | 0.0026094 | 0.0086404 | 0.0024815 | 0.99991 | 2.7079e-06 | 0.0001 | 0.00023795 | OK | 0.99995 | 2.663e-06 | -6.3872e-05 | 6.1379e-05 | OK | 0.99992 | 2.663e-06 | 3.6379e-05 | 0.0001694 | 0.044938 | 4.4076e-06 | 2.9481 | 0.00019013 | -0.0441 | 4.6268e-06 | 0.00864 | 8.4276e-06
1 | m | 0.1 | -0.36772 | 0.0010775 | 0.54262 | 0.00022082 | m | -0.00072677 | 0.0008776 | 0.00061536 | 0.00087307 | m | -0.367 | 0.0028944 | 0.542 | 0.002522 | m | OK | -0.367 |  | 0.542 |  | m | -4.0506e-06 | 0.0028944 | -2.7429e-06 | 0.002522 | m | 7.3243e-06 |  | 5.7125e-06 |  | m | 100 | mHz | 0.001 | -0.00036699 | 1.3897e-06 | 0.00054201 | 9.0056e-07 | 0.00065456 | 1.0785e-06 | 2.166 | 0.0019198 | 1 | 0m | 0 | 4 | 1 | 1m | 22 | 1 | 0m | 1m | 0 | 0.001 | 100mHz1m0m | 100mHz1m1m | 4 | 22 | 100mHz1m0m | 0 | 0.00024824 | 0 | 0.0026528 | 100mHz1m1m | -1.4534e-09 | 0.001512 | 2.1634e-08 | 0.002673 | -9.5137e-10 | 0.001512 | 2.1634e-08 | 0.002666 | -0.00036699 | 0.001512 | 0.00054203 | 0.002666 | 1 | 2.6644e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | OK | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | 0.00065456 | 2.157e-06 | 2.166 | 0.0048518 | -0.000367 | 2.8944e-06 | 0.000542 | 2.522e-06
100 | m | 500 | -6.5488 | 0.0018051 | 17.0998 | 0.00055287 | m | -0.00060637 | 0.00025952 | -0.00032618 | 0.0019311 | m | -6.55 | 0.0039361 | 17.1 | 0.0039674 | m | OK | -6.55 |  | 17.1 |  | m | 1.1533e-06 | 0.0039361 | -8.6409e-07 | 0.0039674 | m | 0.0017967 |  | 0.0001175 |  | m | 500 | Hz | 0.001 | -0.0065482 | 1.8236e-06 | 0.0171 | 2.0087e-06 | 0.018311 | 1.986e-06 | 1.9365 | 0.00010094 | 4 | 10m | 0.01 | 123 | 4 | 100m | 141 | 4 | 10m | 100m | 0.01 | 0.1 | 500Hz100m10m | 500Hz100m100m | 123 | 141 | 500Hz100m10m | 3.8497e-07 | 0.0024448 | 1.0047e-06 | 0.00044903 | 500Hz100m100m | -6.7771e-06 | 0.0035729 | 3.3804e-06 | 0.00065542 | -2.764e-07 | 0.0035729 | 1.2241e-06 | 0.00046809 | -0.0065485 | 0.0035729 | 0.017101 | 0.0004681 | 1 | 3.9204e-06 | 9.9977e-05 | 7.4691e-05 | OK | 0.99993 | 4.1405e-06 | 3.4002e-05 | 5.5705e-05 | OK | 1 | 4.1405e-06 | 9.3885e-05 | 7.2937e-05 | 0.018312 | 3.9728e-06 | 1.9366 | 0.00021465 | -0.00655 | 3.9361e-06 | 0.0171 | 3.9674e-06
3 | m | 10 | 0.99901 | 0.00013122 | 0.012266 | 0.0019302 | m | 0.0010451 | 0.0017071 | 0.0010662 | 0.00042099 | m | 0.998 | 0.0034246 | 0.0113 | 0.0051001 | m | Extrapolated | 0.998 |  | 0.0113 |  | m | -1.069e-09 | 0.0034246 | -1.0777e-09 | 0.0051001 | m | -3.6551e-05 |  | -0.00010023 |  | m | 10 | Hz | 0.001 | 0.00099796 | 1.7121e-06 | 1.12e-05 | 1.9756e-06 | 0.00099803 | 1.7122e-06 | 0.011222 | 0.0019794 | 1 | 0m | 0 | 0 | 2 | 3m | 64 | 2 | 1m | 3m | 0.001 | 0.003 | 10Hz3m1m | 10Hz3m3m | 46 | 64 | 10Hz3m1m | 3.7737e-08 | 0.0024236 | 1.0002e-07 | 0.0016399 | 10Hz3m3m | 7.8446e-08 | 0.0025183 | 2.6572e-07 | 0.0036641 | 3.7697e-08 | 0.0025183 | 9.9856e-08 | 0.0016379 | 0.000998 | 0.0025183 | 1.13e-05 | 0.0016379 | 1 | 3.7848e-06 | 9.9999e-05 | 0.0032292 | OK | 1 | 3.3627e-06 | 8.8572e-05 | 0.0012379 | OK | 1 | 3.3627e-06 | 0.00010001 | 0.0032312 | 0.00099806 | 3.4244e-06 | 0.011322 | 0.0051101 | 0.000998 | 3.4246e-06 | 1.13e-05 | 5.1001e-06
3 | m | 1 | 0.99577 | 0.0014645 | -0.69965 | 0.0003578 | m | 0.0018367 | 0.0010323 | -0.0015707 | 0.00052686 | m | 0.994 | 0.0037292 | -0.698 | 0.0043373 | m | OK | 0.994 |  | -0.698 |  | m | -1.5625e-08 | 0.0037292 | -4.72e-09 | 0.0043373 | m | -6.3341e-05 |  | -8.3895e-05 |  | m | 1 | Hz | 0.001 | 0.00099394 | 1.7918e-06 | -0.00069808 | 6.3687e-07 | 0.0012146 | 1.5113e-06 | -0.6123 | 0.00095028 | 2 | 1m | 0.001 | 43 | 2 | 3m | 61 | 2 | 1m | 3m | 0.001 | 0.003 | 1Hz3m1m | 1Hz3m3m | 43 | 61 | 1Hz3m1m | 1.5144e-08 | 0.00095295 | 1.0001e-07 | 0.0031072 | 1Hz3m3m | -2.9468e-07 | 0.0018178 | -7.7678e-08 | 0.002125 | -1.8098e-08 | 0.0018178 | 8.0941e-08 | 0.0030018 | 0.00099392 | 0.0018178 | -0.000698 | 0.0030018 | 1 | 3.8175e-06 | 9.9988e-05 | 0.0038541 | OK | 0.9999 | 1.9897e-06 | -2.585e-05 | 0.00071823 | OK | 1 | 1.9897e-06 | 8.6486e-05 | 0.0035177 | 0.0012146 | 3.0226e-06 | -0.61221 | 0.0039983 | 0.000994 | 3.7292e-06 | -0.000698 | 4.3373e-06
10 | m | 2000 | -0.60718 | 0.00040155 | -2.6194 | 0.00017677 | m | -0.0019258 | 0.0003138 | 0.00051185 | 0.00038242 | m | -0.605 | 0.0033785 | -2.62 | 0.0011387 | m | Extrapolated | -0.605 |  | -2.62 |  | m | -7.9416e-07 | 0.0033785 | -5.4375e-08 | 0.0011387 | m | -0.00024989 |  | 0.00012944 |  | m | 2 | kHz | 0.001 | -0.00060525 | 5.0962e-07 | -0.0026199 | 4.213e-07 | 0.0026889 | 4.2622e-07 | -1.7978 | 0.000188 | 2 | 1m | 0.001 | 0 | 3 | 10m | 107 | 3 | 3m | 10m | 0.003 | 0.01 | 2kHz10m3m | 2kHz10m10m | 89 | 107 | 2kHz10m3m | 7.5924e-08 | 0.0020259 | 3.0754e-07 | 0.0017468 | 2kHz10m10m | 3.5784e-07 | 0.0017525 | 7.0124e-07 | 0.0035239 | 6.3394e-08 | 0.0017525 | 2.9004e-07 | 0.0016678 | -0.00060519 | 0.0017525 | -0.0026196 | 0.0016678 | 1 | 2.6566e-06 | 9.9684e-05 | 0.0011987 | OK | 1 | 3.1544e-06 | 6.7873e-05 | 0.00046414 | OK | 1 | 3.1544e-06 | 0.0001011 | 0.0012313 | 0.0026889 | 8.5248e-07 | -1.7977 | 0.0012874 | -0.000605 | 3.3785e-06 | -0.00262 | 1.1387e-06
10 | m | 0.2 | 3.0398 | 0.0017767 | 8.8996 | 0.00053803 | m | -0.00040961 | 0.00074725 | -0.00012437 | 0.00096356 | m | 3.04 | 0.0048502 | 8.9 | 0.0028233 | m | OK | 3.04 |  | 8.9 |  | m | -7.1208e-08 | 0.0048502 | 9.6729e-09 | 0.0028233 | m | 0.00018137 |  | -0.00032265 |  | m | 200 | mHz | 0.001 | 0.0030402 | 1.9275e-06 | 0.0088997 | 1.1036e-06 | 0.0094046 | 1.2161e-06 | 1.2416 | 0.00019762 | 3 | 3m | 0.003 | 77 | 3 | 10m | 95 | 3 | 3m | 10m | 0.003 | 0.01 | 200mHz10m3m | 200mHz10m10m | 77 | 95 | 200mHz10m3m | 3.6511e-08 | 0.0027181 | 2.9999e-07 | 0.0017498 | 200mHz10m10m | 2.7533e-07 | 0.0036655 | 2.278e-07 | 0.00043877 | 2.5501e-07 | 0.0036655 | 2.3394e-07 | 0.00055028 | 0.0030404 | 0.0036655 | 0.0088999 | 0.00055029 | 1 | 2.9826e-06 | 0.0001 | 0.0010669 | OK | 1 | 3.7536e-06 | 2.2779e-05 | 0.00029937 | OK | 1 | 3.7536e-06 | 2.9347e-05 | 0.00036465 | 0.0094049 | 2.4325e-06 | 1.2416 | 0.00053776 | 0.00304 | 4.8502e-06 | 0.0089 | 2.8233e-06
3 | m | 1000 | 0.92573 | 2.4682e-05 | 1.0404 | 0.00098096 | m | 0.00069857 | 0.0013547 | 0.00054767 | 0.00074202 | m | 0.925 | 0.0039536 | 1.04 | 0.003708 | m | OK | 0.925 |  | 1.04 |  | m | 2.6822e-07 | 0.0039536 | -1.2824e-06 | 0.003708 | m | 3.5175e-05 |  | -0.00011616 |  | m | 1 | kHz | 0.001 | 0.00092504 | 1.3549e-06 | 0.0010399 | 1.23e-06 | 0.0013918 | 1.2867e-06 | 0.84378 | 0.00093491 | 2 | 1m | 0.001 | 52 | 2 | 3m | 70 | 2 | 1m | 3m | 0.001 | 0.003 | 1kHz3m1m | 1kHz3m3m | 52 | 70 | 1kHz3m1m | 6.383e-08 | 0.0014878 | 1.0087e-07 | 0.0033407 | 1kHz3m3m | -1.0147e-07 | 0.0035415 | -9.9348e-08 | 0.00031018 | 3.1449e-08 | 0.0035415 | 6.1647e-08 | 0.002747 | 0.00092507 | 0.0035415 | 0.0010399 | 0.002747 | 1.0001 | 3.1667e-06 | 9.9451e-05 | 0.0033148 | OK | 0.99997 | 3.6177e-06 | -3.2539e-05 | 0.0010573 | OK | 1 | 3.6177e-06 | 7.3596e-05 | 0.0028726 | 0.0013918 | 2.5733e-06 | 0.84385 | 0.0034275 | 0.000925 | 3.9536e-06 | 0.00104 | 3.708e-06
10 | m | 1000 | 5.8102 | 0.0012739 | 0.79993 | 0.00018029 | m | -0.00015559 | 0.00020626 | -0.0009018 | 0.00083166 | m | 5.81 | 0.002594 | 0.80099 | 0.003718 | m | OK | 5.81 |  | 0.801 |  | m | 4.6904e-06 | 0.002594 | -5.7516e-06 | 0.003718 | m | 0.00035912 |  | -0.00016462 |  | m | 1 | kHz | 0.001 | 0.0058104 | 1.2905e-06 | 0.00080084 | 8.5098e-07 | 0.0058653 | 1.2836e-06 | 0.13697 | 0.00014683 | 3 | 3m | 0.003 | 88 | 3 | 10m | 106 | 3 | 3m | 10m | 0.003 | 0.01 | 1kHz10m3m | 1kHz10m10m | 88 | 106 | 1kHz10m3m | -2.908e-07 | 0.0034509 | 2.9819e-07 | 0.0014362 | 1kHz10m10m | 7.4889e-09 | 0.002247 | -5.8276e-07 | 0.003049 | -1.687e-07 | 0.002247 | -6.2404e-08 | 0.0020964 | 0.0058102 | 0.002247 | 0.00080077 | 0.0020964 | 0.9999 | 3.9777e-06 | 9.9453e-05 | 0.00062772 | OK | 1 | 2.29e-06 | -5.78e-05 | 0.00047536 | OK | 0.99994 | 2.29e-06 | 3.5085e-05 | 0.00056535 | 0.005865 | 2.5673e-06 | 0.137 | 0.00063708 | 0.00581 | 2.594e-06 | 0.00080099 | 3.718e-06
100 | m | 10 | -25.9982 | -6.6747e-05 | -61.3946 | -0.00034512 | m | -0.0015692 | 0.00018016 | 0.0014734 | 0.00034331 | m | -26 | 0.0076732 | -61.4 | 0.0033611 | m | OK | -26 |  | -61.4 |  | m | 4.9046e-07 | 0.0076732 | -3.0817e-07 | 0.0033611 | m | 0.0033194 |  | 0.0039531 |  | m | 10 | Hz | 0.001 | -0.025997 | 1.9213e-07 | -0.061396 | 4.8679e-07 | 0.066673 | 4.5448e-07 | -1.9713 | 3.8917e-06 | 4 | 10m | 0.01 | 118 | 4 | 100m | 136 | 4 | 10m | 100m | 0.01 | 0.1 | 10Hz100m10m | 10Hz100m100m | 118 | 136 | 10Hz100m10m | 9.1346e-07 | 0.00024891 | 9.9995e-07 | 0.00018397 | 10Hz100m100m | 6.3813e-06 | 0.0038509 | -9.4869e-06 | 0.0017151 | 4.3566e-06 | 0.0038509 | -5.6037e-06 | 0.0011481 | -0.025992 | 0.0038509 | -0.061402 | 0.0011481 | 1.0001 | 1.6713e-06 | 9.9999e-05 | 0.00024636 | OK | 1.0001 | 4.1119e-06 | -9.488e-05 | 5.2973e-05 | OK | 1.0001 | 4.1119e-06 | -2.2717e-05 | 0.00012458 | 0.066678 | 9.494e-07 | -1.9714 | 0.00012483 | -0.026 | 7.6732e-06 | -0.0614 | 3.3611e-06
3 | m | 2000 | -0.15446 | 0.00066371 | 1.2393 | 0.00019635 | m | 0.0014206 | 0.0011839 | -0.00062525 | 0.0011374 | m | -0.156 | 0.0042107 | 1.24 | 0.0023572 | m | OK | -0.156 |  | 1.24 |  | m | 1.9497e-06 | 0.0042107 | -1.5236e-07 | 0.0023572 | m | 0.00012045 |  | -8.5473e-05 |  | m | 2 | kHz | 0.001 | -0.00015588 | 1.3573e-06 | 0.0012399 | 1.1542e-06 | 0.0012497 | 1.1576e-06 | 1.6959 | 0.0010837 | 2 | 1m | 0.001 | 53 | 2 | 3m | 71 | 2 | 1m | 3m | 0.001 | 0.003 | 2kHz3m1m | 2kHz3m3m | 53 | 71 | 2kHz3m1m | 8.2417e-08 | 0.0022211 | 1.0261e-07 | 0.0022059 | 2kHz3m3m | 1.8611e-07 | 0.0015749 | -1.8404e-08 | 0.00086881 | 9.5362e-08 | 0.0015749 | 8.7499e-08 | 0.002039 | -0.00015578 | 0.0015749 | 0.00124 | 0.002039 | 1.0001 | 2.3327e-06 | 9.8773e-05 | 0.0028087 | OK | 1.0001 | 3.026e-06 | -7.077e-06 | 0.0011652 | OK | 1.0001 | 3.026e-06 | 8.5559e-05 | 0.0026035 | 0.0012498 | 2.3153e-06 | 1.6959 | 0.0033877 | -0.000156 | 4.2107e-06 | 0.00124 | 2.3572e-06
10 | m | 200 | -4.6811 | 0.0015411 | -4.6781 | 0.00029721 | m | -0.0005273 | 0.0019993 | 0.0014964 | 0.0017005 | m | -4.68 | 0.0058829 | -4.68 | 0.0058829 | m | OK | -4.68 |  | -4.68 |  | m | 3.5593e-07 | 0.0058829 | 3.9034e-07 | 0.0058829 | m | -0.00053533 |  | 0.00037578 |  | m | 200 | Hz | 0.001 | -0.0046805 | 2.5243e-06 | -0.0046796 | 1.7263e-06 | 0.0066186 | 2.1625e-06 | -2.3563 | 0.00032671 | 3 | 3m | 0.003 | 86 | 3 | 10m | 104 | 3 | 3m | 10m | 0.003 | 0.01 | 200Hz10m3m | 200Hz10m10m | 86 | 104 | 200Hz10m3m | 1.2877e-07 | 0.0013144 | 3.0043e-07 | 0.0037855 | 200Hz10m10m | -7.3614e-07 | 0.0025075 | 9.4798e-07 | 0.00019174 | -3.1834e-07 | 0.0025075 | 6.3518e-07 | 0.0019277 | -0.0046809 | 0.0025075 | -0.004679 | 0.0019277 | 1 | 1.8789e-06 | 9.9998e-05 | 0.001369 | OK | 0.99993 | 3.7485e-06 | 9.4847e-05 | 0.00036911 | OK | 0.99998 | 3.7485e-06 | 9.7335e-05 | 0.00085213 | 0.0066185 | 4.3251e-06 | -2.3562 | 0.0010738 | -0.00468 | 5.8829e-06 | -0.00468 | 5.8829e-06

## Cal Data
RLC bridge calibration data: measured values of reference standards.
1) Make list of ranges, nominal values of standards and frequencies to be calibrated following the example:
Range |  | f | Std. nominal Z
Ohm |  | Hz | Ohm
1 | m | f1 | 0 | m | < bottom of range 1 mOhm
1 | m | f2 | 0 | m
1 | m | … | 0 | m
1 | m | f_max | 0 | m
1 | m | f1 | 1 | m | < top of range 1 mOhm
1 | m | f2 | 1 | m
1 | m | … | 1 | m
1 | m | f_max | 1 | m
3 | m | f1 | 1 | m | < bottom of range 3 mOhm
3 | m | f2 | 1 | m
3 | m | … | 1 | m
3 | m | f_max | 1 | m
3 | m | f1 | 3 | m | < top of range 3 mOhm
3 | m | f2 | 3 | m
3 | m | … | 3 | m
3 | m | f_max | 3 | m
10 | m | f1 | 3 | m | < bottom of range 10 mOhm
Select only frequencies and standard nominal impedances you have available and filled in the sheet "Ref Z". This sheet will automatically try to obtain matching rows from the sheet "Ref Z", but it will not interpolate, so the frequencies and nominal Z must match exactly. It will show warning in case of missing data.
2) Take the SHORT or NULL standard that was used for calibration of your reference impedances, connect to the RLC/EIS meter and perform SHORT correction. Most modern bridges should perform the SHORTing on all ranges and frequencies automatically.
3) Measure residual impedance of the connected SHORT standard to the "sh Rs-Xs" columns and include type A uncertainty (averaging is needed for best accuracy). This step is crucial, because RLC/EIS meters perform their SHORTing with one reading per frequency and range and this process is very noisy. Even just after SHORTing the instrument often shows residual Rs-Xs in tens of µOhms or even more depending on the measurement current. Therefore, for best performance, it is necessary to measure residual short Rs-Xs after SHORTing procedure to this sheet and the sheet will subtract it from the main measurement in columns Rs-Xs. The residual SHORT measurement is not performed for the 0 Ohm range and 0 Ohm nominal (gray area). It is performed only for the yellow areas. Note it is enough to measure the SHORT residuals only once per range and frequency and it can be copied to the other nominal impedance rows of the same range as long as the measurement is performed fast in sequence.
4) Connect first calibrating standard and measure Rs-Xs for each frequency of interest and each physical range of instrument where the standard is being used (manual ranging). Fill in the Rs-Xs and associated uncertainties to the columns Rs, Xs, ua(Rs) and ua(Xs). The averaging of measurements is crucial to reach maximum accuracy of the calibration data.
5) Repeat process for all used calibration standards from step 4).
Notes: Use Rs-Xs suffix "m" for milliOhms or nothing for Ohms. The sheet will recalculate. The same applies for nominal Z and range. Add or remove lines to/from bottom of the table, but always copy formulas in whole width of the sheet (the gray area is referenced by other sheets!) Order of frequencies and standards and ranges is irrelevant. Other sheets will identify positions automaticaly.
[5 empty] | Measured calibration Rs-Xs |  |  |  |  | Measured residual SHORT Rs-Xs | [5 empty] | Processing stuff | [12 empty] | Reference Z (auto search from Ref Z sheet) | [9 empty] | Absolute correction |  |  |  |  | Aux stuff to generate zero Z gain correction |  |  | Gain-phase correction
Range |  | f | Std. nominal Z |  | Rs | ua(Rs) | Xs | ua(Xs) | Rs-Xs unit | sh Rs | sh ua(Rs) | sh Xs | sh ua(Xs) | Rs-Xs unit | Status | Range Z | f |  | Nom Z | Rs-Xs mult | Rs | ua(Rs) | Xs | ua(Xs) | ¦Z¦ | U(Zr) | phi | U(phir) | ref tag | ref ID | Rs | U(Rs) | Xs | U(Xs) | Zr | U(Zr) | phir | U(phir) | tag | cor(Rs) | U(cor(Rs)) | cor(Xs) | U(cor(Xs)) | tag f-rng | tag f-rng' | zero id' | cor(gain) | U(cor(gain)) | cor(phi) | U(cor(phi)) |  | #octave_header#
Ohm |  | Hz | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | - | Ohm | - |  | Ohm | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | - | Ohm | Ohm | Ohm | Ohm | Ohm | Ohm | rad | rad | - | Ohm | Ohm | Ohm | Ohm | - | - | - | - | - | rad | rad
1 | m | 0.01 | 0 | m | 0 | 0.0019664 |  | 0.00060013 | m |  |  |  |  | m | OK | 0.001 | 10 | mHz | 0 | 0.001 | 0 | 1.9664e-06 | 0 | 6.0013e-07 | 0 | 2.0559e-06 | #DIV/0! | 0 | 10mHz0m | 1 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10mHz1m0m | 0 | 0.0039328 | 0 | 0.0012003 | 10mHz1m |  | 19 | 1 | 4.1119e-06 | 1.8536e-05 | 0.0022323 | OK
1 | m | 0.02 | 0 | m | 0 | 0.001322 |  | 0.00081866 | m |  |  |  |  | m | OK | 0.001 | 20 | mHz | 0 | 0.001 | 0 | 1.322e-06 | 0 | 8.1866e-07 | 0 | 1.555e-06 | #DIV/0! | 0 | 20mHz0m | 2 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20mHz1m0m | 0 | 0.002644 | 0 | 0.0016373 | 20mHz1m |  | 20 | 1 | 3.1099e-06 | 1.9197e-05 | 0.003812 | OK
1 | m | 0.05 | 0 | m | 0 | 0.00036972 |  | 0.0010239 | m |  |  |  |  | m | OK | 0.001 | 50 | mHz | 0 | 0.001 | 0 | 3.6972e-07 | 0 | 1.0239e-06 | 0 | 1.0886e-06 | #DIV/0! | 0 | 50mHz0m | 3 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50mHz1m0m | 0 | 0.00073944 | 0 | 0.0020477 | 50mHz1m |  | 21 | 1 | 2.1771e-06 | -1.0247e-05 | 0.0035261 | OK
1 | m | 0.1 | 0 | m | 0 | 0.00012412 |  | 0.0013264 | m |  |  |  |  | m | OK | 0.001 | 100 | mHz | 0 | 0.001 | 0 | 1.2412e-07 | 0 | 1.3264e-06 | 0 | 1.3322e-06 | #DIV/0! | 0 | 100mHz0m | 4 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100mHz1m0m | 0 | 0.00024824 | 0 | 0.0026528 | 100mHz1m |  | 22 | 1 | 2.6644e-06 | 2.1632e-05 | 0.002966 | OK
1 | m | 0.2 | 0 | m | 0 | 0.001961 |  | 0.0014392 | m |  |  |  |  | m | OK | 0.001 | 200 | mHz | 0 | 0.001 | 0 | 1.961e-06 | 0 | 1.4392e-06 | 0 | 2.4324e-06 | #DIV/0! | 0 | 200mHz0m | 5 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200mHz1m0m | 0 | 0.003922 | 0 | 0.0028784 | 200mHz1m |  | 23 | 1 | 4.8649e-06 | 9.281e-05 | 0.0032143 | OK
1 | m | 0.5 | 0 | m | 0 | 0.0017178 |  | 0.0017723 | m |  |  |  |  | m | OK | 0.001 | 500 | mHz | 0 | 0.001 | 0 | 1.7178e-06 | 0 | 1.7723e-06 | 0 | 2.4682e-06 | #DIV/0! | 0 | 500mHz0m | 6 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500mHz1m0m | 0 | 0.0034357 | 0 | 0.0035446 | 500mHz1m |  | 24 | 1 | 4.9364e-06 | -9.923e-05 | 0.0041759 | OK
1 | m | 1 | 0 | m | 0 | 0.0014433 |  | 0.0017452 | m |  |  |  |  | m | OK | 0.001 | 1 | Hz | 0 | 0.001 | 0 | 1.4433e-06 | 0 | 1.7452e-06 | 0 | 2.2647e-06 | #DIV/0! | 0 | 1Hz0m | 7 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1Hz1m0m | 0 | 0.0028866 | 0 | 0.0034903 | 1Hz1m |  | 25 | 1 | 4.5293e-06 | 8.4505e-05 | 0.0036842 | OK
1 | m | 2 | 0 | m | 0 | 0.00050727 |  | 0.0017841 | m |  |  |  |  | m | OK | 0.001 | 2 | Hz | 0 | 0.001 | 0 | 5.0727e-07 | 0 | 1.7841e-06 | 0 | 1.8548e-06 | #DIV/0! | 0 | 2Hz0m | 8 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2Hz1m0m | 0 | 0.0010145 | 0 | 0.0035682 | 2Hz1m |  | 26 | 1 | 3.7096e-06 | -4.7365e-06 | 0.0035658 | OK
1 | m | 5 | 0 | m | 0 | 0.0006287 |  | 0.0016536 | m |  |  |  |  | m | OK | 0.001 | 5 | Hz | 0 | 0.001 | 0 | 6.287e-07 | 0 | 1.6536e-06 | 0 | 1.7691e-06 | #DIV/0! | 0 | 5Hz0m | 9 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5Hz1m0m | 0 | 0.0012574 | 0 | 0.0033072 | 5Hz1m |  | 27 | 1 | 3.5381e-06 | -3.5195e-05 | 0.0033992 | OK
1 | m | 10 | 0 | m | 0 | 0.0014838 |  | 0.0015231 | m |  |  |  |  | m | OK | 0.001 | 10 | Hz | 0 | 0.001 | 0 | 1.4838e-06 | 0 | 1.5231e-06 | 0 | 2.1264e-06 | #DIV/0! | 0 | 10Hz0m | 10 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 10Hz1m0m | 0 | 0.0029676 | 0 | 0.0030462 | 10Hz1m |  | 28 | 1 | 4.2528e-06 | -9.9797e-05 | 0.0025687 | OK
1 | m | 20 | 0 | m | 0 | 0.0016961 |  | 0.0013935 | m |  |  |  |  | m | OK | 0.001 | 20 | Hz | 0 | 0.001 | 0 | 1.6961e-06 | 0 | 1.3935e-06 | 0 | 2.1951e-06 | #DIV/0! | 0 | 20Hz0m | 11 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 20Hz1m0m | 0 | 0.0033921 | 0 | 0.002787 | 20Hz1m |  | 29 | 1 | 4.3902e-06 | 8.1744e-05 | 0.0053652 | OK
1 | m | 50 | 0 | m | 0 | 0.0014121 |  | 0.00010602 | m |  |  |  |  | m | OK | 0.001 | 50 | Hz | 0 | 0.001 | 0 | 1.4121e-06 | 0 | 1.0602e-07 | 0 | 1.4161e-06 | #DIV/0! | 0 | 50Hz0m | 12 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 50Hz1m0m | 0 | 0.0028243 | 0 | 0.00021204 | 50Hz1m |  | 30 | 1 | 2.8322e-06 | 6.8977e-05 | 0.0044647 | OK
1 | m | 100 | 0 | m | 0 | 0.0015582 |  | 0.0017772 | m |  |  |  |  | m | OK | 0.001 | 100 | Hz | 0 | 0.001 | 0 | 1.5582e-06 | 0 | 1.7772e-06 | 0 | 2.3636e-06 | #DIV/0! | 0 | 100Hz0m | 13 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 100Hz1m0m | 0 | 0.0031165 | 0 | 0.0035543 | 100Hz1m |  | 31 | 1 | 4.7271e-06 | 8.0636e-05 | 0.0029048 | OK
1 | m | 200 | 0 | m | 0 | 1.6266e-05 |  | 0.000111 | m |  |  |  |  | m | OK | 0.001 | 200 | Hz | 0 | 0.001 | 0 | 1.6266e-08 | 0 | 1.11e-07 | 0 | 1.1218e-07 | #DIV/0! | 0 | 200Hz0m | 14 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 200Hz1m0m | 0 | 3.2533e-05 | 0 | 0.00022199 | 200Hz1m |  | 32 | 1 | 2.2436e-07 | -2.5193e-05 | 0.00078379 | OK
1 | m | 500 | 0 | m | 0 | 0.00080071 |  | 0.00034347 | m |  |  |  |  | m | OK | 0.001 | 500 | Hz | 0 | 0.001 | 0 | 8.0071e-07 | 0 | 3.4347e-07 | 0 | 8.7126e-07 | #DIV/0! | 0 | 500Hz0m | 15 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 500Hz1m0m | 0 | 0.0016014 | 0 | 0.00068694 | 500Hz1m |  | 33 | 1 | 1.7425e-06 | -9.1937e-05 | 0.002059 | OK
1 | m | 1000 | 0 | m | 0 | 0.001589 |  | 0.00063166 | m |  |  |  |  | m | OK | 0.001 | 1 | kHz | 0 | 0.001 | 0 | 1.589e-06 | 0 | 6.3166e-07 | 0 | 1.71e-06 | #DIV/0! | 0 | 1kHz0m | 16 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 1kHz1m0m | 0 | 0.003178 | 0 | 0.0012633 | 1kHz1m |  | 34 | 1 | 3.4199e-06 | -1.7154e-05 | 0.0028616 | OK
1 | m | 2000 | 0 | m | 0 | 0.00079749 |  | 0.0015912 | m |  |  |  |  | m | OK | 0.001 | 2 | kHz | 0 | 0.001 | 0 | 7.9749e-07 | 0 | 1.5912e-06 | 0 | 1.7799e-06 | #DIV/0! | 0 | 2kHz0m | 17 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 2kHz1m0m | 0 | 0.001595 | 0 | 0.0031824 | 2kHz1m |  | 35 | 1 | 3.5597e-06 | -6.5297e-05 | 0.0045475 | OK
1 | m | 5000 | 0 | m |  | 0.0019168 |  | 0.0010078 | m |  |  |  |  | m | OK | 0.001 | 5 | kHz | 0 | 0.001 | 0 | 1.9168e-06 | 0 | 1.0078e-06 | 0 | 2.1656e-06 | #DIV/0! | 0 | 5kHz0m | 18 | 0 | 0 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | 5kHz1m0m | 0 | 0.0038335 | 0 | 0.0020156 | 5kHz1m |  | 36 | 1 | 4.3311e-06 | -9.5518e-05 | 0.0023963 | OK
1 | m | 0.01 | 1 | m | 0.9999 | 0.00094556 | 0.00031496 | 0.00092696 | m | -0.00025832 | 0.0002374 | 0.00040507 | 0.00062154 | m | OK | 0.001 | 10 | mHz | 0.001 | 0.001 | 0.0010002 | 9.7491e-07 | -9.0109e-08 | 1.116e-06 | 0.0010002 | 9.7491e-07 | -9.0094e-05 | 0.0011159 | 10mHz1m | 19 | 0.0010001 | 1e-08 | -7.1568e-08 | 5e-08 | 0.0010001 | 1e-08 | -7.1558e-05 | 4.9993e-05 | 10mHz1m1m | -3.1471e-08 | 0.0018911 | 1.8541e-08 | 0.0018539 | 10mHz1m | 10mHz1m | 19 | 0.99997 | 1.9498e-06 | 1.8536e-05 | 0.0022323 | OK
1 | m | 0.02 | 1 | m | 1.0009 | 0.0014625 | 0.001397 | 0.0009026 | m | 0.00095294 | 0.0012699 | 0.0015485 | 0.0016784 | m | OK | 0.001 | 20 | mHz | 0.001 | 0.001 | 0.00099992 | 1.9369e-06 | -1.5144e-07 | 1.9057e-06 | 0.00099992 | 1.9369e-06 | -0.00015145 | 0.0019058 | 20mHz1m | 20 | 0.001 | 1e-08 | -1.3225e-07 | 5e-08 | 0.001 | 1e-08 | -0.00013225 | 5e-05 | 20mHz1m1m | 7.21e-08 | 0.002925 | 1.9186e-08 | 0.0018052 | 20mHz1m | 20mHz1m | 20 | 1.0001 | 3.8738e-06 | 1.9197e-05 | 0.003812 | OK
1 | m | 0.05 | 1 | m | 0.99968 | 0.00077474 | 0.00086859 | 0.00029823 | m | -0.0003047 | 0.00098804 | 0.00094023 | 0.0017374 | m | OK | 0.001 | 50 | mHz | 0.001 | 0.001 | 0.00099998 | 1.2556e-06 | -7.1645e-08 | 1.7628e-06 | 0.00099998 | 1.2556e-06 | -7.1646e-05 | 0.0017629 | 50mHz1m | 21 | 0.00099993 | 1e-08 | -8.1888e-08 | 5e-08 | 0.00099993 | 1e-08 | -8.1894e-05 | 5.0003e-05 | 50mHz1m1m | -5.0084e-08 | 0.0015495 | -1.0243e-08 | 0.00059645 | 50mHz1m | 50mHz1m | 21 | 0.99995 | 2.5111e-06 | -1.0247e-05 | 0.0035261 | OK
1 | m | 0.1 | 1 | m | 1.001 | 0.00075601 | -0.0015311 | 0.0013365 | m | 0.00095598 | 0.00072867 | -0.0014171 | 0.00064247 | m | OK | 0.001 | 100 | mHz | 0.001 | 0.001 | 0.0010001 | 1.05e-06 | -1.1399e-07 | 1.4829e-06 | 0.0010001 | 1.05e-06 | -0.00011398 | 0.0014828 | 100mHz1m | 22 | 0.0010001 | 1e-08 | -9.2355e-08 | 5e-08 | 0.0010001 | 1e-08 | -9.2348e-05 | 4.9996e-05 | 100mHz1m1m | -1.4534e-09 | 0.001512 | 2.1634e-08 | 0.002673 | 100mHz1m | 100mHz1m | 22 | 1 | 2.1e-06 | 2.1632e-05 | 0.002966 | OK
1 | m | 0.2 | 1 | m | 1.0016 | 0.00034591 | -0.0010915 | 0.0012758 | m | 0.0016165 | 0.0016677 | -0.001063 | 0.00097688 | m | OK | 0.001 | 200 | mHz | 0.001 | 0.001 | 0.00099996 | 1.7032e-06 | -2.8521e-08 | 1.6069e-06 | 0.00099996 | 1.7032e-06 | -2.8522e-05 | 0.0016069 | 200mHz1m | 23 | 0.001 | 1e-08 | 6.4289e-08 | 5e-08 | 0.001 | 1e-08 | 6.4288e-05 | 4.9999e-05 | 200mHz1m1m | 6.5082e-08 | 0.00069181 | 9.281e-08 | 0.0025517 | 200mHz1m | 200mHz1m | 23 | 1.0001 | 3.4064e-06 | 9.281e-05 | 0.0032143 | OK
1 | m | 0.5 | 1 | m | 0.99916 | 0.0013003 | 0.00078112 | 0.0010753 | m | -0.00079601 | 0.0016757 | 0.00052493 | 0.0017895 | m | OK | 0.001 | 500 | mHz | 0.001 | 0.001 | 0.00099995 | 2.1211e-06 | 2.5619e-07 | 2.0877e-06 | 0.00099995 | 2.1211e-06 | 0.00025621 | 0.0020878 | 500mHz1m | 24 | 0.0010001 | 1e-08 | 1.5698e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00015698 | 4.9997e-05 | 500mHz1m1m | 9.8767e-08 | 0.0026006 | -9.921e-08 | 0.0021505 | 500mHz1m | 500mHz1m | 24 | 1.0001 | 4.2421e-06 | -9.923e-05 | 0.0041759 | OK
1 | m | 1 | 1 | m | 0.99913 | 0.0011168 | 0.0017621 | 0.0014071 | m | -0.0009985 | 0.00043928 | 0.0018616 | 0.001189 | m | OK | 0.001 | 1 | Hz | 0.001 | 0.001 | 0.0010001 | 1.2e-06 | -9.9479e-08 | 1.8422e-06 | 0.0010001 | 1.2e-06 | -9.9467e-05 | 0.0018419 | 1Hz1m | 25 | 0.0010002 | 1e-08 | -1.4965e-08 | 5e-08 | 0.0010002 | 1e-08 | -1.4962e-05 | 4.9991e-05 | 1Hz1m1m | 6.3792e-08 | 0.0022335 | 8.4515e-08 | 0.0028142 | 1Hz1m | 1Hz1m | 25 | 1.0001 | 2.4001e-06 | 8.4505e-05 | 0.0036842 | OK
1 | m | 2 | 1 | m | 0.9984 | 0.0013308 | -0.00046306 | 0.0013971 | m | -0.0014154 | 0.0014369 | -0.00029096 | 0.0011068 | m | OK | 0.001 | 2 | Hz | 0.001 | 0.001 | 0.00099982 | 1.9585e-06 | -1.721e-07 | 1.7824e-06 | 0.00099982 | 1.9585e-06 | -0.00017213 | 0.0017827 | 2Hz1m | 26 | 0.00099992 | 1e-08 | -1.7685e-07 | 5e-08 | 0.00099992 | 1e-08 | -0.00017686 | 5.0004e-05 | 2Hz1m1m | 9.6997e-08 | 0.0026616 | -4.7528e-09 | 0.0027943 | 2Hz1m | 2Hz1m | 26 | 1.0001 | 3.9169e-06 | -4.7365e-06 | 0.0035658 | OK
1 | m | 5 | 1 | m | 1.0014 | 0.0012859 | -0.00050788 | 0.0016866 | m | 0.0013412 | 0.0013188 | -0.00073887 | 0.00020933 | m | OK | 0.001 | 5 | Hz | 0.001 | 0.001 | 0.0010001 | 1.842e-06 | 2.31e-07 | 1.6995e-06 | 0.0010001 | 1.842e-06 | 0.00023098 | 0.0016994 | 5Hz1m | 27 | 0.0010001 | 1e-08 | 1.958e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00019579 | 4.9997e-05 | 5Hz1m1m | 1.4056e-08 | 0.0025719 | -3.5194e-08 | 0.0033732 | 5Hz1m | 5Hz1m | 27 | 1 | 3.684e-06 | -3.5195e-05 | 0.0033992 | OK
1 | m | 10 | 1 | m | 0.99865 | 0.0010506 | -0.0010612 | 0.00029515 | m | -0.001588 | 0.0018125 | -0.0013053 | 0.0012501 | m | OK | 0.001 | 10 | Hz | 0.001 | 0.001 | 0.0010002 | 2.095e-06 | 2.441e-07 | 1.2844e-06 | 0.0010002 | 2.095e-06 | 0.00024404 | 0.0012841 | 10Hz1m | 28 | 0.0010002 | 1e-08 | 1.4427e-07 | 5e-08 | 0.0010002 | 1e-08 | 0.00014425 | 4.9991e-05 | 10Hz1m1m | -5.3218e-08 | 0.0021012 | -9.9829e-08 | 0.0005903 | 10Hz1m | 10Hz1m | 28 | 0.99995 | 4.19e-06 | -9.9797e-05 | 0.0025687 | OK
1 | m | 20 | 1 | m | 1.0004 | 0.00065325 | 0.0017091 | 0.0018889 | m | 0.00042383 | 0.0003597 | 0.0016378 | 0.0019047 | m | OK | 0.001 | 20 | Hz | 0.001 | 0.001 | 0.001 | 7.4574e-07 | 7.1378e-08 | 2.6825e-06 | 0.001 | 7.4574e-07 | 7.1376e-05 | 0.0026825 | 20Hz1m | 29 | 0.0010001 | 1e-08 | 1.5313e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00015312 | 4.9997e-05 | 20Hz1m1m | 3.5509e-08 | 0.0013065 | 8.1751e-08 | 0.0037779 | 20Hz1m | 20Hz1m | 29 | 1 | 1.4915e-06 | 8.1744e-05 | 0.0053652 | OK
1 | m | 50 | 1 | m | 1.0001 | 0.0010505 | 0.00017803 | 0.001196 | m | -4.297e-05 | 0.0014377 | -0.00010296 | 0.0018851 | m | OK | 0.001 | 50 | Hz | 0.001 | 0.001 | 0.0010001 | 1.7806e-06 | 2.8099e-07 | 2.2325e-06 | 0.0010001 | 1.7806e-06 | 0.00028096 | 0.0022322 | 50Hz1m | 30 | 0.0010001 | 1e-08 | 3.4996e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00034994 | 4.9996e-05 | 50Hz1m1m | -3.5489e-08 | 0.002101 | 6.8972e-08 | 0.002392 | 50Hz1m | 50Hz1m | 30 | 0.99996 | 3.5611e-06 | 6.8977e-05 | 0.0044647 | OK
1 | m | 100 | 1 | m | 0.99956 | 0.00086807 | 0.00018547 | 0.0013857 | m | -0.00062321 | 0.0015787 | -0.0006397 | 0.00043535 | m | OK | 0.001 | 100 | Hz | 0.001 | 0.001 | 0.0010002 | 1.8016e-06 | 8.2517e-07 | 1.4525e-06 | 0.0010002 | 1.8016e-06 | 0.00082501 | 0.0014522 | 100Hz1m | 31 | 0.0010003 | 1e-08 | 9.0589e-07 | 5e-08 | 0.0010003 | 1e-08 | 0.00090565 | 4.9987e-05 | 100Hz1m1m | 8.0017e-08 | 0.0017361 | 8.0723e-08 | 0.0027714 | 100Hz1m | 100Hz1m | 31 | 1.0001 | 3.6032e-06 | 8.0636e-05 | 0.0029048 | OK
1 | m | 200 | 1 | m | 1.0008 | 0.00031791 | 0.0029004 | 0.00034946 | m | 0.00011443 | 0.0017177 | 0.0011072 | 0.00017619 | m | OK | 0.001 | 200 | Hz | 0.001 | 0.001 | 0.0010007 | 1.7468e-06 | 1.7933e-06 | 3.9136e-07 | 0.0010007 | 1.7468e-06 | 0.001792 | 0.0003911 | 200Hz1m | 32 | 0.0010007 | 1e-08 | 1.768e-06 | 5e-08 | 0.0010007 | 1e-08 | 0.0017668 | 4.9965e-05 | 200Hz1m1m | -1.5628e-08 | 0.00063582 | -2.5239e-08 | 0.00069892 | 200Hz1m | 200Hz1m | 32 | 0.99998 | 3.4936e-06 | -2.5193e-05 | 0.00078379 | OK
1 | m | 500 | 1 | m | 1.0034 | 0.0018658 | 0.006102 | 0.00038733 | m | 0.00058338 | 0.0012913 | 0.0014004 | 0.00095666 | m | OK | 0.001 | 500 | Hz | 0.001 | 0.001 | 0.0010028 | 2.2691e-06 | 4.7017e-06 | 1.0321e-06 | 0.0010029 | 2.2691e-06 | 0.0046883 | 0.0010292 | 500Hz1m | 33 | 0.0010029 | 1.5811e-08 | 4.6098e-06 | 5e-08 | 0.0010029 | 1.5813e-08 | 0.0045964 | 4.9854e-05 | 500Hz1m1m | 6.9724e-08 | 0.0037316 | -9.188e-08 | 0.00077467 | 500Hz1m | 500Hz1m | 33 | 1.0001 | 4.5382e-06 | -9.1937e-05 | 0.002059 | OK
1 | m | 1000 | 1 | m | 1.0068 | 0.00071089 | 0.0077361 | 0.0014062 | m | -0.0015135 | 0.0012399 | -0.0015195 | 0.00031918 | m | OK | 0.001 | 1 | kHz | 0.001 | 0.001 | 0.0010083 | 1.4292e-06 | 9.2556e-06 | 1.4419e-06 | 0.0010084 | 1.4292e-06 | 0.0091788 | 0.0014299 | 1kHz1m | 34 | 0.0010083 | 4.4721e-08 | 9.2383e-06 | 1e-07 | 0.0010084 | 4.4729e-08 | 0.0091616 | 9.9165e-05 | 1kHz1m1m | 6.7902e-09 | 0.0014218 | -1.7236e-08 | 0.0028123 | 1kHz1m | 1kHz1m | 34 | 1 | 2.8584e-06 | -1.7154e-05 | 0.0028616 | OK
1 | m | 2000 | 1 | m | 1.0224 | 0.0012358 | 0.02 | 0.001995 | m | -0.0010749 | 0.0012805 | 0.0015532 | 0.001195 | m | OK | 0.001 | 2 | kHz | 0.001 | 0.001 | 0.0010235 | 1.7795e-06 | 1.8447e-05 | 2.3256e-06 | 0.0010237 | 1.7797e-06 | 0.018022 | 0.0022717 | 2kHz1m | 35 | 0.0010236 | 1.2649e-07 | 1.8381e-05 | 2e-07 | 0.0010237 | 1.2652e-07 | 0.017956 | 0.00019534 | 2kHz1m1m | 7.954e-08 | 0.0024715 | -6.5424e-08 | 0.00399 | 2kHz1m | 2kHz1m | 35 | 1.0001 | 3.5594e-06 | -6.5297e-05 | 0.0045475 | OK
1 | m | 5000 | 1 | m | 1.0912 | 0.00045037 | 0.046521 | 0.0012702 | m | -0.0017778 | 0.0019716 | 0.00070481 | 0.00019438 | m | OK | 0.001 | 5 | kHz | 0.001 | 0.001 | 0.001093 | 2.0224e-06 | 4.5816e-05 | 1.285e-06 | 0.001094 | 2.0214e-06 | 0.041893 | 0.0011761 | 5kHz1m | 36 | 0.0010929 | 5e-07 | 4.5708e-05 | 5e-07 | 0.0010939 | 5e-07 | 0.041798 | 0.00045709 | 5kHz1m1m | -7.7064e-08 | 0.00090073 | -1.0781e-07 | 0.0025404 | 5kHz1m | 5kHz1m | 36 | 0.99993 | 4.0427e-06 | -9.5518e-05 | 0.0023963 | OK
3 | m | 0.01 | 1 | m | 1.0017 | 0.0001041 | -0.0016242 | 0.00019839 | m | 0.0016627 | 0.00054962 | -0.0014526 | 0.0017913 | m | OK | 0.003 | 10 | mHz | 0.001 | 0.001 | 0.0010001 | 5.594e-07 | -1.7157e-07 | 1.8022e-06 | 0.0010001 | 5.594e-07 | -0.00017156 | 0.0018021 | 10mHz1m | 19 | 0.0010001 | 1e-08 | -7.1568e-08 | 5e-08 | 0.0010001 | 1e-08 | -7.1558e-05 | 4.9993e-05 | 10mHz3m1m | 6.8576e-08 | 0.0002082 | 1e-07 | 0.00039678 | 10mHz3m | 10mHz3m | 37 | 1.0001 | 1.1188e-06 | 9.9999e-05 | 0.0036046 | OK
3 | m | 0.02 | 1 | m | 1.0005 | 0.0014025 | 0.0012029 | 0.00024149 | m | 0.0005027 | 0.00054115 | 0.0014352 | 0.0011112 | m | OK | 0.003 | 20 | mHz | 0.001 | 0.001 | 0.001 | 1.5033e-06 | -2.3226e-07 | 1.1371e-06 | 0.001 | 1.5033e-06 | -0.00023225 | 0.0011371 | 20mHz1m | 20 | 0.001 | 1e-08 | -1.3225e-07 | 5e-08 | 0.001 | 1e-08 | -0.00013225 | 5e-05 | 20mHz3m1m | -1.1939e-08 | 0.0028051 | 1e-07 | 0.00048298 | 20mHz3m | 20mHz3m | 38 | 0.99999 | 3.0066e-06 | 0.0001 | 0.0022748 | OK
3 | m | 0.05 | 1 | m | 0.99887 | 0.0015542 | -0.0014579 | 0.00070243 | m | -0.00098579 | 0.00068451 | -0.001276 | 0.00053711 | m | OK | 0.003 | 50 | mHz | 0.001 | 0.001 | 0.00099985 | 1.6983e-06 | -1.8187e-07 | 8.8425e-07 | 0.00099985 | 1.6983e-06 | -0.00018189 | 0.00088438 | 50mHz1m | 21 | 0.00099993 | 1e-08 | -8.1888e-08 | 5e-08 | 0.00099993 | 1e-08 | -8.1894e-05 | 5.0003e-05 | 50mHz3m1m | 7.9425e-08 | 0.0031084 | 9.998e-08 | 0.0014049 | 50mHz3m | 50mHz3m | 39 | 1.0001 | 3.3965e-06 | 0.0001 | 0.0017695 | OK
3 | m | 0.1 | 1 | m | 1.0018 | 0.00045622 | -0.0020813 | 0.00036579 | m | 0.0016199 | 0.00082142 | -0.0018889 | 0.0015659 | m | OK | 0.003 | 100 | mHz | 0.001 | 0.001 | 0.0010002 | 9.3961e-07 | -1.9238e-07 | 1.608e-06 | 0.0010002 | 9.3961e-07 | -0.00019235 | 0.0016077 | 100mHz1m | 22 | 0.0010001 | 1e-08 | -9.2355e-08 | 5e-08 | 0.0010001 | 1e-08 | -9.2348e-05 | 4.9996e-05 | 100mHz3m1m | -8.9053e-08 | 0.00091244 | 1.0002e-07 | 0.00073157 | 100mHz3m | 100mHz3m | 40 | 0.99991 | 1.8792e-06 | 9.9997e-05 | 0.0032159 | OK
3 | m | 0.2 | 1 | m | 1.0014 | 0.0017797 | -8.1272e-05 | 0.0014868 | m | 0.0014521 | 0.00055198 | -4.5563e-05 | 0.0017853 | m | OK | 0.003 | 200 | mHz | 0.001 | 0.001 | 0.00099993 | 1.8633e-06 | -3.5708e-08 | 2.3233e-06 | 0.00099993 | 1.8633e-06 | -3.5711e-05 | 0.0023235 | 200mHz1m | 23 | 0.001 | 1e-08 | 6.4289e-08 | 5e-08 | 0.001 | 1e-08 | 6.4288e-05 | 4.9999e-05 | 200mHz3m1m | 9.3668e-08 | 0.0035594 | 9.9998e-08 | 0.0029736 | 200mHz3m | 200mHz3m | 41 | 1.0001 | 3.7266e-06 | 9.9999e-05 | 0.0046473 | OK
3 | m | 0.5 | 1 | m | 1.0005 | 0.0012234 | -0.001509 | 0.0002482 | m | 0.00044384 | 0.00043234 | -0.001566 | 0.0015184 | m | OK | 0.003 | 500 | mHz | 0.001 | 0.001 | 0.001 | 1.2975e-06 | 5.6981e-08 | 1.5386e-06 | 0.001 | 1.2975e-06 | 5.6978e-05 | 0.0015385 | 500mHz1m | 24 | 0.0010001 | 1e-08 | 1.5698e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00015698 | 4.9997e-05 | 500mHz3m1m | 4.9982e-09 | 0.0024467 | 1e-07 | 0.0004964 | 500mHz3m | 500mHz3m | 42 | 1 | 2.595e-06 | 9.9998e-05 | 0.0030775 | OK
3 | m | 1 | 1 | m | 0.99902 | 0.00047647 | 0.00034004 | 0.0015536 | m | -0.0011528 | 0.0018483 | 0.00045501 | 0.0011404 | m | OK | 0.003 | 1 | Hz | 0.001 | 0.001 | 0.0010002 | 1.9088e-06 | -1.1497e-07 | 1.9272e-06 | 0.0010002 | 1.9088e-06 | -0.00011495 | 0.0019269 | 1Hz1m | 25 | 0.0010002 | 1e-08 | -1.4965e-08 | 5e-08 | 0.0010002 | 1e-08 | -1.4962e-05 | 4.9991e-05 | 1Hz3m1m | 1.5144e-08 | 0.00095295 | 1.0001e-07 | 0.0031072 | 1Hz3m | 1Hz3m | 43 | 1 | 3.8175e-06 | 9.9988e-05 | 0.0038541 | OK
3 | m | 2 | 1 | m | 1.0014 | 0.00094474 | 0.00021909 | 0.00070016 | m | 0.0014426 | 0.0019528 | 0.00049595 | 0.00029403 | m | OK | 0.003 | 2 | Hz | 0.001 | 0.001 | 0.00099997 | 2.1693e-06 | -2.7686e-07 | 7.5939e-07 | 0.00099997 | 2.1693e-06 | -0.00027686 | 0.00075941 | 2Hz1m | 26 | 0.00099992 | 1e-08 | -1.7685e-07 | 5e-08 | 0.00099992 | 1e-08 | -0.00017686 | 5.0004e-05 | 2Hz3m1m | -5.7652e-08 | 0.0018895 | 1.0001e-07 | 0.0014003 | 2Hz3m | 2Hz3m | 44 | 0.99994 | 4.3386e-06 | 0.0001 | 0.0015196 | OK
3 | m | 5 | 1 | m | 1.0006 | 0.0017171 | 0.0020809 | 0.0011666 | m | 0.00059384 | 0.0010421 | 0.0019851 | 0.0001108 | m | OK | 0.003 | 5 | Hz | 0.001 | 0.001 | 0.001 | 2.0086e-06 | 9.5793e-08 | 1.1719e-06 | 0.001 | 2.0086e-06 | 9.5792e-05 | 0.0011719 | 5Hz1m | 27 | 0.0010001 | 1e-08 | 1.958e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00019579 | 4.9997e-05 | 5Hz3m1m | 5.9525e-08 | 0.0034342 | 1.0001e-07 | 0.0023333 | 5Hz3m | 5Hz3m | 45 | 1.0001 | 4.0172e-06 | 9.9996e-05 | 0.0023443 | OK
3 | m | 10 | 1 | m | 1.0003 | 0.0012118 | 0.00050416 | 0.00081995 | m | 0.0001644 | 0.0014535 | 0.00045991 | 0.001391 | m | OK | 0.003 | 10 | Hz | 0.001 | 0.001 | 0.0010001 | 1.8924e-06 | 4.4255e-08 | 1.6147e-06 | 0.0010001 | 1.8924e-06 | 4.4248e-05 | 0.0016144 | 10Hz1m | 28 | 0.0010002 | 1e-08 | 1.4427e-07 | 5e-08 | 0.0010002 | 1e-08 | 0.00014425 | 4.9991e-05 | 10Hz3m1m | 3.7737e-08 | 0.0024236 | 1.0002e-07 | 0.0016399 | 10Hz3m | 10Hz3m | 46 | 1 | 3.7848e-06 | 9.9999e-05 | 0.0032292 | OK
3 | m | 20 | 1 | m | 0.9997 | 9.5401e-05 | -0.00073172 | 0.0013347 | m | -0.00044759 | 4.2711e-05 | -0.00078486 | 0.0012791 | m | OK | 0.003 | 20 | Hz | 0.001 | 0.001 | 0.0010001 | 1.0452e-07 | 5.3134e-08 | 1.8486e-06 | 0.0010001 | 1.0452e-07 | 5.3126e-05 | 0.0018484 | 20Hz1m | 29 | 0.0010001 | 1e-08 | 1.5313e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00015312 | 4.9997e-05 | 20Hz3m1m | -8.6689e-08 | 0.0001908 | 9.9995e-08 | 0.0026694 | 20Hz3m | 20Hz3m | 47 | 0.99991 | 2.0905e-07 | 9.9994e-05 | 0.0036971 | OK
3 | m | 50 | 1 | m | 0.99876 | 8.3996e-05 | 0.00084005 | 0.00094205 | m | -0.0013854 | 0.0010434 | 0.00059007 | 0.0018173 | m | OK | 0.003 | 50 | Hz | 0.001 | 0.001 | 0.0010001 | 1.0468e-06 | 2.4998e-07 | 2.047e-06 | 0.0010001 | 1.0468e-06 | 0.00024994 | 0.0020467 | 50Hz1m | 30 | 0.0010001 | 1e-08 | 3.4996e-07 | 5e-08 | 0.0010001 | 1e-08 | 0.00034994 | 4.9996e-05 | 50Hz3m1m | -7.5343e-08 | 0.00016799 | 9.9984e-08 | 0.0018841 | 50Hz3m | 50Hz3m | 48 | 0.99992 | 2.0935e-06 | 9.9996e-05 | 0.0040936 | OK
3 | m | 100 | 1 | m | 1.0001 | 0.00095613 | 0.0015319 | 0.0005222 | m | -0.00018764 | 0.00058735 | 0.00072599 | 0.0018672 | m | OK | 0.003 | 100 | Hz | 0.001 | 0.001 | 0.0010003 | 1.1221e-06 | 8.0587e-07 | 1.9389e-06 | 0.0010003 | 1.1221e-06 | 0.00080566 | 0.0019383 | 100Hz1m | 31 | 0.0010003 | 1e-08 | 9.0589e-07 | 5e-08 | 0.0010003 | 1e-08 | 0.00090565 | 4.9987e-05 | 100Hz3m1m | -5.0985e-09 | 0.0019123 | 1.0002e-07 | 0.0010444 | 100Hz3m | 100Hz3m | 49 | 0.99999 | 2.2442e-06 | 9.9993e-05 | 0.003877 | OK
3 | m | 200 | 1 | m | 1.0005 | 0.001279 | 0.00094895 | 0.00032128 | m | -0.00027026 | 0.0018455 | -0.00071907 | 0.00062235 | m | OK | 0.003 | 200 | Hz | 0.001 | 0.001 | 0.0010007 | 2.2453e-06 | 1.668e-06 | 7.0039e-07 | 0.0010007 | 2.2453e-06 | 0.0016668 | 0.00069988 | 200Hz1m | 32 | 0.0010007 | 1e-08 | 1.768e-06 | 5e-08 | 0.0010007 | 1e-08 | 0.0017668 | 4.9965e-05 | 200Hz3m1m | -4.3583e-08 | 0.002558 | 9.9996e-08 | 0.00064257 | 200Hz3m | 200Hz3m | 50 | 0.99996 | 4.4907e-06 | 9.9999e-05 | 0.0014006 | OK
3 | m | 500 | 1 | m | 1.0011 | 0.00085385 | 0.0050408 | 0.001462 | m | -0.0016742 | 0.0017553 | 0.00053156 | 0.00071488 | m | OK | 0.003 | 500 | Hz | 0.001 | 0.001 | 0.0010028 | 1.952e-06 | 4.5093e-06 | 1.6274e-06 | 0.0010028 | 1.952e-06 | 0.0044966 | 0.0016228 | 500Hz1m | 33 | 0.0010029 | 1.5811e-08 | 4.6098e-06 | 5e-08 | 0.0010029 | 1.5813e-08 | 0.0045964 | 4.9854e-05 | 500Hz3m1m | 9.9295e-08 | 0.0017077 | 1.0053e-07 | 0.0029239 | 500Hz3m | 500Hz3m | 51 | 1.0001 | 3.904e-06 | 9.9792e-05 | 0.003246 | OK
3 | m | 1000 | 1 | m | 1.0086 | 0.00074392 | 0.010465 | 0.0016703 | m | 0.00030048 | 0.0013977 | 0.0013276 | 1.7739e-05 | m | OK | 0.003 | 1 | kHz | 0.001 | 0.001 | 0.0010083 | 1.5834e-06 | 9.1374e-06 | 1.6704e-06 | 0.0010083 | 1.5834e-06 | 0.0090622 | 0.0016566 | 1kHz1m | 34 | 0.0010083 | 4.4721e-08 | 9.2383e-06 | 1e-07 | 0.0010084 | 4.4729e-08 | 0.0091616 | 9.9165e-05 | 1kHz3m1m | 6.383e-08 | 0.0014878 | 1.0087e-07 | 0.0033407 | 1kHz3m | 1kHz3m | 52 | 1.0001 | 3.1667e-06 | 9.9451e-05 | 0.0033148 | OK
3 | m | 2000 | 1 | m | 1.0231 | 0.0011105 | 0.016925 | 0.001103 | m | -0.00040006 | 0.00035619 | -0.0013539 | 0.00091666 | m | OK | 0.003 | 2 | kHz | 0.001 | 0.001 | 0.0010235 | 1.1663e-06 | 1.8279e-05 | 1.4342e-06 | 0.0010236 | 1.1664e-06 | 0.017857 | 0.001401 | 2kHz1m | 35 | 0.0010236 | 1.2649e-07 | 1.8381e-05 | 2e-07 | 0.0010237 | 1.2652e-07 | 0.017956 | 0.00019534 | 2kHz3m1m | 8.2417e-08 | 0.0022211 | 1.0261e-07 | 0.0022059 | 2kHz3m | 2kHz3m | 53 | 1.0001 | 2.3327e-06 | 9.8773e-05 | 0.0028087 | OK
3 | m | 5000 | 1 | m | 1.0947 | 0.00071684 | 0.045025 | 0.00018829 | m | 0.0018331 | 0.0014619 | -0.00057864 | 0.00056195 | m | OK | 0.003 | 5 | kHz | 0.001 | 0.001 | 0.0010928 | 1.6282e-06 | 4.5604e-05 | 5.9266e-07 | 0.0010938 | 1.627e-06 | 0.041705 | 0.00054491 | 5kHz1m | 36 | 0.0010929 | 5e-07 | 4.5708e-05 | 5e-07 | 0.0010939 | 5e-07 | 0.041798 | 0.00045709 | 5kHz3m1m | 6.676e-08 | 0.0014337 | 1.0383e-07 | 0.00037658 | 5kHz3m | 5kHz3m | 54 | 1.0001 | 3.2539e-06 | 9.229e-05 | 0.0011818 | OK
3 | m | 0.01 | 3 | m | 2.9993 | 0.00097866 | 0.00013521 | 0.00076701 | m | -0.00035348 | 0.00054577 | 0.00032957 | 0.00044318 | m | OK | 0.003 | 10 | mHz | 0.003 | 0.001 | 0.0029996 | 1.1206e-06 | -1.9436e-07 | 8.8584e-07 | 0.0029996 | 1.1206e-06 | -6.4794e-05 | 0.00029532 | 10mHz3m | 37 | 0.0029997 | 3e-08 | 3.7684e-08 | 1.5e-07 | 0.0029997 | 3e-08 | 1.2563e-05 | 5.0006e-05 | 10mHz3m3m | 3.4393e-08 | 0.0019573 | 2.3204e-07 | 0.001534 | 10mHz3m | 10mHz3m | 37 | 1 | 2.2411e-06 | 7.7357e-05 | 0.00059275 | OK
3 | m | 0.02 | 3 | m | 2.999 | 0.0018909 | -0.0015722 | 0.00083795 | m | -0.00070608 | 0.0011661 | -0.0018114 | 0.00065948 | m | OK | 0.003 | 20 | mHz | 0.003 | 0.001 | 0.0029997 | 2.2215e-06 | 2.3918e-07 | 1.0663e-06 | 0.0029997 | 2.2215e-06 | 7.9735e-05 | 0.00035548 | 20mHz3m | 38 | 0.0029998 | 3e-08 | -2.539e-08 | 1.5e-07 | 0.0029998 | 3e-08 | -8.4637e-06 | 5.0003e-05 | 20mHz3m3m | 1.277e-07 | 0.0037817 | -2.6457e-07 | 0.0016759 | 20mHz3m | 20mHz3m | 38 | 1 | 4.443e-06 | -8.8199e-05 | 0.00071272 | OK
3 | m | 0.05 | 3 | m | 2.9989 | 0.0013597 | -0.00097077 | 0.001547 | m | -0.0014982 | 0.00033408 | -0.0010882 | 0.00018616 | m | OK | 0.003 | 50 | mHz | 0.003 | 0.001 | 0.0030004 | 1.4001e-06 | 1.1741e-07 | 1.5582e-06 | 0.0030004 | 1.4001e-06 | 3.9131e-05 | 0.00051932 | 50mHz3m | 39 | 0.0030001 | 3e-08 | 3.2533e-07 | 1.5e-07 | 0.0030001 | 3e-08 | 0.00010844 | 4.9998e-05 | 50mHz3m3m | -2.9377e-07 | 0.0027194 | 2.0792e-07 | 0.0030941 | 50mHz3m | 50mHz3m | 39 | 0.9999 | 2.8003e-06 | 6.9306e-05 | 0.0010399 | OK
3 | m | 0.1 | 3 | m | 2.9976 | 0.001584 | 0.0014368 | 0.0001226 | m | -0.0018875 | 0.0016581 | 0.00069602 | 0.0019152 | m | OK | 0.003 | 100 | mHz | 0.003 | 0.001 | 0.0029995 | 2.2931e-06 | 7.4078e-07 | 1.9192e-06 | 0.0029995 | 2.2931e-06 | 0.00024697 | 0.00063983 | 100mHz3m | 40 | 0.0029996 | 3e-08 | 8.0053e-07 | 1.5e-07 | 0.0029996 | 3e-08 | 0.00026688 | 5.0007e-05 | 100mHz3m3m | 9.0138e-08 | 0.0031681 | 5.9749e-08 | 0.00024521 | 100mHz3m | 100mHz3m | 40 | 1 | 4.5862e-06 | 1.9912e-05 | 0.0012806 | OK
3 | m | 0.2 | 3 | m | 3.0014 | 0.00059409 | -0.00094119 | 0.0013423 | m | 0.0013635 | 0.00073519 | -0.00092446 | 0.00029895 | m | OK | 0.003 | 200 | mHz | 0.003 | 0.001 | 0.003 | 9.4522e-07 | -1.6733e-08 | 1.3751e-06 | 0.003 | 9.4522e-07 | -5.5777e-06 | 0.00045838 | 200mHz3m | 41 | 0.003 | 3e-08 | -4.2224e-09 | 1.5e-07 | 0.003 | 3e-08 | -1.4075e-06 | 5.0001e-05 | 200mHz3m3m | -6.8559e-08 | 0.0011882 | 1.2511e-08 | 0.0026845 | 200mHz3m | 200mHz3m | 41 | 0.99998 | 1.8904e-06 | 4.1703e-06 | 0.00091812 | OK
3 | m | 0.5 | 3 | m | 2.9987 | 0.00080152 | 0.0019763 | 0.00075542 | m | -0.0011902 | 0.0017871 | 0.0018248 | 0.0017762 | m | OK | 0.003 | 500 | mHz | 0.003 | 0.001 | 0.0029999 | 1.9586e-06 | 1.515e-07 | 1.9302e-06 | 0.0029999 | 1.9586e-06 | 5.0501e-05 | 0.00064342 | 500mHz3m | 42 | 0.0029998 | 3e-08 | 2.075e-07 | 1.5e-07 | 0.0029998 | 3e-08 | 6.9171e-05 | 5.0003e-05 | 500mHz3m3m | -2.2045e-08 | 0.001603 | 5.6006e-08 | 0.0015108 | 500mHz3m | 500mHz3m | 42 | 0.99999 | 3.9172e-06 | 1.867e-05 | 0.0012878 | OK
3 | m | 1 | 3 | m | 3.0004 | 0.00090892 | -0.00056008 | 0.0010625 | m | 5.2763e-05 | 0.00040451 | -0.0018702 | 0.00016225 | m | OK | 0.003 | 1 | Hz | 0.003 | 0.001 | 0.0030003 | 9.9487e-07 | 1.3101e-06 | 1.0748e-06 | 0.0030003 | 9.9487e-07 | 0.00043666 | 0.00035824 | 1Hz3m | 43 | 0.003 | 3e-08 | 1.2324e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00041081 | 5e-05 | 1Hz3m3m | -2.9468e-07 | 0.0018178 | -7.7678e-08 | 0.002125 | 1Hz3m | 1Hz3m | 43 | 0.9999 | 1.9897e-06 | -2.585e-05 | 0.00071823 | OK
3 | m | 2 | 3 | m | 2.998 | 0.00059551 | -0.00024027 | 0.00098729 | m | -0.0019786 | 0.00084979 | 6.5669e-05 | 0.0016385 | m | OK | 0.003 | 2 | Hz | 0.003 | 0.001 | 0.003 | 1.0377e-06 | -3.0594e-07 | 1.913e-06 | 0.003 | 1.0377e-06 | -0.00010198 | 0.00063767 | 2Hz3m | 44 | 0.003 | 3e-08 | -1.4799e-08 | 1.5e-07 | 0.003 | 3e-08 | -4.9329e-06 | 5e-05 | 2Hz3m3m | 3.3348e-08 | 0.001191 | 2.9114e-07 | 0.0019746 | 2Hz3m | 2Hz3m | 44 | 1 | 2.0754e-06 | 9.7049e-05 | 0.0012763 | OK
3 | m | 5 | 3 | m | 2.9996 | 0.0014965 | -0.0012448 | 0.00072225 | m | -0.00022746 | 0.0011301 | -0.0018157 | 0.0013372 | m | OK | 0.003 | 5 | Hz | 0.003 | 0.001 | 0.0029998 | 1.8753e-06 | 5.7089e-07 | 1.5198e-06 | 0.0029998 | 1.8753e-06 | 0.00019031 | 0.00050662 | 5Hz3m | 45 | 0.003 | 3e-08 | 5.0723e-07 | 1.5e-07 | 0.003 | 3e-08 | 0.00016908 | 5e-05 | 5Hz3m3m | 1.6146e-07 | 0.0029931 | -6.3659e-08 | 0.0014445 | 5Hz3m | 5Hz3m | 45 | 1.0001 | 3.7506e-06 | -2.123e-05 | 0.0010145 | OK
3 | m | 10 | 3 | m | 2.998 | 0.0012592 | -3.1025e-05 | 0.001832 | m | -0.0019881 | 0.0011142 | 0.0001386 | 0.00029307 | m | OK | 0.003 | 10 | Hz | 0.003 | 0.001 | 0.003 | 1.6813e-06 | -1.6963e-07 | 1.8553e-06 | 0.003 | 1.6813e-06 | -5.6541e-05 | 0.00061844 | 10Hz3m | 46 | 0.0030001 | 3e-08 | 9.6095e-08 | 1.5e-07 | 0.0030001 | 3e-08 | 3.2031e-05 | 4.9998e-05 | 10Hz3m3m | 7.8446e-08 | 0.0025183 | 2.6572e-07 | 0.0036641 | 10Hz3m | 10Hz3m | 46 | 1 | 3.3627e-06 | 8.8572e-05 | 0.0012379 | OK
3 | m | 20 | 3 | m | 2.9994 | 0.0017183 | 0.0021913 | 0.00123 | m | -0.00062828 | 0.00055778 | 0.00078235 | 0.0013731 | m | OK | 0.003 | 20 | Hz | 0.003 | 0.001 | 0.003 | 1.8066e-06 | 1.409e-06 | 1.8435e-06 | 0.003 | 1.8066e-06 | 0.00046965 | 0.00061448 | 20Hz3m | 47 | 0.003 | 3e-08 | 1.1185e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00037285 | 5e-05 | 20Hz3m3m | -4.3006e-08 | 0.0034367 | -2.9044e-07 | 0.0024599 | 20Hz3m | 20Hz3m | 47 | 0.99999 | 3.6132e-06 | -9.6807e-05 | 0.00123 | OK
3 | m | 50 | 3 | m | 3.0014 | 0.0019967 | 0.0013141 | 0.0012818 | m | 0.0011793 | 0.00023843 | -0.0006652 | 0.00057917 | m | OK | 0.003 | 50 | Hz | 0.003 | 0.001 | 0.0030002 | 2.0109e-06 | 1.9793e-06 | 1.4066e-06 | 0.0030002 | 2.0109e-06 | 0.00065972 | 0.00046882 | 50Hz3m | 48 | 0.0030003 | 3e-08 | 1.9181e-06 | 1.5e-07 | 0.0030003 | 3e-08 | 0.0006393 | 4.9995e-05 | 50Hz3m3m | 1.1116e-07 | 0.0039934 | -6.1194e-08 | 0.0025636 | 50Hz3m | 50Hz3m | 48 | 1 | 4.0218e-06 | -2.0421e-05 | 0.00093897 | OK
3 | m | 100 | 3 | m | 3.0019 | 0.0018791 | 0.0012778 | 0.0012362 | m | 0.00072655 | 0.0013523 | -0.0012258 | 0.0013235 | m | OK | 0.003 | 100 | Hz | 0.003 | 0.001 | 0.0030012 | 2.3151e-06 | 2.5036e-06 | 1.811e-06 | 0.0030012 | 2.3151e-06 | 0.00083422 | 0.00060344 | 100Hz3m | 49 | 0.0030012 | 3e-08 | 2.6469e-06 | 1.5e-07 | 0.0030012 | 3e-08 | 0.00088195 | 4.9979e-05 | 100Hz3m3m | 7.2517e-08 | 0.0037581 | 1.4332e-07 | 0.0024724 | 100Hz3m | 100Hz3m | 49 | 1 | 4.6302e-06 | 4.7732e-05 | 0.0012079 | OK
3 | m | 200 | 3 | m | 3.0003 | 0.00033839 | 0.0068899 | 3.5785e-05 | m | -0.0018856 | 0.0012185 | 0.001601 | 0.00095516 | m | OK | 0.003 | 200 | Hz | 0.003 | 0.001 | 0.0030022 | 1.2646e-06 | 5.2888e-06 | 9.5583e-07 | 0.0030022 | 1.2646e-06 | 0.0017617 | 0.00031838 | 200Hz3m | 50 | 0.0030021 | 3e-08 | 5.5775e-06 | 1.5e-07 | 0.0030021 | 3.0001e-08 | 0.0018579 | 4.9965e-05 | 200Hz3m3m | -9.4686e-08 | 0.00067678 | 2.887e-07 | 7.157e-05 | 200Hz3m | 200Hz3m | 50 | 0.99997 | 2.5292e-06 | 9.6222e-05 | 0.00063872 | OK
3 | m | 500 | 3 | m | 3.0084 | 0.00033906 | 0.015943 | 0.0012805 | m | -0.00041907 | 0.001126 | 0.0019208 | 0.0013973 | m | OK | 0.003 | 500 | Hz | 0.003 | 0.001 | 0.0030089 | 1.1759e-06 | 1.4023e-05 | 1.8953e-06 | 0.0030089 | 1.1759e-06 | 0.0046604 | 0.00062989 | 500Hz3m | 51 | 0.0030089 | 4.7434e-08 | 1.3893e-05 | 1.5e-07 | 0.003009 | 4.7439e-08 | 0.0046171 | 4.9851e-05 | 500Hz3m3m | 6.0754e-08 | 0.00067813 | -1.3013e-07 | 0.002561 | 500Hz3m | 500Hz3m | 51 | 1 | 2.3519e-06 | -4.3343e-05 | 0.0012608 | OK
3 | m | 1000 | 3 | m | 3.024 | 0.0017707 | 0.028065 | 0.00015509 | m | -0.00055923 | 0.00036949 | 0.00040628 | 0.0015844 | m | OK | 0.003 | 1 | kHz | 0.003 | 0.001 | 0.0030245 | 1.8089e-06 | 2.7659e-05 | 1.5919e-06 | 0.0030247 | 1.8089e-06 | 0.0091446 | 0.00052633 | 1kHz3m | 52 | 0.0030244 | 1.3416e-07 | 2.756e-05 | 3e-07 | 0.0030246 | 1.3419e-07 | 0.0091121 | 9.9184e-05 | 1kHz3m3m | -1.0147e-07 | 0.0035415 | -9.9348e-08 | 0.00031018 | 1kHz3m | 1kHz3m | 52 | 0.99997 | 3.6177e-06 | -3.2539e-05 | 0.0010573 | OK
3 | m | 2000 | 3 | m | 3.0713 | 0.00078744 | 0.056174 | 0.0004344 | m | 0.00091213 | 0.0012918 | 0.0011741 | 0.0017096 | m | OK | 0.003 | 2 | kHz | 0.003 | 0.001 | 0.0030704 | 1.5129e-06 | 5.5e-05 | 1.7639e-06 | 0.0030709 | 1.513e-06 | 0.017911 | 0.00057438 | 2kHz3m | 53 | 0.0030706 | 3.7947e-07 | 5.4982e-05 | 6e-07 | 0.0030711 | 3.7956e-07 | 0.017904 | 0.00019535 | 2kHz3m3m | 1.8611e-07 | 0.0015749 | -1.8404e-08 | 0.00086881 | 2kHz3m | 2kHz3m | 53 | 1.0001 | 3.026e-06 | -7.077e-06 | 0.0011652 | OK
3 | m | 5000 | 3 | m | 3.2776 | 0.00014027 | 0.13635 | 0.00045131 | m | -0.0013945 | 0.00047661 | -0.001426 | 0.0015077 | m | OK | 0.003 | 5 | kHz | 0.003 | 0.001 | 0.003279 | 4.9682e-07 | 0.00013777 | 1.5738e-06 | 0.0032819 | 5.0076e-07 | 0.041992 | 0.00047915 | 5kHz3m | 54 | 0.0032786 | 1.5e-06 | 0.00013747 | 1.5e-06 | 0.0032815 | 1.5e-06 | 0.041905 | 0.00045711 | 5kHz3m3m | -3.4814e-07 | 0.00028055 | -3.0166e-07 | 0.00090262 | 5kHz3m | 5kHz3m | 54 | 0.99989 | 1.0015e-06 | -8.7391e-05 | 0.0010617 | OK
10 | m | 0.01 | 3 | m | 2.9999 | 0.00087554 | -0.0014623 | 0.00143 | m | 7.9156e-05 | 0.0019334 | -0.0012 | 0.00077347 | m | OK | 0.01 | 10 | mHz | 0.003 | 0.001 | 0.0029998 | 2.1224e-06 | -2.623e-07 | 1.6258e-06 | 0.0029998 | 2.1224e-06 | -8.7438e-05 | 0.00054196 | 10mHz3m | 37 | 0.0029997 | 3e-08 | 3.7684e-08 | 1.5e-07 | 0.0029997 | 3e-08 | 1.2563e-05 | 5.0006e-05 | 10mHz10m3m | -1.2909e-07 | 0.0017511 | 2.9998e-07 | 0.00286 | 10mHz10m | 10mHz10m | 73 | 0.99996 | 4.2448e-06 | 0.0001 | 0.0010851 | OK
10 | m | 0.02 | 3 | m | 2.9986 | 4.6807e-05 | -0.00048198 | 0.00097043 | m | -0.0014555 | 0.0014837 | -0.00015658 | 0.00061195 | m | OK | 0.01 | 20 | mHz | 0.003 | 0.001 | 0.0030001 | 1.4844e-06 | -3.254e-07 | 1.1473e-06 | 0.0030001 | 1.4844e-06 | -0.00010846 | 0.00038241 | 20mHz3m | 38 | 0.0029998 | 3e-08 | -2.539e-08 | 1.5e-07 | 0.0029998 | 3e-08 | -8.4637e-06 | 5.0003e-05 | 20mHz10m3m | -2.9375e-07 | 9.3614e-05 | 3.0001e-07 | 0.0019409 | 20mHz10m | 20mHz10m | 74 | 0.9999 | 2.9689e-06 | 0.0001 | 0.00076645 | OK
10 | m | 0.05 | 3 | m | 3.0005 | 0.00010005 | 0.00029447 | 0.0016453 | m | 0.00015926 | 0.0011004 | 0.00026915 | 0.000674 | m | OK | 0.01 | 50 | mHz | 0.003 | 0.001 | 0.0030003 | 1.105e-06 | 2.5317e-08 | 1.778e-06 | 0.0030003 | 1.105e-06 | 8.4382e-06 | 0.00059262 | 50mHz3m | 39 | 0.0030001 | 3e-08 | 3.2533e-07 | 1.5e-07 | 0.0030001 | 3e-08 | 0.00010844 | 4.9998e-05 | 50mHz10m3m | -1.8406e-07 | 0.0002001 | 3.0001e-07 | 0.0032907 | 50mHz10m | 50mHz10m | 75 | 0.99994 | 2.21e-06 | 9.9999e-05 | 0.0011863 | OK
10 | m | 0.1 | 3 | m | 2.9976 | 7.7125e-05 | -0.00074129 | 0.0017433 | m | -0.0018135 | 0.00081317 | -0.0012418 | 0.00037308 | m | OK | 0.01 | 100 | mHz | 0.003 | 0.001 | 0.0029995 | 8.1682e-07 | 5.0055e-07 | 1.7827e-06 | 0.0029995 | 8.1682e-07 | 0.00016688 | 0.00059435 | 100mHz3m | 40 | 0.0029996 | 3e-08 | 8.0053e-07 | 1.5e-07 | 0.0029996 | 3e-08 | 0.00026688 | 5.0007e-05 | 100mHz10m3m | 1.1845e-07 | 0.00015425 | 2.9999e-07 | 0.0034865 | 100mHz10m | 100mHz10m | 76 | 1 | 1.6336e-06 | 0.0001 | 0.0011898 | OK
10 | m | 0.2 | 3 | m | 2.9991 | 0.001359 | 7.6599e-05 | 0.00087492 | m | -0.00081413 | 0.00061402 | 0.00038081 | 0.0013378 | m | OK | 0.01 | 200 | mHz | 0.003 | 0.001 | 0.0029999 | 1.4913e-06 | -3.0422e-07 | 1.5985e-06 | 0.0029999 | 1.4913e-06 | -0.00010141 | 0.00053285 | 200mHz3m | 41 | 0.003 | 3e-08 | -4.2224e-09 | 1.5e-07 | 0.003 | 3e-08 | -1.4075e-06 | 5.0001e-05 | 200mHz10m3m | 3.6511e-08 | 0.0027181 | 2.9999e-07 | 0.0017498 | 200mHz10m | 200mHz10m | 77 | 1 | 2.9826e-06 | 0.0001 | 0.0010669 | OK
10 | m | 0.5 | 3 | m | 3.0012 | 0.00088055 | -0.0011393 | 0.00013322 | m | 0.0013796 | 0.0016315 | -0.0010468 | 0.0015907 | m | OK | 0.01 | 500 | mHz | 0.003 | 0.001 | 0.0029998 | 1.8539e-06 | -9.2488e-08 | 1.5963e-06 | 0.0029998 | 1.8539e-06 | -3.0831e-05 | 0.00053213 | 500mHz3m | 42 | 0.0029998 | 3e-08 | 2.075e-07 | 1.5e-07 | 0.0029998 | 3e-08 | 6.9171e-05 | 5.0003e-05 | 500mHz10m3m | 2.8682e-08 | 0.0017611 | 2.9999e-07 | 0.00026644 | 500mHz10m | 500mHz10m | 78 | 1 | 3.7079e-06 | 0.0001 | 0.0010654 | OK
10 | m | 1 | 3 | m | 2.9987 | 0.0015136 | 0.00091698 | 0.00032096 | m | -0.0012383 | 0.0003398 | -1.5425e-05 | 0.00080838 | m | OK | 0.01 | 1 | Hz | 0.003 | 0.001 | 0.003 | 1.5513e-06 | 9.3241e-07 | 8.6977e-07 | 0.003 | 1.5513e-06 | 0.00031081 | 0.00028993 | 1Hz3m | 43 | 0.003 | 3e-08 | 1.2324e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00041081 | 5e-05 | 1Hz10m3m | 4.3805e-08 | 0.0030273 | 3.0001e-07 | 0.00064192 | 1Hz10m | 1Hz10m | 79 | 1 | 3.1026e-06 | 0.0001 | 0.000582 | OK
10 | m | 2 | 3 | m | 3.0011 | 0.0012006 | -0.0016096 | 0.0011408 | m | 0.001223 | 0.00011801 | -0.0012949 | 0.00040927 | m | OK | 0.01 | 2 | Hz | 0.003 | 0.001 | 0.0029999 | 1.2064e-06 | -3.1479e-07 | 1.212e-06 | 0.0029999 | 1.2064e-06 | -0.00010493 | 0.000404 | 2Hz3m | 44 | 0.003 | 3e-08 | -1.4799e-08 | 1.5e-07 | 0.003 | 3e-08 | -4.9329e-06 | 5e-05 | 2Hz10m3m | 1.0505e-07 | 0.0024012 | 2.9999e-07 | 0.0022815 | 2Hz10m | 2Hz10m | 80 | 1 | 2.4128e-06 | 0.0001 | 0.00080955 | OK
10 | m | 5 | 3 | m | 2.9998 | 0.001333 | -1.7755e-05 | 0.00056745 | m | -0.00015269 | 6.7339e-06 | -0.00022498 | 0.00022231 | m | OK | 0.01 | 5 | Hz | 0.003 | 0.001 | 0.0029999 | 1.333e-06 | 2.0722e-07 | 6.0945e-07 | 0.0029999 | 1.333e-06 | 6.9076e-05 | 0.00020315 | 5Hz3m | 45 | 0.003 | 3e-08 | 5.0723e-07 | 1.5e-07 | 0.003 | 3e-08 | 0.00016908 | 5e-05 | 5Hz10m3m | 8.5356e-08 | 0.002666 | 3.0001e-07 | 0.0011349 | 5Hz10m | 5Hz10m | 81 | 1 | 2.666e-06 | 0.0001 | 0.00040937 | OK
10 | m | 10 | 3 | m | 2.9986 | 0.00071443 | 0.0001644 | 0.00096004 | m | -0.0015056 | 0.0015287 | 0.00036831 | 0.001759 | m | OK | 0.01 | 10 | Hz | 0.003 | 0.001 | 0.0030001 | 1.6874e-06 | -2.0391e-07 | 2.0039e-06 | 0.0030001 | 1.6874e-06 | -6.7967e-05 | 0.00066794 | 10Hz3m | 46 | 0.0030001 | 3e-08 | 9.6095e-08 | 1.5e-07 | 0.0030001 | 3e-08 | 3.2031e-05 | 4.9998e-05 | 10Hz10m3m | -8.7454e-09 | 0.0014289 | 3e-07 | 0.0019201 | 10Hz10m | 10Hz10m | 82 | 1 | 3.3748e-06 | 9.9997e-05 | 0.0013368 | OK
10 | m | 20 | 3 | m | 3.0016 | 0.00067018 | 0.0011057 | 0.0013105 | m | 0.0017226 | 0.0011436 | 0.00028716 | 0.0017526 | m | OK | 0.01 | 20 | Hz | 0.003 | 0.001 | 0.0029999 | 1.3255e-06 | 8.1851e-07 | 2.1883e-06 | 0.0029999 | 1.3255e-06 | 0.00027285 | 0.00072947 | 20Hz3m | 47 | 0.003 | 3e-08 | 1.1185e-06 | 1.5e-07 | 0.003 | 3e-08 | 0.00037285 | 5e-05 | 20Hz10m3m | 7.3898e-08 | 0.0013404 | 3.0002e-07 | 0.0026209 | 20Hz10m | 20Hz10m | 83 | 1 | 2.651e-06 | 0.0001 | 0.0014598 | OK
10 | m | 50 | 3 | m | 3.0004 | 0.00050493 | 0.0027869 | 0.0017099 | m | -0.00015943 | 0.00031363 | 0.0011688 | 0.0018686 | m | OK | 0.01 | 50 | Hz | 0.003 | 0.001 | 0.0030005 | 5.944e-07 | 1.6182e-06 | 2.5329e-06 | 0.0030005 | 5.9441e-07 | 0.0005393 | 0.00084414 | 50Hz3m | 48 | 0.0030003 | 3e-08 | 1.9181e-06 | 1.5e-07 | 0.0030003 | 3e-08 | 0.0006393 | 4.9995e-05 | 50Hz10m3m | -2.1913e-07 | 0.0010099 | 2.9991e-07 | 0.0034198 | 50Hz10m | 50Hz10m | 84 | 0.99993 | 1.1888e-06 | 9.9999e-05 | 0.001689 | OK
10 | m | 100 | 3 | m | 3.0022 | 0.0011763 | 0.0012898 | 0.00037132 | m | 0.0009528 | 0.0012951 | -0.001057 | 0.0015807 | m | OK | 0.01 | 100 | Hz | 0.003 | 0.001 | 0.0030012 | 1.7496e-06 | 2.3468e-06 | 1.6237e-06 | 0.0030012 | 1.7496e-06 | 0.00078195 | 0.00054103 | 100Hz3m | 49 | 0.0030012 | 3e-08 | 2.6469e-06 | 1.5e-07 | 0.0030012 | 3e-08 | 0.00088195 | 4.9979e-05 | 100Hz10m3m | 2.1915e-08 | 0.0023525 | 3.0013e-07 | 0.00074264 | 100Hz10m | 100Hz10m | 85 | 1 | 3.4991e-06 | 9.9996e-05 | 0.0010832 | OK
10 | m | 200 | 3 | m | 3.002 | 0.00065718 | 0.0069311 | 0.0018928 | m | 3.6897e-05 | 0.00067134 | 0.001654 | 0.00079653 | m | OK | 0.01 | 200 | Hz | 0.003 | 0.001 | 0.003002 | 9.3946e-07 | 5.2771e-06 | 2.0535e-06 | 0.003002 | 9.3946e-07 | 0.0017579 | 0.00068407 | 200Hz3m | 50 | 0.0030021 | 3e-08 | 5.5775e-06 | 1.5e-07 | 0.0030021 | 3.0001e-08 | 0.0018579 | 4.9965e-05 | 200Hz10m3m | 1.2877e-07 | 0.0013144 | 3.0043e-07 | 0.0037855 | 200Hz10m | 200Hz10m | 86 | 1 | 1.8789e-06 | 9.9998e-05 | 0.001369 | OK
10 | m | 500 | 3 | m | 3.0081 | 0.0010074 | 0.012045 | 6.1561e-05 | m | -0.00061741 | 0.0018272 | -0.0015458 | 0.00016201 | m | OK | 0.01 | 500 | Hz | 0.003 | 0.001 | 0.0030087 | 2.0865e-06 | 1.3591e-05 | 1.7331e-07 | 0.0030087 | 2.0865e-06 | 0.0045171 | 5.7687e-05 | 500Hz3m | 51 | 0.0030089 | 4.7434e-08 | 1.3893e-05 | 1.5e-07 | 0.003009 | 4.7439e-08 | 0.0046171 | 4.9851e-05 | 500Hz10m3m | 2.0609e-07 | 0.0020148 | 3.0181e-07 | 0.00012312 | 500Hz10m | 500Hz10m | 87 | 1.0001 | 4.173e-06 | 9.9995e-05 | 0.00012568 | OK
10 | m | 1000 | 3 | m | 3.024 | 0.0017254 | 0.028935 | 0.0007181 | m | -0.000724 | 0.00098926 | 0.0016731 | 0.00060242 | m | OK | 0.01 | 1 | kHz | 0.003 | 0.001 | 0.0030247 | 1.9889e-06 | 2.7262e-05 | 9.3732e-07 | 0.0030249 | 1.9888e-06 | 0.0090126 | 0.00030992 | 1kHz3m | 52 | 0.0030244 | 1.3416e-07 | 2.756e-05 | 3e-07 | 0.0030246 | 1.3419e-07 | 0.0091121 | 9.9184e-05 | 1kHz10m3m | -2.908e-07 | 0.0034509 | 2.9819e-07 | 0.0014362 | 1kHz10m | 1kHz10m | 88 | 0.9999 | 3.9777e-06 | 9.9453e-05 | 0.00062772 | OK
10 | m | 2000 | 3 | m | 3.0695 | 0.0010129 | 0.056457 | 0.00087338 | m | -0.0010149 | 0.00085902 | 0.0017827 | 0.0015923 | m | OK | 0.01 | 2 | kHz | 0.003 | 0.001 | 0.0030705 | 1.3281e-06 | 5.4674e-05 | 1.8161e-06 | 0.003071 | 1.3283e-06 | 0.017804 | 0.00059132 | 2kHz3m | 53 | 0.0030706 | 3.7947e-07 | 5.4982e-05 | 6e-07 | 0.0030711 | 3.7956e-07 | 0.017904 | 0.00019535 | 2kHz10m3m | 7.5924e-08 | 0.0020259 | 3.0754e-07 | 0.0017468 | 2kHz10m | 2kHz10m | 89 | 1 | 2.6566e-06 | 9.9684e-05 | 0.0011987 | OK
10 | m | 5000 | 3 | m | 3.2778 | 0.0017408 | 0.13839 | 0.0010205 | m | -0.00078823 | 0.00042269 | 0.0012341 | 0.00060528 | m | OK | 0.01 | 5 | kHz | 0.003 | 0.001 | 0.0032786 | 1.7914e-06 | 0.00013715 | 1.1865e-06 | 0.0032815 | 1.7905e-06 | 0.041808 | 0.00036196 | 5kHz3m | 54 | 0.0032786 | 1.5e-06 | 0.00013747 | 1.5e-06 | 0.0032815 | 1.5e-06 | 0.041905 | 0.00045711 | 5kHz10m3m | -4.7338e-09 | 0.0034816 | 3.1839e-07 | 0.0020409 | 5kHz10m | 5kHz10m | 90 | 1 | 3.581e-06 | 9.7001e-05 | 0.00085617 | OK
10 | m | 0.01 | 10 | m | 10.0014 | 0.0014812 | -0.0013835 | 0.00014192 | m | 0.0012329 | 0.00084688 | 0.00062425 | 0.00087442 | m | OK | 0.01 | 10 | mHz | 0.01 | 0.001 | 0.01 | 1.7062e-06 | -2.0077e-06 | 8.8586e-07 | 0.01 | 1.7062e-06 | -0.00020077 | 8.8584e-05 | 10mHz10m | 55 | 0.010001 | 1e-07 | -1.2341e-06 | 5e-07 | 0.010001 | 1e-07 | -0.0001234 | 4.9994e-05 | 10mHz10m10m | 9.0451e-07 | 0.0029623 | 7.7359e-07 | 0.00028384 | 10mHz10m | 10mHz10m | 73 | 1.0001 | 3.4123e-06 | 7.7369e-05 | 0.00018409 | OK
10 | m | 0.02 | 10 | m | 9.9977 | 0.0011215 | -0.0018829 | 0.00063096 | m | -0.0019064 | 0.001136 | 0.00059866 | 0.0019515 | m | OK | 0.01 | 20 | mHz | 0.01 | 0.001 | 0.0099996 | 1.5963e-06 | -2.4816e-06 | 2.051e-06 | 0.0099996 | 1.5963e-06 | -0.00024817 | 0.00020511 | 20mHz10m | 56 | 0.01 | 1e-07 | -1.6219e-06 | 5e-07 | 0.01 | 1e-07 | -0.00016219 | 4.9999e-05 | 20mHz10m10m | 6.9304e-07 | 0.0022429 | 8.5965e-07 | 0.0012619 | 20mHz10m | 20mHz10m | 74 | 1.0001 | 3.1926e-06 | 8.5979e-05 | 0.00041325 | OK
10 | m | 0.05 | 10 | m | 9.9991 | 0.0002807 | 0.00056723 | 0.00072325 | m | -0.0019769 | 0.0011822 | 0.0008837 | 0.00018995 | m | OK | 0.01 | 50 | mHz | 0.01 | 0.001 | 0.010001 | 1.215e-06 | -3.1647e-07 | 7.4778e-07 | 0.010001 | 1.215e-06 | -3.1643e-05 | 7.477e-05 | 50mHz10m | 57 | 0.01 | 1e-07 | 7.5712e-08 | 5e-07 | 0.01 | 1e-07 | 7.5709e-06 | 4.9998e-05 | 50mHz10m10m | -6.7861e-07 | 0.00056139 | 3.9218e-07 | 0.0014465 | 50mHz10m | 50mHz10m | 75 | 0.99993 | 2.4301e-06 | 3.9214e-05 | 0.00015768 | OK
10 | m | 0.1 | 10 | m | 9.9988 | 9.3467e-05 | 1.262e-05 | 0.00065416 | m | -0.0014881 | 0.0016431 | -9.5481e-06 | 0.00036807 | m | OK | 0.01 | 100 | mHz | 0.01 | 0.001 | 0.01 | 1.6458e-06 | 2.2168e-08 | 7.506e-07 | 0.01 | 1.6458e-06 | 2.2167e-06 | 7.5058e-05 | 100mHz10m | 58 | 0.01 | 1e-07 | 4.3356e-07 | 5e-07 | 0.01 | 1e-07 | 4.3356e-05 | 5e-05 | 100mHz10m10m | -2.0872e-07 | 0.00018693 | 4.114e-07 | 0.0013083 | 100mHz10m | 100mHz10m | 76 | 0.99998 | 3.2915e-06 | 4.1139e-05 | 0.00015822 | OK
10 | m | 0.2 | 10 | m | 9.9983 | 0.0018327 | -0.00011118 | 0.00021938 | m | -0.00064898 | 0.00040435 | -0.0010155 | 0.0014593 | m | OK | 0.01 | 200 | mHz | 0.01 | 0.001 | 0.009999 | 1.8768e-06 | 9.0435e-07 | 1.4757e-06 | 0.009999 | 1.8768e-06 | 9.0444e-05 | 0.00014758 | 200mHz10m | 59 | 0.0099993 | 1e-07 | 1.1321e-06 | 5e-07 | 0.0099993 | 1e-07 | 0.00011322 | 5.0004e-05 | 200mHz10m10m | 2.7533e-07 | 0.0036655 | 2.278e-07 | 0.00043877 | 200mHz10m | 200mHz10m | 77 | 1 | 3.7536e-06 | 2.2779e-05 | 0.00029937 | OK
10 | m | 0.5 | 10 | m | 9.9989 | 0.00061014 | 0.00038543 | 0.0016791 | m | -0.0018445 | 0.0016178 | -0.00025706 | 0.0013191 | m | OK | 0.01 | 500 | mHz | 0.01 | 0.001 | 0.010001 | 1.7291e-06 | 6.4249e-07 | 2.1353e-06 | 0.010001 | 1.7291e-06 | 6.4244e-05 | 0.00021351 | 500mHz10m | 60 | 0.0099999 | 1e-07 | 5.2604e-07 | 5e-07 | 0.0099999 | 1e-07 | 5.2605e-05 | 5.0001e-05 | 500mHz10m10m | -8.3937e-07 | 0.0012203 | -1.1645e-07 | 0.0033582 | 500mHz10m | 500mHz10m | 78 | 0.99992 | 3.4581e-06 | -1.164e-05 | 0.00042994 | OK
10 | m | 1 | 10 | m | 10.0025 | 0.0019629 | -0.00039958 | 0.0011316 | m | 0.001367 | 0.0015838 | -0.00072857 | 0.0019894 | m | OK | 0.01 | 1 | Hz | 0.01 | 0.001 | 0.010001 | 2.5222e-06 | 3.2899e-07 | 2.2887e-06 | 0.010001 | 2.5222e-06 | 3.2895e-05 | 0.00022885 | 1Hz10m | 61 | 0.01 | 1e-07 | 4.2024e-07 | 5e-07 | 0.01 | 1e-07 | 4.2023e-05 | 4.9999e-05 | 1Hz10m10m | -9.2959e-07 | 0.0039258 | 9.1253e-08 | 0.0022633 | 1Hz10m | 1Hz10m | 79 | 0.99991 | 5.0443e-06 | 9.1282e-06 | 0.00046042 | OK
10 | m | 2 | 10 | m | 10.002 | 0.00017667 | 0.0011082 | 0.0011254 | m | 0.001757 | 0.00033064 | 0.001234 | 0.00039654 | m | OK | 0.01 | 2 | Hz | 0.01 | 0.001 | 0.01 | 3.7488e-07 | -1.2582e-07 | 1.1933e-06 | 0.01 | 3.7488e-07 | -1.2581e-05 | 0.00011932 | 2Hz10m | 62 | 0.0099994 | 1e-07 | -1.1907e-07 | 5e-07 | 0.0099994 | 1e-07 | -1.1907e-05 | 5.0003e-05 | 2Hz10m10m | -8.7805e-07 | 0.00035334 | 6.7498e-09 | 0.0022509 | 2Hz10m | 2Hz10m | 80 | 0.99991 | 7.4977e-07 | 6.7391e-07 | 0.00024383 | OK
10 | m | 5 | 10 | m | 10 | 0.0013804 | 0.00037577 | 0.00020191 | m | 0.0018245 | 6.6644e-05 | -0.0010347 | 0.00057105 | m | OK | 0.01 | 5 | Hz | 0.01 | 0.001 | 0.0099981 | 1.382e-06 | 1.4105e-06 | 6.057e-07 | 0.0099981 | 1.382e-06 | 0.00014108 | 6.0581e-05 | 5Hz10m | 63 | 0.009999 | 1e-07 | 9.0385e-07 | 5e-07 | 0.009999 | 1e-07 | 9.0394e-05 | 5.0005e-05 | 5Hz10m10m | 8.9949e-07 | 0.0027607 | -5.0666e-07 | 0.00040381 | 5Hz10m | 5Hz10m | 81 | 1.0001 | 2.7639e-06 | -5.0683e-05 | 0.00013108 | OK
10 | m | 10 | 10 | m | 10.0017 | 0.00075575 | 0.00076635 | 0.00075379 | m | 0.00079649 | 0.0019018 | 0.00012896 | 0.00083976 | m | OK | 0.01 | 10 | Hz | 0.01 | 0.001 | 0.010001 | 2.0465e-06 | 6.374e-07 | 1.1285e-06 | 0.010001 | 2.0465e-06 | 6.3734e-05 | 0.00011283 | 10Hz10m | 64 | 0.010001 | 1e-07 | -7.0354e-08 | 5e-07 | 0.010001 | 1e-07 | -7.0349e-06 | 4.9997e-05 | 10Hz10m10m | -2.676e-07 | 0.0015115 | -7.0775e-07 | 0.0015076 | 10Hz10m | 10Hz10m | 82 | 0.99997 | 4.093e-06 | -7.0769e-05 | 0.00023114 | OK
10 | m | 20 | 10 | m | 10.0001 | 0.001784 | 0.00094065 | 0.00056184 | m | -0.00094527 | 0.00063595 | -9.4174e-05 | 0.0018186 | m | OK | 0.01 | 20 | Hz | 0.01 | 0.001 | 0.010001 | 1.894e-06 | 1.0348e-06 | 1.9034e-06 | 0.010001 | 1.894e-06 | 0.00010347 | 0.00019032 | 20Hz10m | 65 | 0.010001 | 1e-07 | 1.2883e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00012882 | 4.9997e-05 | 20Hz10m10m | -5.06e-07 | 0.003568 | 2.5343e-07 | 0.0011237 | 20Hz10m | 20Hz10m | 83 | 0.99995 | 3.7879e-06 | 2.5347e-05 | 0.00038391 | OK
10 | m | 50 | 10 | m | 10.0015 | 8.2833e-05 | 0.0012957 | 0.00014792 | m | 0.00026598 | 0.00039833 | -0.0018991 | 0.0014157 | m | OK | 0.01 | 50 | Hz | 0.01 | 0.001 | 0.010001 | 4.0685e-07 | 3.1948e-06 | 1.4234e-06 | 0.010001 | 4.0685e-07 | 0.00031944 | 0.00014233 | 50Hz10m | 66 | 0.010001 | 1e-07 | 3.9368e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00039364 | 4.9994e-05 | 50Hz10m10m | -5.033e-08 | 0.00016567 | 7.4204e-07 | 0.00029584 | 50Hz10m | 50Hz10m | 84 | 0.99999 | 8.1371e-07 | 7.4197e-05 | 0.00028901 | OK
10 | m | 100 | 10 | m | 10.0019 | 0.00059707 | 0.0086577 | 0.00095318 | m | -0.00034319 | 0.00064232 | -0.00014699 | 0.0014576 | m | OK | 0.01 | 100 | Hz | 0.01 | 0.001 | 0.010002 | 8.7696e-07 | 8.8046e-06 | 1.7416e-06 | 0.010002 | 8.7697e-07 | 0.00088027 | 0.00017412 | 100Hz10m | 67 | 0.010003 | 1e-07 | 9.6062e-06 | 5e-07 | 0.010003 | 1e-07 | 0.00096036 | 4.9987e-05 | 100Hz10m10m | 4.7711e-07 | 0.0011941 | 8.0154e-07 | 0.0019064 | 100Hz10m | 100Hz10m | 85 | 1 | 1.7539e-06 | 8.009e-05 | 0.00035182 | OK
10 | m | 200 | 10 | m | 10.0082 | 0.0012537 | 0.016891 | 9.5872e-05 | m | -0.00027241 | 0.0013932 | 0.00030786 | 0.0018276 | m | OK | 0.01 | 200 | Hz | 0.01 | 0.001 | 0.010008 | 1.8742e-06 | 1.6583e-05 | 1.8301e-06 | 0.010008 | 1.8742e-06 | 0.0016569 | 0.00018285 | 200Hz10m | 68 | 0.010008 | 1e-07 | 1.7531e-05 | 5e-07 | 0.010008 | 1e-07 | 0.0017517 | 4.9961e-05 | 200Hz10m10m | -7.3614e-07 | 0.0025075 | 9.4798e-07 | 0.00019174 | 200Hz10m | 200Hz10m | 86 | 0.99993 | 3.7485e-06 | 9.4847e-05 | 0.00036911 | OK
10 | m | 500 | 10 | m | 10.0305 | 0.0018719 | 0.047047 | 0.00079574 | m | -0.00012442 | 0.00018116 | 0.00057598 | 0.0012801 | m | OK | 0.01 | 500 | Hz | 0.01 | 0.001 | 0.010031 | 1.8806e-06 | 4.6471e-05 | 1.5073e-06 | 0.010031 | 1.8806e-06 | 0.0046329 | 0.00015027 | 500Hz10m | 69 | 0.010031 | 1.5811e-07 | 4.7435e-05 | 5e-07 | 0.010031 | 1.5813e-07 | 0.0047286 | 4.9843e-05 | 500Hz10m10m | 6.7942e-07 | 0.0037438 | 9.6332e-07 | 0.0015915 | 500Hz10m | 500Hz10m | 87 | 1.0001 | 3.7612e-06 | 9.5716e-05 | 0.00030464 | OK
10 | m | 1000 | 10 | m | 10.0838 | 0.0011235 | 0.090967 | 0.0015245 | m | 0.0011297 | 0.0002199 | 0.00027678 | 0.0017803 | m | OK | 0.01 | 1 | kHz | 0.01 | 0.001 | 0.010083 | 1.1448e-06 | 9.069e-05 | 2.3439e-06 | 0.010083 | 1.145e-06 | 0.0089945 | 0.00023245 | 1kHz10m | 70 | 0.010083 | 4.4721e-07 | 9.0108e-05 | 1e-06 | 0.010083 | 4.4729e-07 | 0.0089367 | 9.9173e-05 | 1kHz10m10m | 7.4889e-09 | 0.002247 | -5.8276e-07 | 0.003049 | 1kHz10m | 1kHz10m | 88 | 1 | 2.29e-06 | -5.78e-05 | 0.00047536 | OK
10 | m | 2000 | 10 | m | 10.2336 | 0.00087625 | 0.18252 | 0.0017619 | m | -6.3509e-05 | 0.0013111 | -0.00065006 | 0.0012404 | m | OK | 0.01 | 2 | kHz | 0.01 | 0.001 | 0.010234 | 1.577e-06 | 0.00018317 | 2.1548e-06 | 0.010235 | 1.5772e-06 | 0.017896 | 0.00021051 | 2kHz10m | 71 | 0.010234 | 1.2649e-06 | 0.00018387 | 2e-06 | 0.010236 | 1.2652e-06 | 0.017964 | 0.00019538 | 2kHz10m10m | 3.5784e-07 | 0.0017525 | 7.0124e-07 | 0.0035239 | 2kHz10m | 2kHz10m | 89 | 1 | 3.1544e-06 | 6.7873e-05 | 0.00046414 | OK
10 | m | 5000 | 10 | m | 10.9265 | 0.0017478 | 0.45544 | 0.00014318 | m | -0.00045419 | 0.0013901 | -0.00042888 | 0.0018871 | m | OK | 0.01 | 5 | kHz | 0.01 | 0.001 | 0.010927 | 2.2332e-06 | 0.00045587 | 1.8925e-06 | 0.010936 | 2.2327e-06 | 0.041696 | 0.0001731 | 5kHz10m | 72 | 0.010928 | 5e-06 | 0.00045609 | 5e-06 | 0.010938 | 5e-06 | 0.041711 | 0.00045714 | 5kHz10m10m | 1.1528e-06 | 0.0034956 | 2.1712e-07 | 0.0002864 | 5kHz10m | 5kHz10m | 90 | 1.0001 | 4.4657e-06 | 1.544e-05 | 0.00057344 | OK
100 | m | 0.01 | 10 | m | 10.0002 | 0.00017356 | -0.0013338 | 0.0014933 | m | 7.5052e-05 | 0.0013567 | 0.0009002 | 0.0019764 | m | OK | 0.1 | 10 | mHz | 0.01 | 0.001 | 0.01 | 1.3678e-06 | -2.234e-06 | 2.4771e-06 | 0.01 | 1.3678e-06 | -0.0002234 | 0.0002477 | 10mHz10m | 55 | 0.010001 | 1e-07 | -1.2341e-06 | 5e-07 | 0.010001 | 1e-07 | -0.0001234 | 4.9994e-05 | 10mHz100m10m | 9.3919e-07 | 0.00034713 | 9.9989e-07 | 0.0029866 | 10mHz100m | 10mHz100m | 109 | 1.0001 | 2.7356e-06 | 9.9999e-05 | 0.00049792 | OK
100 | m | 0.02 | 10 | m | 9.998 | 0.0019989 | -0.0034916 | 3.4048e-05 | m | -0.0014924 | 0.0013809 | -0.00086981 | 0.0018224 | m | OK | 0.1 | 20 | mHz | 0.01 | 0.001 | 0.0099995 | 2.4295e-06 | -2.6218e-06 | 1.8227e-06 | 0.0099995 | 2.4295e-06 | -0.00026219 | 0.00018228 | 20mHz10m | 56 | 0.01 | 1e-07 | -1.6219e-06 | 5e-07 | 0.01 | 1e-07 | -0.00016219 | 4.9999e-05 | 20mHz100m10m | 7.8487e-07 | 0.0039979 | 9.9982e-07 | 6.8098e-05 | 20mHz100m | 20mHz100m | 110 | 1.0001 | 4.8591e-06 | 0.0001 | 0.00036797 | OK
100 | m | 0.05 | 10 | m | 10.0003 | 0.00052282 | -8.2571e-05 | 0.001306 | m | -0.00075269 | 0.0017627 | 0.00084181 | 0.0009255 | m | OK | 0.1 | 50 | mHz | 0.01 | 0.001 | 0.010001 | 1.8386e-06 | -9.2439e-07 | 1.6007e-06 | 0.010001 | 1.8386e-06 | -9.2429e-05 | 0.00016005 | 50mHz10m | 57 | 0.01 | 1e-07 | 7.5712e-08 | 5e-07 | 0.01 | 1e-07 | 7.5709e-06 | 4.9998e-05 | 50mHz100m10m | -6.5259e-07 | 0.0010456 | 1.0001e-06 | 0.002612 | 50mHz100m | 50mHz100m | 111 | 0.99993 | 3.6773e-06 | 0.0001 | 0.00032398 | OK
100 | m | 0.1 | 10 | m | 9.9981 | 0.0015601 | -0.0020221 | 0.0014151 | m | -0.00097849 | 0.001659 | -0.0014558 | 0.0010595 | m | OK | 0.1 | 100 | mHz | 0.01 | 0.001 | 0.0099991 | 2.2773e-06 | -5.6639e-07 | 1.7678e-06 | 0.0099991 | 2.2773e-06 | -5.6644e-05 | 0.00017679 | 100mHz10m | 58 | 0.01 | 1e-07 | 4.3356e-07 | 5e-07 | 0.01 | 1e-07 | 4.3356e-05 | 5e-05 | 100mHz100m10m | 9.6638e-07 | 0.0031203 | 9.9995e-07 | 0.0028302 | 100mHz100m | 100mHz100m | 112 | 1.0001 | 4.5546e-06 | 0.0001 | 0.00035711 | OK
100 | m | 0.2 | 10 | m | 9.9997 | 0.00022074 | 0.00046974 | 0.0011951 | m | 0.00074405 | 0.0014648 | 0.00033752 | 0.00063463 | m | OK | 0.1 | 200 | mHz | 0.01 | 0.001 | 0.0099989 | 1.4813e-06 | 1.3222e-07 | 1.3532e-06 | 0.0099989 | 1.4813e-06 | 1.3223e-05 | 0.00013533 | 200mHz10m | 59 | 0.0099993 | 1e-07 | 1.1321e-06 | 5e-07 | 0.0099993 | 1e-07 | 0.00011322 | 5.0004e-05 | 200mHz100m10m | 3.1379e-07 | 0.00044148 | 9.9993e-07 | 0.0023903 | 200mHz100m | 200mHz100m | 113 | 1 | 2.9627e-06 | 0.0001 | 0.00027525 | OK
100 | m | 0.5 | 10 | m | 10.0004 | 0.001881 | 0.00080274 | 0.001982 | m | 0.0007926 | 0.00080972 | 0.0012767 | 0.00048887 | m | OK | 0.1 | 500 | mHz | 0.01 | 0.001 | 0.0099997 | 2.0479e-06 | -4.7394e-07 | 2.0414e-06 | 0.0099997 | 2.0479e-06 | -4.7395e-05 | 0.00020415 | 500mHz10m | 60 | 0.0099999 | 1e-07 | 5.2604e-07 | 5e-07 | 0.0099999 | 1e-07 | 5.2605e-05 | 5.0001e-05 | 500mHz100m10m | 2.1434e-07 | 0.0037621 | 9.9998e-07 | 0.0039641 | 500mHz100m | 500mHz100m | 114 | 1 | 4.0958e-06 | 0.0001 | 0.00041135 | OK
100 | m | 1 | 10 | m | 10.0003 | 0.0019794 | -0.0016053 | 0.0010636 | m | -0.0002819 | 9.3376e-05 | -0.0010255 | 8.9828e-05 | m | OK | 0.1 | 1 | Hz | 0.01 | 0.001 | 0.010001 | 1.9816e-06 | -5.798e-07 | 1.0674e-06 | 0.010001 | 1.9816e-06 | -5.7977e-05 | 0.00010673 | 1Hz10m | 61 | 0.01 | 1e-07 | 4.2024e-07 | 5e-07 | 0.01 | 1e-07 | 4.2023e-05 | 4.9999e-05 | 1Hz100m10m | -3.7732e-07 | 0.0039589 | 1e-06 | 0.0021272 | 1Hz100m | 1Hz100m | 115 | 0.99996 | 3.9633e-06 | 0.0001 | 0.00021924 | OK
100 | m | 2 | 10 | m | 10 | 0.00045781 | -0.0030933 | 0.00097205 | m | -0.00028696 | 0.0012742 | -0.0019742 | 0.00063891 | m | OK | 0.1 | 2 | Hz | 0.01 | 0.001 | 0.01 | 1.3539e-06 | -1.1191e-06 | 1.1632e-06 | 0.01 | 1.3539e-06 | -0.00011191 | 0.00011632 | 2Hz10m | 62 | 0.0099994 | 1e-07 | -1.1907e-07 | 5e-07 | 0.0099994 | 1e-07 | -1.1907e-05 | 5.0003e-05 | 2Hz100m10m | -9.2929e-07 | 0.00091562 | 1.0001e-06 | 0.0019441 | 2Hz100m | 2Hz100m | 116 | 0.99991 | 2.7079e-06 | 0.0001 | 0.00023795 | OK
100 | m | 5 | 10 | m | 9.9977 | 0.0019395 | 0.00094106 | 0.00025341 | m | -0.0015865 | 0.0014176 | 0.0010371 | 0.0019447 | m | OK | 0.1 | 5 | Hz | 0.01 | 0.001 | 0.0099993 | 2.4024e-06 | -9.606e-08 | 1.9612e-06 | 0.0099993 | 2.4024e-06 | -9.6067e-06 | 0.00019613 | 5Hz10m | 63 | 0.009999 | 1e-07 | 9.0385e-07 | 5e-07 | 0.009999 | 1e-07 | 9.0394e-05 | 5.0005e-05 | 5Hz100m10m | -2.1599e-07 | 0.003879 | 9.9991e-07 | 0.00050682 | 5Hz100m | 5Hz100m | 117 | 0.99998 | 4.8047e-06 | 0.0001 | 0.00039544 | OK
100 | m | 10 | 10 | m | 9.9982 | 0.00012446 | -0.0019949 | 9.1985e-05 | m | -0.0014794 | 0.00082635 | -0.0009246 | 0.0012026 | m | OK | 0.1 | 10 | Hz | 0.01 | 0.001 | 0.0099997 | 8.3567e-07 | -1.0703e-06 | 1.2061e-06 | 0.0099997 | 8.3567e-07 | -0.00010703 | 0.00012061 | 10Hz10m | 64 | 0.010001 | 1e-07 | -7.0354e-08 | 5e-07 | 0.010001 | 1e-07 | -7.0349e-06 | 4.9997e-05 | 10Hz100m10m | 9.1346e-07 | 0.00024891 | 9.9995e-07 | 0.00018397 | 10Hz100m | 10Hz100m | 118 | 1.0001 | 1.6713e-06 | 9.9999e-05 | 0.00024636 | OK
100 | m | 20 | 10 | m | 10.0003 | 0.00052623 | -5.507e-06 | 0.0010544 | m | -0.00018133 | 0.00047282 | -0.00029371 | 0.0011248 | m | OK | 0.1 | 20 | Hz | 0.01 | 0.001 | 0.01 | 7.0744e-07 | 2.882e-07 | 1.5417e-06 | 0.01 | 7.0744e-07 | 2.8818e-05 | 0.00015416 | 20Hz10m | 65 | 0.010001 | 1e-07 | 1.2883e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00012882 | 4.9997e-05 | 20Hz100m10m | 5.5052e-08 | 0.0010525 | 1.0001e-06 | 0.0021087 | 20Hz100m | 20Hz100m | 119 | 1 | 1.4149e-06 | 0.0001 | 0.00031235 | OK
100 | m | 50 | 10 | m | 10.0031 | 0.001185 | 0.0018547 | 0.0016496 | m | 0.0012881 | 0.00065976 | -0.0010822 | 0.00031574 | m | OK | 0.1 | 50 | Hz | 0.01 | 0.001 | 0.010002 | 1.3563e-06 | 2.9369e-06 | 1.6796e-06 | 0.010002 | 1.3563e-06 | 0.00029364 | 0.00016793 | 50Hz10m | 66 | 0.010001 | 1e-07 | 3.9368e-06 | 5e-07 | 0.010001 | 1e-07 | 0.00039364 | 4.9994e-05 | 50Hz100m10m | -6.4421e-07 | 0.0023701 | 9.9992e-07 | 0.0032992 | 50Hz100m | 50Hz100m | 120 | 0.99994 | 2.7126e-06 | 9.9999e-05 | 0.00033955 | OK
100 | m | 100 | 10 | m | 10.0033 | 0.0011609 | 0.0099435 | 0.0016495 | m | 0.0013863 | 0.0016884 | 0.0013382 | 0.00068289 | m | OK | 0.1 | 100 | Hz | 0.01 | 0.001 | 0.010002 | 2.049e-06 | 8.6053e-06 | 1.7853e-06 | 0.010002 | 2.049e-06 | 0.00086036 | 0.0001785 | 100Hz10m | 67 | 0.010003 | 1e-07 | 9.6062e-06 | 5e-07 | 0.010003 | 1e-07 | 0.00096036 | 4.9987e-05 | 100Hz100m10m | 7.5569e-07 | 0.0023218 | 1.0009e-06 | 0.0032991 | 100Hz100m | 100Hz100m | 121 | 1.0001 | 4.0981e-06 | 9.9998e-05 | 0.00036047 | OK
100 | m | 200 | 10 | m | 10.0084 | 0.00059584 | 0.014587 | 0.0019943 | m | 0.00097416 | 0.0011512 | -0.0019429 | 0.00022175 | m | OK | 0.1 | 200 | Hz | 0.01 | 0.001 | 0.010007 | 1.2963e-06 | 1.653e-05 | 2.0066e-06 | 0.010007 | 1.2963e-06 | 0.0016518 | 0.00020051 | 200Hz10m | 68 | 0.010008 | 1e-07 | 1.7531e-05 | 5e-07 | 0.010008 | 1e-07 | 0.0017517 | 4.9961e-05 | 200Hz100m10m | 2.9465e-07 | 0.0011917 | 1.0011e-06 | 0.0039887 | 200Hz100m | 200Hz100m | 122 | 1 | 2.5925e-06 | 9.9984e-05 | 0.00040413 | OK
100 | m | 500 | 10 | m | 10.0294 | 0.0012224 | 0.047566 | 0.00022451 | m | -0.0015026 | 0.0015324 | 0.0011363 | 0.00016539 | m | OK | 0.1 | 500 | Hz | 0.01 | 0.001 | 0.010031 | 1.9602e-06 | 4.643e-05 | 2.7885e-07 | 0.010031 | 1.9602e-06 | 0.0046287 | 2.7813e-05 | 500Hz10m | 69 | 0.010031 | 1.5811e-07 | 4.7435e-05 | 5e-07 | 0.010031 | 1.5813e-07 | 0.0047286 | 4.9843e-05 | 500Hz100m10m | 3.8497e-07 | 0.0024448 | 1.0047e-06 | 0.00044903 | 500Hz100m | 500Hz100m | 123 | 1 | 3.9204e-06 | 9.9977e-05 | 7.4691e-05 | OK
100 | m | 1000 | 10 | m | 10.0841 | 0.00021512 | 0.090011 | 0.0018265 | m | 0.0011165 | 6.0662e-05 | 0.00090749 | 0.0012769 | m | OK | 0.1 | 1 | kHz | 0.01 | 0.001 | 0.010083 | 2.2351e-07 | 8.9103e-05 | 2.2285e-06 | 0.010083 | 2.2436e-07 | 0.0088368 | 0.000221 | 1kHz10m | 70 | 0.010083 | 4.4721e-07 | 9.0108e-05 | 1e-06 | 0.010083 | 4.4729e-07 | 0.0089367 | 9.9173e-05 | 1kHz100m10m | -2.8772e-07 | 0.00043023 | 1.0046e-06 | 0.0036529 | 1kHz100m | 1kHz100m | 124 | 0.99997 | 4.4875e-07 | 9.9884e-05 | 0.000453 | OK
100 | m | 2000 | 10 | m | 10.234 | 0.001286 | 0.18386 | 0.00081744 | m | 0.00045609 | 0.00019602 | 0.0010238 | 0.0016843 | m | OK | 0.1 | 2 | kHz | 0.01 | 0.001 | 0.010234 | 1.3008e-06 | 0.00018284 | 1.8721e-06 | 0.010235 | 1.3011e-06 | 0.017865 | 0.0001829 | 2kHz10m | 71 | 0.010234 | 1.2649e-06 | 0.00018387 | 2e-06 | 0.010236 | 1.2652e-06 | 0.017964 | 0.00019538 | 2kHz100m10m | 4.63e-07 | 0.002572 | 1.0273e-06 | 0.0016349 | 2kHz100m | 2kHz100m | 125 | 1 | 2.6022e-06 | 9.9542e-05 | 0.0004147 | OK
100 | m | 5000 | 10 | m | 10.9273 | 0.0017421 | 0.45542 | 0.0019871 | m | -1.183e-05 | 0.00015789 | 0.00046003 | 0.0019171 | m | OK | 0.1 | 5 | kHz | 0.01 | 0.001 | 0.010927 | 1.7492e-06 | 0.00045496 | 2.7611e-06 | 0.010937 | 1.7515e-06 | 0.041611 | 0.00025233 | 5kHz10m | 72 | 0.010928 | 5e-06 | 0.00045609 | 5e-06 | 0.010938 | 5e-06 | 0.041711 | 0.00045714 | 5kHz100m10m | 7.5599e-07 | 0.0034842 | 1.1254e-06 | 0.0039742 | 5kHz100m | 5kHz100m | 126 | 1.0001 | 3.5034e-06 | 9.9926e-05 | 0.00068093 | OK
100 | m | 0.01 | 100 | m | 100.0168 | 0.0023875 | -0.022076 | 0.0007995 | m | 0.00081364 | 0.001913 | 0.00043867 | 0.00091466 | m | OK | 0.1 | 10 | mHz | 0.1 | 0.001 | 0.10002 | 3.0594e-06 | -2.2515e-05 | 1.2148e-06 | 0.10002 | 3.0594e-06 | -0.00022511 | 1.2146e-05 | 10mHz100m | 73 | 0.10001 | 1e-06 | -2.0285e-05 | 5e-06 | 0.10001 | 1e-06 | -0.00020282 | 4.9993e-05 | 10mHz100m100m | -2.1481e-06 | 0.004775 | 2.2302e-06 | 0.001599 | 10mHz100m | 10mHz100m | 109 | 0.99998 | 6.1195e-06 | 2.2295e-05 | 5.5583e-05 | OK
100 | m | 0.02 | 100 | m | 99.9884 | 0.0013728 | -0.0018258 | 0.00083705 | m | 0.00081551 | 0.0005896 | -0.0011228 | 0.00012085 | m | OK | 0.1 | 20 | mHz | 0.1 | 0.001 | 0.099988 | 1.4941e-06 | -7.0301e-07 | 8.4573e-07 | 0.099988 | 1.4941e-06 | -7.0309e-06 | 8.4584e-06 | 20mHz100m | 74 | 0.099994 | 1e-06 | -5.5451e-06 | 5e-06 | 0.099994 | 1e-06 | -5.5454e-05 | 5.0003e-05 | 20mHz100m100m | 6.6645e-06 | 0.0027457 | -4.842e-06 | 0.0016741 | 20mHz100m | 20mHz100m | 110 | 1.0001 | 2.9899e-06 | -4.8423e-05 | 5.2787e-05 | OK
100 | m | 0.05 | 100 | m | 99.9933 | 0.0017397 | -0.018042 | 0.00054431 | m | -1.1857e-05 | 0.0003151 | -0.00079281 | 0.0019275 | m | OK | 0.1 | 50 | mHz | 0.1 | 0.001 | 0.099993 | 1.768e-06 | -1.7249e-05 | 2.0028e-06 | 0.099993 | 1.768e-06 | -0.00017251 | 2.003e-05 | 50mHz100m | 75 | 0.099997 | 1e-06 | -9.6195e-06 | 5e-06 | 0.099997 | 1e-06 | -9.6199e-05 | 5.0002e-05 | 50mHz100m100m | 3.3757e-06 | 0.0034795 | 7.6298e-06 | 0.0010886 | 50mHz100m | 50mHz100m | 111 | 1 | 3.5375e-06 | 7.6307e-05 | 6.407e-05 | OK
100 | m | 0.1 | 100 | m | 100.0002 | 0.0014058 | -0.0028062 | 0.00049525 | m | -0.00024961 | 3.2239e-05 | 0.0016069 | 0.0017914 | m | OK | 0.1 | 100 | mHz | 0.1 | 0.001 | 0.1 | 1.4062e-06 | -4.4131e-06 | 1.8586e-06 | 0.1 | 1.4062e-06 | -4.413e-05 | 1.8586e-05 | 100mHz100m | 76 | 0.1 | 1e-06 | -4.6986e-06 | 5e-06 | 0.1 | 1e-06 | -4.6986e-05 | 5e-05 | 100mHz100m100m | 4.9194e-07 | 0.0028116 | -2.8555e-07 | 0.00099051 | 100mHz100m | 100mHz100m | 112 | 1 | 2.8142e-06 | -2.8553e-06 | 6.2304e-05 | OK
100 | m | 0.2 | 100 | m | 100.013 | 0.0021262 | -0.012123 | 0.00021143 | m | 0.00019673 | 0.0016034 | 4.3007e-05 | 0.00022675 | m | OK | 0.1 | 200 | mHz | 0.1 | 0.001 | 0.10001 | 2.663e-06 | -1.2166e-05 | 3.1003e-07 | 0.10001 | 2.663e-06 | -0.00012164 | 3.0999e-06 | 200mHz100m | 77 | 0.10001 | 1e-06 | -3.7512e-06 | 5e-06 | 0.10001 | 1e-06 | -3.7508e-05 | 4.9995e-05 | 200mHz100m100m | -3.2931e-06 | 0.0042524 | 8.4146e-06 | 0.00042289 | 200mHz100m | 200mHz100m | 113 | 0.99997 | 5.327e-06 | 8.4134e-05 | 5.0378e-05 | OK
100 | m | 0.5 | 100 | m | 100.0058 | 0.002236 | 0.012813 | 0.00070673 | m | 0.0014974 | 0.00050447 | -0.0017384 | 0.0015205 | m | OK | 0.1 | 500 | mHz | 0.1 | 0.001 | 0.1 | 2.2922e-06 | 1.4551e-05 | 1.6767e-06 | 0.1 | 2.2922e-06 | 0.00014551 | 1.6766e-05 | 500mHz100m | 78 | 0.099995 | 1e-06 | 1.1572e-05 | 5e-06 | 0.099995 | 1e-06 | 0.00011573 | 5.0003e-05 | 500mHz100m100m | -9.3361e-06 | 0.004472 | -2.9789e-06 | 0.0014135 | 500mHz100m | 500mHz100m | 114 | 0.99991 | 4.5855e-06 | -2.9777e-05 | 6.0205e-05 | OK
100 | m | 1 | 100 | m | 100.0113 | 0.00086533 | 0.0021442 | 4.6706e-05 | m | -5.2384e-05 | 0.0012449 | -0.00055709 | 0.0016174 | m | OK | 0.1 | 1 | Hz | 0.1 | 0.001 | 0.10001 | 1.5161e-06 | 2.7013e-06 | 1.618e-06 | 0.10001 | 1.5161e-06 | 2.701e-05 | 1.6179e-05 | 1Hz100m | 79 | 0.10001 | 1e-06 | 1.1529e-05 | 5e-06 | 0.10001 | 1e-06 | 0.00011528 | 4.9994e-05 | 1Hz100m100m | 3.0133e-07 | 0.0017307 | 8.828e-06 | 9.3545e-05 | 1Hz100m | 1Hz100m | 115 | 1 | 3.0339e-06 | 8.8269e-05 | 5.9552e-05 | OK
100 | m | 2 | 100 | m | 100.0129 | 0.0013047 | -0.0033478 | 0.0016641 | m | 4.1177e-05 | 0.00026126 | 0.0017223 | 0.00063326 | m | OK | 0.1 | 2 | Hz | 0.1 | 0.001 | 0.10001 | 1.3306e-06 | -5.0701e-06 | 1.7805e-06 | 0.10001 | 1.3306e-06 | -5.0694e-05 | 1.7803e-05 | 2Hz100m | 80 | 0.10001 | 1e-06 | -1.1457e-05 | 5e-06 | 0.10001 | 1e-06 | -0.00011457 | 4.9996e-05 | 2Hz100m100m | -5.2275e-06 | 0.0026094 | -6.3874e-06 | 0.0033282 | 2Hz100m | 2Hz100m | 116 | 0.99995 | 2.663e-06 | -6.3872e-05 | 6.1379e-05 | OK
100 | m | 5 | 100 | m | 100.0004 | 0.0020113 | 5.1302e-05 | 0.00060119 | m | -0.0012679 | 0.00057024 | -0.0010675 | 0.0012978 | m | OK | 0.1 | 5 | Hz | 0.1 | 0.001 | 0.1 | 2.0906e-06 | 1.1188e-06 | 1.4303e-06 | 0.1 | 2.0906e-06 | 1.1188e-05 | 1.4302e-05 | 5Hz100m | 81 | 0.1 | 1e-06 | 3.4954e-06 | 5e-06 | 0.1 | 1e-06 | 3.4954e-05 | 5e-05 | 5Hz100m100m | -1.552e-06 | 0.0040226 | 2.3766e-06 | 0.0012024 | 5Hz100m | 5Hz100m | 117 | 0.99998 | 4.1824e-06 | 2.3766e-05 | 5.7604e-05 | OK
100 | m | 10 | 100 | m | 99.9864 | 0.0019255 | 0.0086103 | 0.00085754 | m | -0.00060155 | 0.00071913 | 0.0016898 | 0.00016996 | m | OK | 0.1 | 10 | Hz | 0.1 | 0.001 | 0.099987 | 2.0554e-06 | 6.9205e-06 | 8.7422e-07 | 0.099987 | 2.0554e-06 | 6.9214e-05 | 8.7434e-06 | 10Hz100m | 82 | 0.099993 | 1e-06 | -2.5665e-06 | 5e-06 | 0.099993 | 1e-06 | -2.5666e-05 | 5.0003e-05 | 10Hz100m100m | 6.3813e-06 | 0.0038509 | -9.4869e-06 | 0.0017151 | 10Hz100m | 10Hz100m | 118 | 1.0001 | 4.1119e-06 | -9.488e-05 | 5.2973e-05 | OK
100 | m | 20 | 100 | m | 100.0023 | 0.0020061 | 0.005818 | 0.0015702 | m | -0.0010987 | 0.00021933 | 0.0015877 | 0.00041379 | m | OK | 0.1 | 20 | Hz | 0.1 | 0.001 | 0.1 | 2.018e-06 | 4.2303e-06 | 1.6238e-06 | 0.1 | 2.018e-06 | 4.2302e-05 | 1.6238e-05 | 20Hz100m | 83 | 0.1 | 1e-06 | 1.1555e-05 | 5e-06 | 0.1 | 1e-06 | 0.00011555 | 5e-05 | 20Hz100m100m | -3.2387e-06 | 0.0040121 | 7.3249e-06 | 0.0031404 | 20Hz100m | 20Hz100m | 119 | 0.99997 | 4.0373e-06 | 7.325e-05 | 5.9621e-05 | OK
100 | m | 50 | 100 | m | 100.023 | 0.0016939 | 0.034078 | 0.00028647 | m | 0.001702 | 0.0018789 | -6.7182e-06 | 0.00010468 | m | OK | 0.1 | 50 | Hz | 0.1 | 0.001 | 0.10002 | 2.5297e-06 | 3.4084e-05 | 3.0499e-07 | 0.10002 | 2.5297e-06 | 0.00034077 | 3.0493e-06 | 50Hz100m | 84 | 0.10002 | 1e-06 | 3.6969e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00036962 | 4.999e-05 | 50Hz100m100m | -1.6148e-06 | 0.0033878 | 2.8843e-06 | 0.00057296 | 50Hz100m | 50Hz100m | 120 | 0.99998 | 5.0604e-06 | 2.8843e-05 | 5.0361e-05 | OK
100 | m | 100 | 100 | m | 100.028 | 0.0013842 | 0.079968 | 0.0019446 | m | 9.7947e-05 | 0.0013299 | -0.00029773 | 0.001132 | m | OK | 0.1 | 100 | Hz | 0.1 | 0.001 | 0.10003 | 1.9195e-06 | 8.0266e-05 | 2.25e-06 | 0.10003 | 1.9195e-06 | 0.00080244 | 2.2494e-05 | 100Hz100m | 85 | 0.10002 | 1e-06 | 8.5079e-05 | 5e-06 | 0.10002 | 1e-06 | 0.00085059 | 4.9989e-05 | 100Hz100m100m | -5.0441e-06 | 0.0027683 | 4.8126e-06 | 0.0038891 | 100Hz100m | 100Hz100m | 121 | 0.99995 | 3.8403e-06 | 4.8155e-05 | 6.7252e-05 | OK
100 | m | 200 | 100 | m | 100.0662 | 0.0023776 | 0.17547 | 0.0015126 | m | 0.0013513 | 0.00073148 | 0.0012551 | 0.0016645 | m | OK | 0.1 | 200 | Hz | 0.1 | 0.001 | 0.10006 | 2.4876e-06 | 0.00017422 | 2.2492e-06 | 0.10007 | 2.4876e-06 | 0.0017411 | 2.2477e-05 | 200Hz100m | 86 | 0.10007 | 1e-06 | 0.00016547 | 5e-06 | 0.10007 | 1e-06 | 0.0016536 | 4.9966e-05 | 200Hz100m100m | 2.8943e-06 | 0.0047552 | -8.7494e-06 | 0.0030253 | 200Hz100m | 200Hz100m | 122 | 1 | 4.9761e-06 | -8.7485e-05 | 6.7212e-05 | OK
100 | m | 500 | 100 | m | 100.3197 | 0.0017864 | 0.44872 | 0.0003277 | m | 0.00085752 | 0.0010432 | -0.0016526 | 0.0012038 | m | OK | 0.1 | 500 | Hz | 0.1 | 0.001 | 0.10032 | 2.0687e-06 | 0.00045037 | 1.2476e-06 | 0.10032 | 2.0687e-06 | 0.0044893 | 1.2437e-05 | 500Hz100m | 87 | 0.10031 | 1.5811e-06 | 0.00045375 | 5e-06 | 0.10031 | 1.5813e-06 | 0.0045233 | 4.9843e-05 | 500Hz100m100m | -6.7771e-06 | 0.0035729 | 3.3804e-06 | 0.00065542 | 500Hz100m | 500Hz100m | 123 | 0.99993 | 4.1405e-06 | 3.4002e-05 | 5.5705e-05 | OK
100 | m | 1000 | 100 | m | 100.8184 | 0.0020286 | 0.91895 | 8.5535e-05 | m | -0.0008507 | 0.00061474 | -0.0012897 | 0.0014619 | m | OK | 0.1 | 1 | kHz | 0.1 | 0.001 | 0.10082 | 2.1197e-06 | 0.00092024 | 1.4644e-06 | 0.10082 | 2.1197e-06 | 0.0091274 | 1.4525e-05 | 1kHz100m | 88 | 0.10083 | 4.4721e-06 | 0.00091764 | 1e-05 | 0.10083 | 4.4729e-06 | 0.0091008 | 9.9172e-05 | 1kHz100m100m | 7.7068e-06 | 0.0040572 | -2.6087e-06 | 0.00017136 | 1kHz100m | 1kHz100m | 124 | 1.0001 | 4.2632e-06 | -2.6568e-05 | 0.00010334 | OK
100 | m | 2000 | 100 | m | 102.3546 | 0.0021432 | 1.8435 | 0.00021502 | m | 0.00080972 | 0.00021036 | 0.00018396 | 0.0017533 | m | OK | 0.1 | 2 | kHz | 0.1 | 0.001 | 0.10235 | 2.1535e-06 | 0.0018433 | 1.7664e-06 | 0.10237 | 2.1534e-06 | 0.018007 | 1.7256e-05 | 2kHz100m | 89 | 0.10234 | 1.2649e-05 | 0.001835 | 2e-05 | 0.10236 | 1.2652e-05 | 0.017927 | 0.00019537 | 2kHz100m100m | -9.4999e-06 | 0.0042864 | -8.3259e-06 | 0.0004305 | 2kHz100m | 2kHz100m | 125 | 0.99991 | 4.4973e-06 | -7.9654e-05 | 0.00019839 | OK
100 | m | 5000 | 100 | m | 109.3057 | 0.0012781 | 4.5435 | 0.00042342 | m | -0.0016987 | 0.0018891 | 0.0016489 | 0.001494 | m | OK | 0.1 | 5 | kHz | 0.1 | 0.001 | 0.10931 | 2.2808e-06 | 0.0045419 | 1.5529e-06 | 0.1094 | 2.2797e-06 | 0.041528 | 1.4208e-05 | 5kHz100m | 90 | 0.1093 | 5e-05 | 0.0045518 | 5e-05 | 0.1094 | 5e-05 | 0.04162 | 0.00045705 | 5kHz100m100m | -4.1147e-06 | 0.0025566 | 9.9168e-06 | 0.00084832 | 5kHz100m | 5kHz100m | 126 | 0.99997 | 7.1211e-06 | 9.2132e-05 | 0.00045793 | OK

## Ref Z
This is sheet containing calibration data of reference standards to be used for calibration of RLC meter.
1) Select resistance standards to be used for calibration. Typical procedure is to pick two resistors for each physical range of instrument. That is one for bottom of the range and one for full range. Exception is the lowest range, where the bottom of the range is expected to be zero (SHORT standard).
Example:
Range | Standard
Ohm | Ohm
1m | SHORT
1m
3m | 1m
3m
10m | 3m
10m
The short can be considered "ideal", so the values of impedance may remain zero and only the uncertainty estimates may be used to include them to final uncertainty budget.
2) Select frequencies at which you want to measure and have available ref-Z data. This sheet contains selection made in scope of LiBforSecUse project for EIS spectroscopy.
3) Fill in Rs-Xs data and associated expanded uncertainties (k = 2) for each standard. Fill in "Nominal Z" with nominal value of standard, i.e. 0m for SHORT, 1m for roughly 1mOhm standard, etc. The eventual residual value of SHORT standard (nominal 0 Ohm) is given as absolute value if available or at least some uncertainty should be estimated. The values of all other standards are given relative to SHORT or NULL standard against which they were calibrated. This should be mentioned in the calibration certificate of the standards. If the information whether it is absolute impedance or relative to SHORT is not available, then expand uncertainty by estimate of worst case residual impedance of your SHORT standard.
4) Fill in the list of chosen nominal calibration impedances to sheet "Ref Z list".
Notes: Use Rs-Xs suffix "m" for milliOhms or nothing for Ohms. The sheet will recalculate. The same applies for nominal Z. Add or remove lines to/from bottom of the tables, but always copy formulas in whole width of the sheet (the gray area is referenced by other sheets!) Order of frequencies and standards and ranges is irrelevant. Other sheets will identify positions automaticaly.
Calibration data of ref. Z standards
f | Nominal Z |  | Rs | U(Rs) | Xs | U(Xs) | Rs-Xs unit | f |  | Nom Z | Rs-Xs mult | Rs | U(Rs) | Xs | U(Xs) | tag | #octave_header#
Hz | Ohm |  | Ohm | Ohm | Ohm | Ohm | Ohm | - |  | Ohm | - | Ohm | Ohm | Ohm | Ohm | -
0.01 | 0 | m | 0 | 0 | 0 | 0 | m | 10 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 10mHz0m
0.02 | 0 | m | 0 | 0 | 0 | 0 | m | 20 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 20mHz0m
0.05 | 0 | m | 0 | 0 | 0 | 0 | m | 50 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 50mHz0m
0.1 | 0 | m | 0 | 0 | 0 | 0 | m | 100 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 100mHz0m
0.2 | 0 | m | 0 | 0 | 0 | 0 | m | 200 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 200mHz0m
0.5 | 0 | m | 0 | 0 | 0 | 0 | m | 500 | mHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 500mHz0m
1 | 0 | m | 0 | 0 | 0 | 0 | m | 1 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 1Hz0m
2 | 0 | m | 0 | 0 | 0 | 0 | m | 2 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 2Hz0m
5 | 0 | m | 0 | 0 | 0 | 0 | m | 5 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 5Hz0m
10 | 0 | m | 0 | 0 | 0 | 0 | m | 10 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 10Hz0m
20 | 0 | m | 0 | 0 | 0 | 0 | m | 20 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 20Hz0m
50 | 0 | m | 0 | 0 | 0 | 0 | m | 50 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 50Hz0m
100 | 0 | m | 0 | 0 | 0 | 0 | m | 100 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 100Hz0m
200 | 0 | m | 0 | 0 | 0 | 0 | m | 200 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 200Hz0m
500 | 0 | m | 0 | 0 | 0 | 0 | m | 500 | Hz | 0 | 0.001 | 0 | 0 | 0 | 0 | 500Hz0m
1000 | 0 | m | 0 | 0 | 0 | 0 | m | 1 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 1kHz0m
2000 | 0 | m | 0 | 0 | 0 | 0 | m | 2 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 2kHz0m
5000 | 0 | m | 0 | 0 | 0 | 0 | m | 5 | kHz | 0 | 0.001 | 0 | 0 | 0 | 0 | 5kHz0m
0.01 | 1 | m | 1.0001 | 1e-05 | -7.1568e-05 | 5e-05 | m | 10 | mHz | 0.001 | 0.001 | 0.0010001 | 1e-08 | -7.1568e-08 | 5e-08 | 10mHz1m
0.02 | 1 | m | 1 | 1e-05 | -0.00013225 | 5e-05 | m | 20 | mHz | 0.001 | 0.001 | 0.001 | 1e-08 | -1.3225e-07 | 5e-08 | 20mHz1m
0.05 | 1 | m | 0.99993 | 1e-05 | -8.1888e-05 | 5e-05 | m | 50 | mHz | 0.001 | 0.001 | 0.00099993 | 1e-08 | -8.1888e-08 | 5e-08 | 50mHz1m
0.1 | 1 | m | 1.0001 | 1e-05 | -9.2355e-05 | 5e-05 | m | 100 | mHz | 0.001 | 0.001 | 0.0010001 | 1e-08 | -9.2355e-08 | 5e-08 | 100mHz1m
0.2 | 1 | m | 1 | 1e-05 | 6.4289e-05 | 5e-05 | m | 200 | mHz | 0.001 | 0.001 | 0.001 | 1e-08 | 6.4289e-08 | 5e-08 | 200mHz1m
0.5 | 1 | m | 1.0001 | 1e-05 | 0.00015698 | 5e-05 | m | 500 | mHz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 1.5698e-07 | 5e-08 | 500mHz1m
1 | 1 | m | 1.0002 | 1e-05 | -1.4965e-05 | 5e-05 | m | 1 | Hz | 0.001 | 0.001 | 0.0010002 | 1e-08 | -1.4965e-08 | 5e-08 | 1Hz1m
2 | 1 | m | 0.99992 | 1e-05 | -0.00017685 | 5e-05 | m | 2 | Hz | 0.001 | 0.001 | 0.00099992 | 1e-08 | -1.7685e-07 | 5e-08 | 2Hz1m
5 | 1 | m | 1.0001 | 1e-05 | 0.0001958 | 5e-05 | m | 5 | Hz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 1.958e-07 | 5e-08 | 5Hz1m
10 | 1 | m | 1.0002 | 1e-05 | 0.00014427 | 5e-05 | m | 10 | Hz | 0.001 | 0.001 | 0.0010002 | 1e-08 | 1.4427e-07 | 5e-08 | 10Hz1m
20 | 1 | m | 1.0001 | 1e-05 | 0.00015313 | 5e-05 | m | 20 | Hz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 1.5313e-07 | 5e-08 | 20Hz1m
50 | 1 | m | 1.0001 | 1e-05 | 0.00034996 | 5e-05 | m | 50 | Hz | 0.001 | 0.001 | 0.0010001 | 1e-08 | 3.4996e-07 | 5e-08 | 50Hz1m
100 | 1 | m | 1.0003 | 1e-05 | 0.00090589 | 5e-05 | m | 100 | Hz | 0.001 | 0.001 | 0.0010003 | 1e-08 | 9.0589e-07 | 5e-08 | 100Hz1m
200 | 1 | m | 1.0007 | 1e-05 | 0.001768 | 5e-05 | m | 200 | Hz | 0.001 | 0.001 | 0.0010007 | 1e-08 | 1.768e-06 | 5e-08 | 200Hz1m
500 | 1 | m | 1.0029 | 1.5811e-05 | 0.0046098 | 5e-05 | m | 500 | Hz | 0.001 | 0.001 | 0.0010029 | 1.5811e-08 | 4.6098e-06 | 5e-08 | 500Hz1m
1000 | 1 | m | 1.0083 | 4.4721e-05 | 0.0092383 | 0.0001 | m | 1 | kHz | 0.001 | 0.001 | 0.0010083 | 4.4721e-08 | 9.2383e-06 | 1e-07 | 1kHz1m
2000 | 1 | m | 1.0236 | 0.00012649 | 0.018381 | 0.0002 | m | 2 | kHz | 0.001 | 0.001 | 0.0010236 | 1.2649e-07 | 1.8381e-05 | 2e-07 | 2kHz1m
5000 | 1 | m | 1.0929 | 0.0005 | 0.045708 | 0.0005 | m | 5 | kHz | 0.001 | 0.001 | 0.0010929 | 5e-07 | 4.5708e-05 | 5e-07 | 5kHz1m
0.01 | 3 | m | 2.9997 | 3e-05 | 3.7684e-05 | 0.00015 | m | 10 | mHz | 0.003 | 0.001 | 0.0029997 | 3e-08 | 3.7684e-08 | 1.5e-07 | 10mHz3m
0.02 | 3 | m | 2.9998 | 3e-05 | -2.539e-05 | 0.00015 | m | 20 | mHz | 0.003 | 0.001 | 0.0029998 | 3e-08 | -2.539e-08 | 1.5e-07 | 20mHz3m
0.05 | 3 | m | 3.0001 | 3e-05 | 0.00032533 | 0.00015 | m | 50 | mHz | 0.003 | 0.001 | 0.0030001 | 3e-08 | 3.2533e-07 | 1.5e-07 | 50mHz3m
0.1 | 3 | m | 2.9996 | 3e-05 | 0.00080053 | 0.00015 | m | 100 | mHz | 0.003 | 0.001 | 0.0029996 | 3e-08 | 8.0053e-07 | 1.5e-07 | 100mHz3m
0.2 | 3 | m | 3 | 3e-05 | -4.2224e-06 | 0.00015 | m | 200 | mHz | 0.003 | 0.001 | 0.003 | 3e-08 | -4.2224e-09 | 1.5e-07 | 200mHz3m
0.5 | 3 | m | 2.9998 | 3e-05 | 0.0002075 | 0.00015 | m | 500 | mHz | 0.003 | 0.001 | 0.0029998 | 3e-08 | 2.075e-07 | 1.5e-07 | 500mHz3m
1 | 3 | m | 3 | 3e-05 | 0.0012324 | 0.00015 | m | 1 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | 1.2324e-06 | 1.5e-07 | 1Hz3m
2 | 3 | m | 3 | 3e-05 | -1.4799e-05 | 0.00015 | m | 2 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | -1.4799e-08 | 1.5e-07 | 2Hz3m
5 | 3 | m | 3 | 3e-05 | 0.00050723 | 0.00015 | m | 5 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | 5.0723e-07 | 1.5e-07 | 5Hz3m
10 | 3 | m | 3.0001 | 3e-05 | 9.6095e-05 | 0.00015 | m | 10 | Hz | 0.003 | 0.001 | 0.0030001 | 3e-08 | 9.6095e-08 | 1.5e-07 | 10Hz3m
20 | 3 | m | 3 | 3e-05 | 0.0011185 | 0.00015 | m | 20 | Hz | 0.003 | 0.001 | 0.003 | 3e-08 | 1.1185e-06 | 1.5e-07 | 20Hz3m
50 | 3 | m | 3.0003 | 3e-05 | 0.0019181 | 0.00015 | m | 50 | Hz | 0.003 | 0.001 | 0.0030003 | 3e-08 | 1.9181e-06 | 1.5e-07 | 50Hz3m
100 | 3 | m | 3.0012 | 3e-05 | 0.0026469 | 0.00015 | m | 100 | Hz | 0.003 | 0.001 | 0.0030012 | 3e-08 | 2.6469e-06 | 1.5e-07 | 100Hz3m
200 | 3 | m | 3.0021 | 3e-05 | 0.0055775 | 0.00015 | m | 200 | Hz | 0.003 | 0.001 | 0.0030021 | 3e-08 | 5.5775e-06 | 1.5e-07 | 200Hz3m
500 | 3 | m | 3.0089 | 4.7434e-05 | 0.013893 | 0.00015 | m | 500 | Hz | 0.003 | 0.001 | 0.0030089 | 4.7434e-08 | 1.3893e-05 | 1.5e-07 | 500Hz3m
1000 | 3 | m | 3.0244 | 0.00013416 | 0.02756 | 0.0003 | m | 1 | kHz | 0.003 | 0.001 | 0.0030244 | 1.3416e-07 | 2.756e-05 | 3e-07 | 1kHz3m
2000 | 3 | m | 3.0706 | 0.00037947 | 0.054982 | 0.0006 | m | 2 | kHz | 0.003 | 0.001 | 0.0030706 | 3.7947e-07 | 5.4982e-05 | 6e-07 | 2kHz3m
5000 | 3 | m | 3.2786 | 0.0015 | 0.13747 | 0.0015 | m | 5 | kHz | 0.003 | 0.001 | 0.0032786 | 1.5e-06 | 0.00013747 | 1.5e-06 | 5kHz3m
0.01 | 10 | m | 10.0011 | 0.0001 | -0.0012341 | 0.0005 | m | 10 | mHz | 0.01 | 0.001 | 0.010001 | 1e-07 | -1.2341e-06 | 5e-07 | 10mHz10m
0.02 | 10 | m | 10.0003 | 0.0001 | -0.0016219 | 0.0005 | m | 20 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | -1.6219e-06 | 5e-07 | 20mHz10m
0.05 | 10 | m | 10.0004 | 0.0001 | 7.5712e-05 | 0.0005 | m | 50 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | 7.5712e-08 | 5e-07 | 50mHz10m
0.1 | 10 | m | 10.0001 | 0.0001 | 0.00043356 | 0.0005 | m | 100 | mHz | 0.01 | 0.001 | 0.01 | 1e-07 | 4.3356e-07 | 5e-07 | 100mHz10m
0.2 | 10 | m | 9.9993 | 0.0001 | 0.0011321 | 0.0005 | m | 200 | mHz | 0.01 | 0.001 | 0.0099993 | 1e-07 | 1.1321e-06 | 5e-07 | 200mHz10m
0.5 | 10 | m | 9.9999 | 0.0001 | 0.00052604 | 0.0005 | m | 500 | mHz | 0.01 | 0.001 | 0.0099999 | 1e-07 | 5.2604e-07 | 5e-07 | 500mHz10m
1 | 10 | m | 10.0002 | 0.0001 | 0.00042024 | 0.0005 | m | 1 | Hz | 0.01 | 0.001 | 0.01 | 1e-07 | 4.2024e-07 | 5e-07 | 1Hz10m
2 | 10 | m | 9.9994 | 0.0001 | -0.00011907 | 0.0005 | m | 2 | Hz | 0.01 | 0.001 | 0.0099994 | 1e-07 | -1.1907e-07 | 5e-07 | 2Hz10m
5 | 10 | m | 9.999 | 0.0001 | 0.00090385 | 0.0005 | m | 5 | Hz | 0.01 | 0.001 | 0.009999 | 1e-07 | 9.0385e-07 | 5e-07 | 5Hz10m
10 | 10 | m | 10.0006 | 0.0001 | -7.0354e-05 | 0.0005 | m | 10 | Hz | 0.01 | 0.001 | 0.010001 | 1e-07 | -7.0354e-08 | 5e-07 | 10Hz10m
20 | 10 | m | 10.0005 | 0.0001 | 0.0012883 | 0.0005 | m | 20 | Hz | 0.01 | 0.001 | 0.010001 | 1e-07 | 1.2883e-06 | 5e-07 | 20Hz10m
50 | 10 | m | 10.0012 | 0.0001 | 0.0039368 | 0.0005 | m | 50 | Hz | 0.01 | 0.001 | 0.010001 | 1e-07 | 3.9368e-06 | 5e-07 | 50Hz10m
100 | 10 | m | 10.0027 | 0.0001 | 0.0096062 | 0.0005 | m | 100 | Hz | 0.01 | 0.001 | 0.010003 | 1e-07 | 9.6062e-06 | 5e-07 | 100Hz10m
200 | 10 | m | 10.0077 | 0.0001 | 0.017531 | 0.0005 | m | 200 | Hz | 0.01 | 0.001 | 0.010008 | 1e-07 | 1.7531e-05 | 5e-07 | 200Hz10m
500 | 10 | m | 10.0313 | 0.00015811 | 0.047435 | 0.0005 | m | 500 | Hz | 0.01 | 0.001 | 0.010031 | 1.5811e-07 | 4.7435e-05 | 5e-07 | 500Hz10m
1000 | 10 | m | 10.0827 | 0.00044721 | 0.090108 | 0.001 | m | 1 | kHz | 0.01 | 0.001 | 0.010083 | 4.4721e-07 | 9.0108e-05 | 1e-06 | 1kHz10m
2000 | 10 | m | 10.234 | 0.0012649 | 0.18387 | 0.002 | m | 2 | kHz | 0.01 | 0.001 | 0.010234 | 1.2649e-06 | 0.00018387 | 2e-06 | 2kHz10m
5000 | 10 | m | 10.9281 | 0.005 | 0.45609 | 0.005 | m | 5 | kHz | 0.01 | 0.001 | 0.010928 | 5e-06 | 0.00045609 | 5e-06 | 5kHz10m
0.01 | 100 | m | 100.0138 | 0.001 | -0.020285 | 0.005 | m | 10 | mHz | 0.1 | 0.001 | 0.10001 | 1e-06 | -2.0285e-05 | 5e-06 | 10mHz100m
0.02 | 100 | m | 99.9942 | 0.001 | -0.0055451 | 0.005 | m | 20 | mHz | 0.1 | 0.001 | 0.099994 | 1e-06 | -5.5451e-06 | 5e-06 | 20mHz100m
0.05 | 100 | m | 99.9967 | 0.001 | -0.0096195 | 0.005 | m | 50 | mHz | 0.1 | 0.001 | 0.099997 | 1e-06 | -9.6195e-06 | 5e-06 | 50mHz100m
0.1 | 100 | m | 100.001 | 0.001 | -0.0046986 | 0.005 | m | 100 | mHz | 0.1 | 0.001 | 0.1 | 1e-06 | -4.6986e-06 | 5e-06 | 100mHz100m
0.2 | 100 | m | 100.0095 | 0.001 | -0.0037512 | 0.005 | m | 200 | mHz | 0.1 | 0.001 | 0.10001 | 1e-06 | -3.7512e-06 | 5e-06 | 200mHz100m
0.5 | 100 | m | 99.995 | 0.001 | 0.011572 | 0.005 | m | 500 | mHz | 0.1 | 0.001 | 0.099995 | 1e-06 | 1.1572e-05 | 5e-06 | 500mHz100m
1 | 100 | m | 100.0117 | 0.001 | 0.011529 | 0.005 | m | 1 | Hz | 0.1 | 0.001 | 0.10001 | 1e-06 | 1.1529e-05 | 5e-06 | 1Hz100m
2 | 100 | m | 100.0076 | 0.001 | -0.011457 | 0.005 | m | 2 | Hz | 0.1 | 0.001 | 0.10001 | 1e-06 | -1.1457e-05 | 5e-06 | 2Hz100m
5 | 100 | m | 100.0001 | 0.001 | 0.0034954 | 0.005 | m | 5 | Hz | 0.1 | 0.001 | 0.1 | 1e-06 | 3.4954e-06 | 5e-06 | 5Hz100m
10 | 100 | m | 99.9933 | 0.001 | -0.0025665 | 0.005 | m | 10 | Hz | 0.1 | 0.001 | 0.099993 | 1e-06 | -2.5665e-06 | 5e-06 | 10Hz100m
20 | 100 | m | 100.0002 | 0.001 | 0.011555 | 0.005 | m | 20 | Hz | 0.1 | 0.001 | 0.1 | 1e-06 | 1.1555e-05 | 5e-06 | 20Hz100m
50 | 100 | m | 100.0197 | 0.001 | 0.036969 | 0.005 | m | 50 | Hz | 0.1 | 0.001 | 0.10002 | 1e-06 | 3.6969e-05 | 5e-06 | 50Hz100m
100 | 100 | m | 100.0228 | 0.001 | 0.085079 | 0.005 | m | 100 | Hz | 0.1 | 0.001 | 0.10002 | 1e-06 | 8.5079e-05 | 5e-06 | 100Hz100m
200 | 100 | m | 100.0678 | 0.001 | 0.16547 | 0.005 | m | 200 | Hz | 0.1 | 0.001 | 0.10007 | 1e-06 | 0.00016547 | 5e-06 | 200Hz100m
500 | 100 | m | 100.3121 | 0.0015811 | 0.45375 | 0.005 | m | 500 | Hz | 0.1 | 0.001 | 0.10031 | 1.5811e-06 | 0.00045375 | 5e-06 | 500Hz100m
1000 | 100 | m | 100.827 | 0.0044721 | 0.91764 | 0.01 | m | 1 | kHz | 0.1 | 0.001 | 0.10083 | 4.4721e-06 | 0.00091764 | 1e-05 | 1kHz100m
2000 | 100 | m | 102.3443 | 0.012649 | 1.835 | 0.02 | m | 2 | kHz | 0.1 | 0.001 | 0.10234 | 1.2649e-05 | 0.001835 | 2e-05 | 2kHz100m
5000 | 100 | m | 109.3033 | 0.05 | 4.5518 | 0.05 | m | 5 | kHz | 0.1 | 0.001 | 0.1093 | 5e-05 | 0.0045518 | 5e-05 | 5kHz100m

## Ref Z list
List of used reference impedances.
Fill in only the standards used in the sheet "Ref Z"! Fill them in order including 0m for NULL or SHORT standard.
Use "m" suffix to identify milliOhms or remove it to enter data in Ohms directly.
Add rows if necessary, but always extend all columns formulas (the gray area).
Edit only the yellow cells. Grays are referenced by other sheets.
¦Z¦ |  | ¦Z¦ | Z tag | #octave_header#
Ohm |  | Ohm | -
0 | m | 0 | 0m
1 | m | 0.001 | 1m
3 | m | 0.003 | 3m
10 | m | 0.01 | 10m
100 | m | 0.1 | 100m